## Cover Page
AlloSeq HCT
Workbook
IFU100-1
Version Number: 3.0
Issue Date: October 2023
ASHCT.1(96)-IVD
| CareDx Pty Ltd |  |  | Qarad BV,
| 20 Collie St. |  |  | Cipalstraat 3
| Fremantle 6160 |  |  | 2440 Geel
| Western Australia |  |  | Belgium
VERSION | REVISION DATE (DDMMMYY ) |  | REVISION DESCRIPTION
3 | 2023-10-12 |  | Workbook update for consistency with AlloSeq HCT Software v2.2.
© 2023 CareDx, Inc. All service marks or trademarks are owned or licensed by CareDx, Inc. or its affiliates. All rights reserved.

## Sample Sheet
[Paste Sample Sheet here]
Text to Columns
Note on A1: Instructions: 1) Copy text from SampleSheet.csv (generated by HCT software) 2) Paste text into cell A1 3) Run Text to Columns Macro OR if macros are disabled use the Data¦Text to Columns tool in excel. Choose commas as a delimiter. 4) Print Plate map from Plate Map tab

## Plate Map
RUN LEVEL INFORMATION
Investigator: |  | 0 |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
Experiment: |  | 0 |  |  |  | NOTE: ORANGE fields must be selected from drop down before printing.
Date: |  | 0 |  |  |  | NOTE: YELLOW fields must be completed before printing.
Description: |  | 0
# Samples: |  | 0
Sample Plate:
Sequencer:
SAMPLE ID AND INDEX PLATE LAYOUT
|  | A701 | A702 | A703 | A704 | A705 | A706 | A707 | A708 | A709 | A710 | A711 | A712
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A501 | A
A502 | B
A503 | C
A504 | D
A505 | E
A506 | F
A507 | G
A508 | H

## Calculations
PREPARING SAMPLES FOR ONE-STEP MULTIPLEX PCR AMPLIFICATION
Use this section to perform calculations to input 10ng DNA for each sample into the PCR amplification.
(1) Enter the Sample_ID in column B or I.
(2) Enter the stock concentration in ng/uL (determined by Qubit High-Sensitivity assay or equivalent) for each sample in column C or J.
(3) The total recommended volume to be prepared in uL is set at 100uL for stock concentration greater than 10ng/uL in column E and L
(4) The required volume of stock sample in uL will be automatically calculated in column F or M.
(5) The required volume of PCR-grade water in uL will be automatically calculated in column G or N.
(6) In a clean tube, combine the appropriate volumes of stock sample and PCR-grade water. Mix well and store on ice while assembling the amplification reaction.
| Enter Sample_ID | Enter stock conc (ng/uL) | Final DNA Concentration (in ng/ul)* | Recommended Final DNA Volume Needed (ul)* | Volume of stock sample (uL) | Volume of PCR-grade H2O (uL) | Enter Sample_ID | Enter stock conc (ng/uL) | Final DNA Concentration (in ng/ul)* | Recommended Final DNA Volume Needed (ul)* | Volume of stock sample (uL) | Volume of PCR-grade H2O (uL)
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
|  |  | 0.625 | 20 | #DIV/0! | #DIV/0! |  |  | 0.625 | 20 | #DIV/0! | #DIV/0!
* This value remains constant

## Index Layouts (hidden)
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
Product Code | A | B | C | D | E | F | G | H | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 2.0 Library_Prep Lookup
Select | [20 empty] | Select index set on 1.0 Sample_Prep tab.
96 Test Kit | A501 | A502 | A503 | A504 | A505 | A506 | A507 | A508 | A701 | A702 | A703 | A704 | A705 | A706 | A707 | A708 | A709 | A710 | A711 | A712 | Indexes: A501, A502, A503, A504, A505, A506, A507, A508, A701, A702, A703, A704, A705, A706, A707, A708, A709, A710, A711, A712

## DropDowns (hidden)
i7_IndexNames | i7_IndexSequences | i5_IndexNames | i5_IndexSequences
A701 | ATCACGAC | A501 | TGAACCTT
A702 | ACAGTGGT | A502 | TGCTAAGT
A703 | CAGATCCA | A503 | TGTTCTCT
A704 | ACAAACGG | A504 | TAAGACAC
A705 | ACCCAGCA | A505 | CTAATCGA
A706 | AACCCCTC | A506 | CTAGAACA
A707 | CCCAACCT | A507 | TAAGTTCC
A708 | CACCACAC | A508 | TAGACCTA
A709 | GAAACCCA
A710 | TGTGACCA
A711 | AGGGTCAA
A712 | AGGAGTGG

## Lib Prep
ALLOSEQ HCT- LIBRARY PREPARATION
| Experiment ID #: |  |  |  | 0
| Test Kit |  |  |  | 96 Test Kit
| # of Samples |  |  |  | 0
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Sample (µL) | Volume for # of Samples Required (µL): | 0 | Preparation Required
|  |  |  | PCR Mix |  |  | -15°C to -25°C BOX 1 | 13 | 0 |  | Thaw and keep on ice.
|  |  |  | PCR Enzyme |  |  | -15°C to -25°C BOX 1 | 0.8 | 0 |  | Keep on ice.
|  |  |  | SNP Primer Pool |  |  | -15°C to -25°C BOX 1 | 2.2 | 0 |  | Thaw and keep on ice.
|  |  |  | HCT Index Plate |  |  | -15°C to -25°C BOX 2 | 8 | 0 |  | Thaw and keep on ice.
|  | 2 | Briefly vortex and quick spin PCR Mix, SNP Primer Pool and the Indices before use. Do not vortex the PCR Enzyme.
|  | 3 | Gather required consumables listed below:
96-Well Half-Skirted Plate,
Microseal B seal,
1.5 mL microcentrifuge tube,
Index Plate Fixture.
LIBRARY PREPARATION PROTOCOL
NOTE: Turn on the thermal cycler instrument to ensure it is at the right temperature after the PCR plate is set up and program the cycling protocol in advance.
Delays of ≥ 1 hour in starting thermal cycling protocol of the PCR plate has shown to decrease Q30 scores.
|  | 1 | Ensure DNA samples to be used in the library preparation are normalized to a final concentration of 0.625 ng/μl and at a minimum volume of 18 μl to allow
pipetting of 16 μl into the final PCR plate.
|  | 2 | Place a fresh PCR plate on ice (if using the TruSeq Index Plate Fixture, keep it on ice).
|  | 3 | Confirm correct index plate orientation with A1 well being the top left. Pierce foil for desired index combination using a pipette tip. Ensure a different pipette tip is used per well pierced. When piercing the well ensure not to distrub the index solution.
|  | 4 | Add 8 ul of premixed i7 and i5 dual index mixtures from the 96 indices plate to the fresh PCR plate on ice using a single P10-P20 or multichannel P10-P20 pipette. Ensure the bottom of the well is touched and aspirate entire volume. Confirm with visual check.
|  | 5 | Add 16 μl of your prepared DNA to each sample well.
|  | 6 | Prepare PCR master mix according to the following volumes:
|  |  |  | Reagent |  | Volume per Sample (µL) | Volume for Number of Samples Required +12% Excess (µL): | 0
|  |  |  | PCR Mix |  | 13 | 0
|  |  |  | PCR Enzyme |  | 0.8 | 0
|  |  |  | SNP Primer Pool |  | 2.2 | 0
|  |  |  | Total |  | 16 | 0
|  | 7 | Mix the above master mix by vortexing for 5 seconds and centrifuge briefly.
|  | 8 | Add 16 μl of the Master Mix to each sample well containing the index primers and DNA samples. Make sure to change tips after pipetting into each well.
|  | 9 | Apply Microseal B and immediately vortex the plate for 5 seconds to mix.
|  | 10 | Immediately centrifuge plate at 1000 x g for 30 seconds.
|  | 11 | Immediately place the plate on the thermal cycler in a post-amplification area and run the PCR cycling protocol. The PCR Cycling Protocol takes
approximately 1 hour.
Thermocycler Used:
Program Name:
Indexing PCR program:
Reaction Volume: 40 µL
Lid temperature: 100°C
|  |  |  | # | Temperature | Time | Ramp Rate | Number of Cycles
|  |  |  | 1 | 98°C | 3 minutes | default | 1
|  |  |  | 2 | 96°C | 15 seconds | default | 8
|  |  |  | 3 | 70°C | 5 seconds | default
|  |  |  | 4 | 57°C | 60 seconds | refer to ramp table
|  |  |  | 5 | 72°C | 30 seconds | refer to ramp table
|  |  |  | 6 | 96°C | 15 seconds | default | 12
|  |  |  | 7 | 72°C | 60 seconds | default
|  |  |  | 8 | 72°C | 2 minutes | default | 1
|  |  |  | 9 | 10°C | Hold | default | 1
[6 empty] | Supported Thermal Cyclers | Ramp Rate Equivalent
[6 empty] | ABI GeneAmp PCR System 9700 | 25%
[6 empty] | ABI Veriti Thermal Cycler | 18%
[6 empty] | ABI SimpliAmp | 0.7°C/sec
[6 empty] | Eppendorf Nexus | 0.4°C/sec
NOTE: Safe stopping point. You can leave the plate on the thermal cycler up to 48hrs at 10°C hold step, if needed.
NOTE: If proceeding immediately to cleanup, ensure Purification Beads are brought to room temperature before use.
Finish time:
PROTOCOL END. PROCEED TO CLEANUP STEP.

## Cleanup and Qubit
ALLOSEQ HCT - CLEANUP AND CONCENTRATION DETERMINATION
| Experiment ID #: |  |  |  | 0
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Library Pool (µL) |  | Preparation Required
|  |  |  | Purification Beads |  |  | 2°C to 8°C BOX 3 | 150 |  | Bring to room temperature, at least 30 mins.
|  |  |  | Resuspension Buffer |  |  | 2°C to 8°C BOX 3 | 100 |  | Bring to room temperature.
|  |  |  | PCR-grade Water |  |  | 15°C to 30°C User supplied | 600 |  | No preparation required.
|  |  |  | 100% Ethanol |  |  | 15°C to 30°C User supplied | 2400 |  | Prepare 80% ethanol, as described below.
|  |  |  | Qubit dsDNA HS assay kit |  |  | 15°C to 30°C User supplied |  |  | Bring Standards to room temperature.
|  | 2 | Vortex and spin all reagents before use.
|  | 3 | Gather required consumables listed below:
DynaMag-2 magnet,
1.5 mL microcentrifuge tubes,
5 mL tube,
Qubit tubes,
Qubit Fluorometer.
|  | 3 | Prepare fresh 80% ethanol (2 washes):
|  |  |  | Reagent |  | Volume per Sample (µL)
|  |  |  | 100% Ethanol |  | 2400
|  |  |  | PCR-grade Water |  | 600
|  |  |  | Total |  | 3000
CLEANUP PROTOCOL
|  | 1 | Centrifuge the indexing PCR plate at 1000 x g for 30 seconds.
|  | 2 | Pool the appropriate volume of each sample defined in the table below into a single 1.5 mL Eppendorf tube, to achieve a final volume of 120 μl.
|  |  |  | # of samples |  | Volume to pool per Sample (µL)
|  |  |  | 0 |  | #DIV/0!
|  | 3 | Vortex Purification Beads until complete resuspension (visually check that there is no pellet left). Briefly spin down to get the bead suspension at
the bottom of the tube
|  | 4 | Add 100 μl of the Purification Beads to the 120 μl of pooled library.
|  | 5 | Vortex the tube for 5 seconds on full speed, then incubate for 5 minutes at room temperature.
|  | 6 | Briefly spin down, then place the tube on DynaMag-2 magnet and allow beads to pellet for 5 minutes
|  | 7 | Keep the tube on the magnet and remove all supernatant with a P200 pipette. Use a P20 pipette to remove any remaining supernatant, if necessary.
|  | 8 | Add 65 µL of the AlloSeq HCT Resuspension Buffer.
|  | 9 | Vortex the tube for 5 seconds on full speed and incubate tube for 5 minutes at room temperature
|  | 10 | Briefly spin down, then place the tube on the magnet and allow beads to pellet for 5 minutes
|  | 11 | Keep the tube on the magnet and, without touching or disturbing the beads, transfer 60 µL of supernatant to a new 1.5 mL Eppendorf tube.
|  | 12 | Ensure AlloSeq HCT Purification Beads are thoroughly resuspended (see step 3). Add 50 µL of the beads to the 60 µL library supernatant.
|  | 13 | Vortex the tube for 5 seconds on full speed and incubate for 5 minutes at room temperature.
|  | 14 | Briefly spin down, then place the tube on the magnet and allow beads to pellet for 5 minutes.
|  | 15 | Keep the tube on the magnet and remove all supernatant with a P200 pipette.
|  | 16 | Wash with 80% Ethanol two times:
|  |  | a) | Add 1 mL of 80% ethanol without disturbing the beads,
|  |  | b) | Incubate at room temperature for 30 seconds,
|  |  | c) | Using a pipette, aspirate and discard all superanatant,
|  |  | d) | Repeat steps a) through c) for a total of 2 washes."
|  | 17 | Remove all remaining supernatant with P20 pipette.
|  | 18 | Keep the tube open and air dry the the tube for 5- 10 minutes (or until no EtOH remains) at room temperature, to allow residual ethanol to evaporate.
|  | 19 | Add 35 µL of Resuspension Buffer to each tube to elute target fragments.
|  | 20 | Vortex the tube for 5 seconds on full speed, then incubate tube for 5 minutes at room temperature.
|  | 21 | Briefly spin down, then place the tube on the magnet and allow beads to pellet for 5 minutes
|  | 22 | Keep tube on the magnet and without touching or disturbing the beads transfer 32 μl of supernatant to a new 1.5 mL Eppendorf tube
NOTE: Safe stopping point. Cleaned up library may be stored at -15°C to -25°C for up to 1 week.
Finish time:
QUANTIFICATION OF CLEANED-UP LIBRARY
| 1 | Set up two assay tubes for the standards and two for the library and one duplicate.
| 2 | Prepare the QuBit working solution using the following volumes:
|  |  |  | Reagent |  | Voulme (µL)
|  |  |  | QuBit dsDNA HS buffer |  | 995
|  |  |  | QuBit dsDNA HS dye |  | 5
|  |  |  | Total |  | 1000
NOTE: This volume of QuBit working solution includes excess.
| 3 | Vortex working solution for 2-3 seconds, and then pulse-spin.
| 4 | Aliquot 190 µL of working solution into each of the Standard tubes.
| 5 | Aliquot 198 µL of working solution into each of the Library tubes.
| 6 | Aliquot 10 µL of Standard solution into each of the respective Standard tubes.
| 7 | Aliquot 2 µL of the library into the respective tubes.
| 8 | Vortex all tubes for 2-3 seconds, and then pulse-spin.
| 9 | Incubate the tubes for 2 minutes at room temperature.
| 10 | Insert the tubes in the QuBit Fluorometer and take readings (see QuBit protocol for further information).
| 11 | Record QuBit readings in the table below.
|  |  |  | 1st read (ng/µL) |  | 2nd read (ng/µL) | Average (ng/µL) |  | Standard 1 read | Standard 2 read
0
PROTOCOL END. PROCEED TO SEQUENCING.

## TapeStation(Optional)
ALLOSEQ HCT - TAPESTATION VISUALISATION
| Experiment ID #: |  |  |  | 0
Number of pools:
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 0 | Preparation Required
|  |  |  | High Sensitivity D1000 Ladder |  |  | 2°C to 8°C User supplied | 2 | 2 |  | Bring to room temperature.
|  |  |  | High Sensitivity D1000 Sample Buffer |  |  | 2°C to 8°C User supplied | 2 | 2 |  | Bring to room temperature.
|  | 2 | Gather required consumables listed below:
High Sensitivity D1000 ScreenTape,
TapeStation Optical tube strips (8x Strip),
TapeStation Optical tube strip caps (8x Strip)
SIZE DISTRIBUTION VISUALISATION OF CLEANED LIBRARY
|  | 1 | Add 2 µL of High Sensitivity D1000 Ladder into a reference well.
|  | 2 | Add 2 µL of cleaned library to a well on the strip.
|  | 3 | Add 2 µL of High Sensitivity D1000 sample buffer to the cleaned library and ladder well.
|  | 4 | Seal all tubes with caps.
|  | 5 | Vortex all tubes thoroughly using IKA vortex at 2000 rpm for 1 minute.
|  | 6 | Centrifuge briefly to ensure that all samples are at the bottom of the tubes.
|  | 7 | Uncap and load sample tubes into the TapeStation instrument.
|  | 8 | Select the required tubes on the TapeStation Controller Software and run the samples.
|  | 9 | Once the run is complete, start TapeStation Analysis Software to view results.
|  | 10 | Record results in a table below:
|  |  | Pool # |  | Tube # | Cleaned Library ID | Peak Fragment Size (bp)
1
Finish time:
PROTOCOL END. PROCEED TO SEQUENCING

## Sequencing
ALLOSEQ HCT - DILUTE AND DENATURE FOR MISEQ
Flow Cell ID #:
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) |  | Preparation Required
|  |  |  | AlloSeq HCT Cleaned-up Library Pool |  | -15°C to -25°C User prepared | Do not dilute |  | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 1 | 4 |  | Bring to room temperature.
|  |  |  | PCR-grade Water |  | 15°C to 30°C User supplied | 36 |  | No preparation required.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | Min 588 |  | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 3 | Do not dilute |  | Bring to room temperature.
|  |  |  | 10 nM PhiX (optional) |  | -15°C to -25°C User supplied | 2 |  | Thaw and then keep on ice.
|  | 2 | Briefky vortex and spin all reagents before use.
|  | 3 | Thaw and prepare Illumina sequencing reagents according to manufacturer’s specifications.
|  | 4 | Gather required consumables listed below:
1.5 microcentrifuge tubes
|  | 5 | Prepare 0.2N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | PCR-grade Water | 36
|  |  |  | 2N NaOH | 4
DILUTE ALLOSEQ HCT LIBRARY TO 1.33 nM
|  | 1 | Dilute the Library Pool to: |  |  | 1.33 | nM
|  |  |  | Information Required |  | Value
|  |  |  | Library Pool Concentration (ng/µL) |  | 0
|  |  |  | Pool Concentration (nM) |  | 0
|  |  |  | Reagent |  | Volume (µL) |  | Calculation |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | Do not dilute |  | Concentration (ng/ul) |  | 0
|  |  |  | Library Pool |  | Do not dilute |  | Concentration (nM) |  | 0
|  |  |  | Total |  | 0
DENATURE & DILUTE FINAL LIBRARY (WITHOUT PHIX)
|  | 1 | In a new 1.5 mL tube combine 1.33nM library pool and 0.2 N NaOH.
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Library Pool (ul) |  | #DIV/0! | Note: Ensure sufficient volume is available. If not, reclean of library required.
|  |  |  | 0.2 N NaOH(ul) |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex, spin down briefly, and incubate at room temperature for 5 minutes.
|  | 3 | Dilute Final library to: |  |  | 20 | pM
|  |  |  | Reagent |  | Volume (µL) |  | Calculation |  | Volume (µL)
|  |  |  | Denatured Library (ul) |  | #DIV/0! |  | Concentration (ng/ul) |  | 0.0024013
|  |  |  | HT1 |  | #DIV/0! |  | Input (ng) |  | 1.4408
|  |  |  | Total |  | 600
|  | 4 | Vortex briefly and spin down.
|  | 5 | Load 600 μl of the final library to the Illumina MiSeq sequencing cartridge previously thawed and prepared according to
manufacturer’s recommendation.
|  | 6 | Proceed according to the Illumina MiSeq user’s manual to load the flow cell and reagent cartridge into the instrument to start the sequencing run.
NOTE: Make sure to use a sample sheet prepared with this workbook according to the instructions in the AlloSeq HCT Software IFU.
DENATURE & DILUTE FINAL LIBRARY (WITH PHIX)
|  | 1 | In a new 1.5 mL tube combine 1.33nM library pool and 0.2 N NaOH.
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Library Pool (ul) |  | #DIV/0! | Note: Ensure sufficient volume is available. If not, reclean of library required.
|  |  |  | 0.2 N NaOH(ul) |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | In a new 1.5 mL tube combine 2 μl of 10 nM PhiX and 2 μl of 0.2 N NaOH.
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | PhiX-10nM (ul) |  | 2
|  |  |  | 0.2 N NaOH(ul) |  | 2
|  |  |  | Total |  | 4
|  | 3 | Vortex both tubes, spin down briefly, and incubate at room temperature for 5 minutes.
|  | 4 | Dilute PhiX library to 20pM
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Denatured 5nM PhiX library (ul) |  | 4
|  |  |  | HT1 |  | 996
|  |  |  | Total |  | 1000
|  | 5 | Dilute Final library to: |  |  | 20 | pM
|  |  |  | Reagent |  | Volume (µL) |  | Calculation |  | Volume (µL)
|  |  |  | Denatured Library (ul) |  | #DIV/0! |  | Concentration (ng/ul) |  | 0.0024013
|  |  |  | Diluted 20pM PhiX library |  | 6 |  | Input (ng) |  | 1.4408
|  |  |  | HT1 |  | #DIV/0!
|  |  |  | Total |  | 600
NOTE: Final PhiX concentration in the library is 1%.
|  | 4 | Vortex briefly and spin down. Store on Ice
|  | 5 | Load 600 μl of the final library to the Illumina MiSeq sequencing cartridge previously thawed and prepared according to
manufacturer’s recommendation.
|  | 6 | Proceed according to the Illumina MiSeq user’s manual to load the flow cell and reagent cartridge into the instrument to start the sequencing run.
NOTE: Make sure to use a sample sheet prepared with this workbook according to the instructions in the AlloSeq HCT Software IFU.
PROTOCOL END.
Finish time:

## VersionHistory
VERSION | REVISION DATE (DDMMMYY ) | REVISION DESCRIPTION
1 | 10 May 22 | First version of the AlloSeq HCT 96 Workbook – CE IVD. Refer to CR-2021-035.
2 | 18 Nov 22 | Cover Page Tab: Updated version # and Issue Date, Revision date and Revision Description. Sample Sheet Tab: Updated the instruction box connected to cell A1. Removed v2 from step 4. Plate Map Tab: Updated all relevant formulas in the Plate layout table to remove additional zeros eg. =IFNA(INDEX('Sample Sheet'!$A$1:$J$116,MATCH("A01",'Sample Sheet'!$D$1:$D$116,0),1),"") changed to =IFNA(INDEX('Sample Sheet'!$A$1:$J$116,MATCH("A1",'Sample Sheet'!$D$1:$D$116,0),1),"") Calculations Tab: In step 3, for stock concentrations >10ng/µL, changed the volume prepared from “400µL” to “100µL” Lib Prep Tab: Unlocked cells F6-F8 and G14-G18. Cleanup and Qubit Tab: Unlocked cells F4-F6 and G12-G16. Added instruction to steps 6, 10, 14 and 21 to briefly spin down the tube content prior to placing on the magnet. Tapestation Tab: Unlocked cells F5-F7 and G13 & G14 Sequencing Tab: Unlocked cells F3-F6 and G12-G17 Workbook Verification Tab: Corrected Errors, included missing formula, included section for verification of the Index Layouts Tab and updated to reflect changes listed above where relevant. Updated Plate Map tab to include a count of the number of samples based on the sample sheet tab. Added the generic colour coding key for Fremantle workbooks. Updated formatting of cells F3 and F5 of the Lib Prep tab. Workbook verification tab updated to reflect this change.
3 | 12 Oct 23 | Plate Map Tab: Updated all relevant formulas in the Plate layout table to remove additional zeros eg. =IFNA(INDEX('Sample Sheet'!$A$1:$J$116,MATCH("A01",'Sample Sheet'!$D$1:$D$116,0),1),"") changed to =IFNA(INDEX('Sample Sheet'!$A$1:$J$116,MATCH("A1",'Sample Sheet'!$D$1:$D$116,0),1),"")