## CE Cover
AlloSeq Tx
Assay Workbook
IFU095-5-CS
Číslo verze: 1.0
Revidováno: září 2023
ASTX17.1(24)-IVD
ASTX17.1(24)-B-IVD
ASTX17.1(96)-A-IVD
ASTX17.1(96)-B-IVD
ASTX9.1(96)-A-IVD
ASTX9.1(96)-B-IVD
| CareDx Pty Ltd |  | CareDx AB, | Qarad BV,
| 20 Collie St. |  | Franzéngatan 5 | Cipalstraat 3
| Fremantle WA 6160 |  | SE-112 51 | 2440 Geel
| Australia |  | Stockholm, Sweden | Belgium
© 2021-2024 CareDx, Inc. Veškeré služební značky nebo ochranné známky jsou ve vlastnictví nebo jsou předmětem licence společnosti CareDx, Inc. nebo jejími pobočkami.

## DropDowns (hidden)
Probe_Set | Sequencer | i7_IndexNames | i7_IndexSequences | i5_IndexNames | i5_IndexSequences | Wells 96 | Cell Reference | Wells 24 | Cell Reference | Kit
Select | Select |  |  |  |  | A1 | G24 | A1 | G24 | 24
AlloSeqTx17.1 | MiSeq | H705 | GGACTCCT | H503 | TATCCTCT | B1 | G25 | A2 | H24 | 96
AlloSeqTx9.1 | MiniSeq | H706 | TAGGCATG | H505 | GTAAGGAG | C1 | G26 | A3 | I24
| iSeq | H707 | CTCTCTAC | H506 | ACTGCATA | D1 | G27 | A4 | J24
|  | H710 | CGAGGCTG | H517 | GCGTAAGA | E1 | G28 | A5 | K24
|  | H711 | AAGAGGCA | H502 | CTCTCTAT | F1 | G29 | A6 | L24
|  | H714 | GCTCATGA | H507 | AAGGAGTA | G1 | G30 | B1 | G25
|  | H701 | TAAGGCGA | H508 | CTAAGCCT | H1 | G31 | B2 | H25
AlloSeqTx17.2 |  | H702 | CGTACTAG | H521 | GAGCCTTA | A2 | H24 | B3 | I25
AlloSeqTx8.1 |  | H703 | AGGCAGAA | H510 | CGTCTAAT | B2 | H25 | B4 | J25
AlloSeqTxKIR |  | H704 | TCCTGAGC | H511 | TCTCTCCG | C2 | H26 | B5 | K25
|  | H712 | GTAGAGGA | H513 | TCGACTAG | D2 | H27 | B6 | L25
|  | H715 | ATCTCAGG | H522 | TTATGCGA | E2 | H28 | C1 | G26
|  | H716 | ACTCGCTA | H515 | TTCTAGCT | F2 | H29 | C2 | H26
|  | H718 | GGAGCTAC | H516 | CCTAGAGT | G2 | H30 | C3 | I26
|  | H719 | GCGTAGTA | H518 | CTATTAAG | H2 | H31 | C4 | J26
|  | H720 | CGGAGCCT | H520 | AAGGCTAT | A3 | I24 | C5 | K26
|  | H721 | TACGCTGC |  |  | B3 | I25 | C6 | L26
|  | H722 | ATGCGCAG |  |  | C3 | I26 | D1 | G27
|  | H723 | TAGCGCTC |  |  | D3 | I27 | D2 | H27
|  | H724 | ACTGAGCG |  |  | E3 | I28 | D3 | I27
|  | H726 | CCTAAGAC |  |  | F3 | I29 | D4 | J27
|  | H727 | CGATCAGT |  |  | G3 | I30 | D5 | K27
|  | H728 | TGCAGCTA |  |  | H3 | I31 | D6 | L27
|  | H729 | TCGACGTC |  |  | A4 | J24
[6 empty] | B4 | J25
[6 empty] | C4 | J26
[6 empty] | D4 | J27
[6 empty] | E4 | J28
[6 empty] | F4 | J29
[6 empty] | G4 | J30
[6 empty] | H4 | J31
[6 empty] | A5 | K24
[6 empty] | B5 | K25
[6 empty] | C5 | K26
[6 empty] | D5 | K27
[6 empty] | E5 | K28
[6 empty] | F5 | K29
[6 empty] | G5 | K30
[6 empty] | H5 | K31
[6 empty] | A6 | L24
[6 empty] | B6 | L25
[6 empty] | C6 | L26
[6 empty] | D6 | L27
[6 empty] | E6 | L28
[6 empty] | F6 | L29
[6 empty] | G6 | L30
[6 empty] | H6 | L31
[6 empty] | A7 | M24
[6 empty] | B7 | M25
[6 empty] | C7 | M26
[6 empty] | D7 | M27
[6 empty] | E7 | M28
[6 empty] | F7 | M29
[6 empty] | G7 | M30
[6 empty] | H7 | M31
[6 empty] | A8 | N24
[6 empty] | B8 | N25
[6 empty] | C8 | N26
[6 empty] | D8 | N27
[6 empty] | E8 | N28
[6 empty] | F8 | N29
[6 empty] | G8 | N30
[6 empty] | H8 | N31
[6 empty] | A9 | O24
[6 empty] | B9 | O25
[6 empty] | C9 | O26
[6 empty] | D9 | O27
[6 empty] | E9 | O28
[6 empty] | F9 | O29
[6 empty] | G9 | O30
[6 empty] | H9 | O31
[6 empty] | A10 | P24
[6 empty] | B10 | P25
[6 empty] | C10 | P26
[6 empty] | D10 | P27
[6 empty] | E10 | P28
[6 empty] | F10 | P29
[6 empty] | G10 | P30
[6 empty] | H10 | P31
[6 empty] | A11 | Q24
[6 empty] | B11 | Q25
[6 empty] | C11 | Q26
[6 empty] | D11 | Q27
[6 empty] | E11 | Q28
[6 empty] | F11 | Q29
[6 empty] | G11 | Q30
[6 empty] | H11 | Q31
[6 empty] | A12 | R24
[6 empty] | B12 | R25
[6 empty] | C12 | R26
[6 empty] | D12 | R27
[6 empty] | E12 | R28
[6 empty] | F12 | R29
[6 empty] | G12 | R30
[6 empty] | H12 | R31

## IndexLayout (hidden)
Product Code | A | B | C | D | E | F | G | H | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 2.0 Library_Prep Lookup
Select | [20 empty] | Select index set on 1.0 Sample_Prep tab.
24 Set A Indexes | H503 | H505 | H506 | H517 |  |  |  |  | H705 | H706 | H707 | H710 | H711 | H714 | [6 empty] | Indexes: H503, H505, H506, H517, H705, H706, H707, H710, H711, H714
24 Set B Indexes | H502 | H507 | H508 | H521 |  |  |  |  | H701 | H702 | H703 | H704 | H712 | H715 | [6 empty] | Indexes: H502, H507, H508, H521, H701, H702, H703, H704, H712, H715
96 Set A Indexes | H503 | H505 | H506 | H517 | H502 | H507 | H508 | H521 | H705 | H706 | H707 | H710 | H711 | H714 | H701 | H702 | H703 | H704 | H712 | H715 | Indexes: H503, H505, H506, H517, H502, H507, H508, H521, H705, H706, H707, H710, H711, H714, H701, H702, H703, H704, H712, H715
96 Set B Indexes | H510 | H511 | H513 | H522 | H515 | H516 | H518 | H520 | H716 | H718 | H719 | H720 | H721 | H722 | H723 | H724 | H726 | H727 | H728 | H729 | Indexes: H510, H511, H513, H522, H515, H516, H518, H520, H716, H718, H719, H720, H721, H722, H723, H724, H726, H727, H728, H729
| H503 | H505 | H506 | H517 | H502 | H507 | H508 | H521 | H716 | H718 | H719 | H720 | H721 | H722 | H723 | H724 | H726 | H727 | H728 | H729 | Indexes: H503, H505, H506, H517, H502, H507, H508, H521, H716, H718, H719, H720, H721, H722, H723, H724, H726, H727, H728, H729
| H510 | H511 | H513 | H522 | H515 | H516 | H518 | H520 | H705 | H706 | H707 | H710 | H711 | H714 | H701 | H702 | H703 | H704 | H712 | H715 | Indexes: H510, H511, H513, H522, H515, H516, H518, H520, H705, H706, H707, H710, H711, H714, H701, H702, H703, H704, H712, H715
96 Set C Indexes
96 Set D Indexes

## 0.1 Import
IMPORT SEZNAMU VZORKŮ | [5 empty] | Seznam vzorků:
1
|  | Níže uvedené makro lze použít k importu seznamů vzorků z excelového souboru. |  |  | 2
|  | POZNÁMKA: Seznam vzorků musí být umístěn ve sloupci A s počáteční buňkou A1. Nezačleňujte záhlaví sloupce. |  |  | 3
| 1 | V případě potřeby klikněte na tlačítko „Enable Makros“ (povolit makra) nacházející se ve žluté liště výše. |  |  | 4
| 2 | Klikněte na tlačítko „Import Sample List“ (importovat seznam vzorků) vpravo. |  |  | 5
| 3 | Vyhledejte seznam a klikněte na tlačítko Otevřít. |  |  | 6
| 4 | Zkontrolujte názvy vzorků v seznamu vpravo. |  |  | 7
| 5 | Vyberte Typ kitu níže. |  | 1 | 8
| 6 | Vyberte počáteční jamku. POZNÁMKA: 24vzorkové kity zaujímají pozice zleva doprava: A1, A2, A3. |  |  | 9
|  | 96vzorkové kity zaujímají pozice shora dolů: A1, B1, C1. |  |  | 10
| 7 | Klikněte na tlačítko „Build Plate Map“ (vytvořit rozvržení destičky) pro vygenerování rozvržení destičky |  |  | 11
| 8 | Klikněte na kartu „1.0 Sample_Prep" a zkontrolujte Plate Layout (rozvržení destičky). |  |  | 12
13
|  | Vyberte typ kitu: | Vyberte počáteční jamku: |  | 14
15
16
17
18
KONEC IMPORTU SEZNAMU VZORKŮ. POKRAČUJTE PODLE PROTOKOLU. |  |  |  |  | 19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93
94
95
96

## 1.0 Sample_Prep
ALLOSEQ TX - PŘÍPRAVA VZORKŮ
| ID experimentu: |  |  | Zadejte |  |  |  | POZNÁMKA: ZELENÁ pole musí vyplnit operátor během nastavení.
| Popis: |  |  | Zadejte |  |  |  | POZNÁMKA: ORANŽOVÁ pole je nutno před tiskem vybrat z rozevírací nabídky.
| Sada prób/specifikace kitu: |  |  | Select |  |  |  | POZNÁMKA: ŽLUTÁ pole je nutno před tiskem vyplnit.
| Sekvenátor: |  |  | Select
| Počet vzorků: |  |  | 0
| Operátor: |  |  | Zadejte
| Datum: |  |  | Zadejte
Nastavení experimentu
|  | 1 | Do výše uvedených žlutých polí zadejte ID experimentu, popis, jméno operátora a datum.
|  | 2 | Z rozevírací nabídky uvedené výše vyberte sadu prób, která má být použita v experimentu.
|  | 3 | Z výše uvedeného rozevíracího seznamu vyberte typ sekvenátoru, který chcete použít. Po výběru typu sekvenátoru se v sešitu zobrazí pokyny pro sekvenování.
|  | 4 | Zadejte ID vzorků, které mají být testovány, do žluté části Plate Layout uvedené níže v souladu s požadovanou konfigurací. POZNÁMKA: Je možno zadat pouze alfanumerické znaky. POZNÁMKA: Duplicitní ID vzorků budou označeny červeně proto, aby byl uživatel upozorněn na nutnost provedení opravy. Požadavkem softwaru sekvenátoru je, aby měl při spuštění pouze jedinečná ID vzorků. POZNÁMKA: Nezadávejte informace do pozic, které nebudou obsahovat vzorky.
|  | 5 | Vyberte požadovanou indexovou sadu z oranžové rozevírací nabídky pod položkou Plate Layout.
|  | 6 | Pokud je to třeba, lze upravit pořadí indexu i7 tak, že vyberete alternativní index i7z možností rozevírací nabídky pro požadovaný sloupec. POZNÁMKA: Pokud jsou vybrány duplicitní indexy i7, budou buňky označeny červeně tak, aby byl uživatel upozorněn na nutnost provedení opravy.
|  | 7 | Pokud je to třeba, lze upravit pořadí indexu i5 tak, že vyberete alternativní index i5 z možností rozevírací nabídky pro požadovaný řádek. POZNÁMKA: Pokud jsou vybrány duplicitní indexy i5, budou buňky označeny červeně tak, aby byl uživatel upozorněn na nutnost provedení opravy.
|  | 8 | Po dokončení všech výše uvedených činností klikněte na kartu 1.2 SampleSheet a proveďte kontrolu. Červený text označuje místa, kde je nutné informace doplnit. Pokud není patrný žádný červený text, pak lze kartu SampleSheet uložit jako CSV (Comma delimited) (*.csv). Spusťte makro exportu kliknutím na tlačítko „Export“ a vygenerujte soubor csv. Případně postupujte podle následujících kroků a uložte soubor ručně: Uložte soubor jako „SampleSheet.csv“. Uložením ve formátu csv se uloží pouze aktivní karta. Otevřete uložený soubor SampleSheet.csv v Excelu a odstraňte všechny prázdné řádky (sekce [Data])v oblasti řádků 22 až 117, které neobsahují informace o vzorku. Jakmile to uděláte, soubor uložte. Seznam vzorků je poté připraven k importu pro sekvenování na sekvenátoru Illumina.
|  |  |  | PLATE LAYOUT |  | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12
|  |  |  | Select |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | A | 0
|  |  |  | B | 0
|  |  |  | C | 0
|  |  |  | D | 0
|  |  |  | E | 0
|  |  |  | F | 0
|  |  |  | G | 0
|  |  |  | H | 0
POKRAČUJTE V PŘÍPRAVĚ KNIHOVNY.
ImportData!A1
ImportData!A2
ImportData!A3
ImportData!A4
ImportData!A5
ImportData!A6
ImportData!A7
ImportData!A8
ImportData!A9
ImportData!A10
ImportData!A11
ImportData!A12
ImportData!A13
ImportData!A14
ImportData!A15
ImportData!A16
ImportData!A17
ImportData!A18
ImportData!A19
ImportData!A20
ImportData!A21
ImportData!A22
ImportData!A23
ImportData!A24

## 1.1 LinearView
ROZVRŽENÍ VZORKŮ – LINEÁRNÍ POHLED
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
| Sample # | Well # | Sample_ID | Sample_Name | Sample_Plate | Sample_Well | I7_Index_ID | index | I5_Index_ID | index2 | Sample_Project
| 1 | A01
| 2 | B01
| 3 | C01
| 4 | D01
| 5 | E01
| 6 | F01
| 7 | G01
| 8 | H01
| 9 | A02
| 10 | B02
| 11 | C02
| 12 | D02
| 13 | E02
| 14 | F02
| 15 | G02
| 16 | H02
| 17 | A03
| 18 | B03
| 19 | C03
| 20 | D03
| 21 | E03
| 22 | F03
| 23 | G03
| 24 | H03
| 25 | A04
| 26 | B04
| 27 | C04
| 28 | D04
| 29 | E04
| 30 | F04
| 31 | G04
| 32 | H04
| 33 | A05
| 34 | B05
| 35 | C05
| 36 | D05
| 37 | E05
| 38 | F05
| 39 | G05
| 40 | H05
| 41 | A06
| 42 | B06
| 43 | C06
| 44 | D06
| 45 | E06
| 46 | F06
| 47 | G06
| 48 | H06
| 49 | A07
| 50 | B07
| 51 | C07
| 52 | D07
| 53 | E07
| 54 | F07
| 55 | G07
| 56 | H07
| 57 | A08
| 58 | B08
| 59 | C08
| 60 | D08
| 61 | E08
| 62 | F08
| 63 | G08
| 64 | H08
| 65 | A09
| 66 | B09
| 67 | C09
| 68 | D09
| 69 | E09
| 70 | F09
| 71 | G09
| 72 | H09
| 73 | A10
| 74 | B10
| 75 | C10
| 76 | D10
| 77 | E10
| 78 | F10
| 79 | G10
| 80 | H10
| 81 | A11
| 82 | B11
| 83 | C11
| 84 | D11
| 85 | E11
| 86 | F11
| 87 | G11
| 88 | H11
| 89 | A12
| 90 | B12
| 91 | C12
| 92 | D12
| 93 | E12
| 94 | F12
| 95 | G12
| 96 | H12
UKONČIT ROZVRŽENÍ VZORKŮ

## 1.2 SampleSheet
[Header]
IEMFileVersion | 5
Investigator Name | Zadejte
Experiment Name | Zadejte
Date | Zadejte
Workflow | GenerateFASTQ
Application | FASTQ Only
Instrument Type | Select
Assay | Select
Description | Zadejte
Chemistry | Amplicon
[Settings]
[Reads]
151
151
[Data]
Sample_ID | Sample_Name | Sample_Plate | Sample_Well | I7_Index_ID | index | I5_Index_ID | index2 | Sample_Project | Autolauncher_Project
Kliknutím na tlačítko níže vygenerujete seznam vzorků ve formátu csv. Po kliknutí na tlačítko se zobrazí výzva k zadání názvu a umístění souboru.
Export

## 2.0 Library_Prep
ALLOSEQ TX – PŘÍPRAVA KNIHOVNY
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Vezměte požadované reagencie (objem je uveden v tabulce níže) a připravte je podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 vzorek (μl) | Objem pro # požadovaných vzorků (μl): | 0 | Požadovaná příprava
|  |  |  | Tagmentation Buffer |  |  | –15 °C až –25 °C KRABIČKA 1 BOX 1 | 10 | 0 |  | Dejte do pokojové teploty
|  |  |  | Sterile Water |  |  | 15 °C až 30 °C Dodáno uživatelem User supplied | 30 | 0 |  | Žádná příprava se nepožaduje.
|  |  |  | Tagmentation Beads |  |  | –15 °C až –25 °C KRABIČKA 1 BOX 1 | 10 | 0 |  | Nechte při pokojové teplotě po dobu alespoň 30 minut.
|  |  |  | Stop Buffer |  |  | 15 °C až 30 °C KRABIČKA 2 BOX 2 | 10 | 0 |  | Žádná příprava se nepožaduje.
|  |  |  | Tagmentation Wash Buffer |  |  | 15 °C až 30 °C KRABIČKA 2 BOX 2 | 300 | 0 |  | Žádná příprava se nepožaduje.
|  |  |  | PCR Mix (nebo PCR Mix-1 u kitů na 96 reakcí) |  |  | –15 °C až –25 °C KRABIČKA 4 (nebo KRABIČKA 1) BOX 4 (or BOX 1) | 20 | 0 |  | Rozmrazte a poté mějte na ledu. Vraťte do mrazničky pro další použití.
|  |  |  | Select index set on 1.0 Sample_Prep tab. |  |  | –15 °C až –25 °C KRABIČKA 3 BOX 3 | 10 | 0 |  | Rozmrazte a poté uchovávejte na ledu.
|  | 2 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 3 | Připravte si požadovaný spotřební materiál uvedený níže:
96-jamková destička,
lepící folie Microseal B,
1,5ml mikrocentrifugační zkumavky
PROTOKOL PRO PŘÍPRAVU KNIHOVNY
POZNÁMKA: Neexistuje žádný bod bezpečného zastavení až do indexing PCR. Proces trvá přibližně 1:50 hodiny (včetně 50 min v PCR termocykleru).
|  | 1 | Do příslušné jamky PCR destičky podle rozpisu na listu 1.0 Sample_Prep dáme 10 μl vzorku o koncentraci 10–100 ng/μl.
|  | 2 | Připravte Tagmentation master mix (tagmentační master mix) z následujících reagencií:
|  |  |  | Reagencie |  | Objem pro 1 vzorek (μl) | Objem pro požadovaný počet vzorků + 10% rezerva (μl): | 0
|  |  |  | Tagmentation Buffer |  | 10 | 0
|  |  |  | Sterilní voda |  | 20 | 0
|  |  |  | Tagmentation Beads |  | 10 | 0
|  |  |  | Celkem |  | 40 | 0
|  | 3 | Smíchejte výše uvedený master mix vortexováním a poté krátce odstřeďte.
|  | 4 | Napipetujte pipetou 40 μl tagmentačního master mixu do každé jamky, která obsahuje vzorek DNA.
|  | 5 | Přelepte destičku fólií Microseal B.
|  | 6 | Centrifugujte destičku při 100 x g po dobu 10 sekund, abyste shromáždili veškerý obsah na dně jamky.
|  | 7 | Pomocí třepačky při 1800 otáčkách za minutu třepejte po dobu 1 minuty.
|  | 8 | Vizuálně zkontrolujte destičku:
|  |  | a) | Pokud kuličky nejsou v jamce rovnoměrně rozloženy, opakujte protřepávání podle kroku 7,
|  |  | b) | Pokud materiál není na dně jamky nebo se rozstříkl na fólii Microseal B, krátce stočte a opakujte krok protřepání 7.
|  | 9 | Umístěte destičku do termocykleru a spusťte program Tagmentation.
Použitý termocykler:
Název programu:
Program Tagmentation:
Reakční objem: 50 µL
Teplota víka: 105 °C
|  |  |  | Teplota |  | Čas
|  |  |  | 55°C |  | 5 minut
|  |  |  | 10°C |  | 2 minuty
|  | 10 | Po dokončení programu okamžitě vyjměte destičku z termocykleru a nechte ji 2 minuty při pokojové teplotě.
|  | 11 | Odstraňte fólii Microseal B.
|  | 12 | Do každé jamky přidejte 10 μl pufru Stop Buffer.
|  | 13 | Přelepte destičku novou fólií Microseal B.
|  | 14 | Pomocí třepačky při 1800 otáčkách za minutu třepejte po dobu 1 minuty.
|  | 15 | Destičku inkubujte dalších 5 minut při pokojové teplotě.
|  | 16 | Během inkubace vyjměte PCR Mix a indexy z mrazničky, aby se rozmrazily, a poté je umístěte na led / do chladícího stojanu.
|  | 17 | Vizuálně destičku zkontrolujte. Pokud materiál není na dně jamky nebo je rozstříknutý na fólii Microseal B, destičku krátce stočte.
|  | 18 | Promyjte třikrát pomocí Tagmentation Wash Buffer (promývací pufr), jak je popsáno níže:
|  |  |  | a) | Odstraňte fólii Microseal B a umístěte destičku na magnetický stojanna 30 sekund, aby se kuličky shromáždily v jamkách souběžně s magnetem,
|  |  |  | b) | Pomocí pipety odsajte a zlikvidujte supernatant, přičemž kuličky ponechte v jamkách souběžně s magnetem,
|  |  |  | c) | Pomalu přidejte 100 μl pufru Tagmentation Wash Buffer do každé jamky,
|  |  |  | d) | Přelepte destičku novou fólií Microseal B a ujistěte se, že je fólie řádně připevněna,
|  |  |  | e) | Pomocí třepačky při 1800 otáčkách za minutu třepejte 2 minuty při pokojové teplotě.
|  |  |  | f) | UPOZORNĚNÍ: Pokud vzorky vystříkly na folii, je nutné před jejím sejmutím destičku odstředit při 100 x g 10 sekund,
|  |  |  | g) | Opakujte kroky a) až f) ještě dvakrát, k dosažení celkem 3 promytí.
POZNÁMKA: Před odstraněním supernatantu ze třetího promytí připravte PCR master mix, jak je popsáno níže.
|  | 19 | Připravte PCR master mix v souladu s následujícími objemy:
|  |  |  | Reagencie |  | Objem pro 1 vzorek (μl) | Objem pro požadovaný počet vzorků + 10% rezerva (μl): | 0
|  |  |  | Sterilní voda |  | 20 | 0
|  |  |  | PCR Mix (nebo PCR mix - 1) |  | 20 | 0
|  |  |  | Celkem |  | 40 | 0
POZNÁMKA: PCR Mix se používá v dalších krocích protokolu. Nevyhazujte tuto lahvičku.
|  | 20 | Odstraňte fólii Microseal B a umístěte destičku na magnetický stojanna 30 sekund, aby se kuličky shromáždily v jamkách souběžně s magnetem,
|  | 21 | Pomocí pipety nastavené na 100 μl odsajte a zlikvidujte supernatant z posledního promytí pufrem Tagmentation Wash Buffer.
|  | 22 | Sejměte destičku z magnetického stojanu.
|  | 23 | Do každé jamky přidejte 40 μl PCR master mixu.
|  | 24 | Přelepte destičku fólií Microseal B.
|  | 25 | Pomocí třepačky při 1800 otáčkách za minutu třepejte při pokojové teplotě 2 minuty.
|  | 26 | Odstřeďujte destičku při 100 x g po dobu 10 sekund, aby byly všechny kuličky suspendovány v PCR master mixu.
Pro zkumavky s indexy (T),
|  |  | 27 | (T) Vortexujte a krátce odstřeďte zkumavky s indexy, aby byl veškerý obsah na dně zkumavky.
|  |  | 28 | (T) Odstraňte fólii Microseal B z destičky se vzorky.
|  |  | 29 | (T) Do každé jamky se přidá 5 μl indexu i7 podle Plate Layout na listu 1.0 Sample_Prep.
|  |  | 30 | (T) Do každé jamky se přidá 5 μl indexu i5 podle Plate Layout na listu 1.0 Sample_Prep.
Pokračujte krokem 31.
Pro destičku s indexy (P),
|  |  | 27 | (P) Třepejte na třepačce destičku s indexy při 1800 ot./min po dobu 1 minuty.
|  |  | 28 | (P) Centrifugujte indexovou destičku při 100 x g po dobu 10 sekund, abyste zajistili, že veškerý objem je na dně jamky.
|  |  | 29 | (P) Odstraňte fólii Microseal B z destičky se vzorky.
|  |  | 30 | (P) Zkontrolujte správnou orientaci destičky a správnou sadu indexů. Fólii nestrhávejte. Špičkou propíchněte fólii destičky s indexy. Pomocí nové špičky, přeneste 10 μl kombinovaných indexů z destičky s indexy do každé jamky v destičce se vzorky v souladu s Plate Layout na listu 1.0 Sample_Prep.
Pokračujte krokem 31.
|  | 31 | Přelepte destičku novou fólií Microseal B.
|  | 32 | Při pokojové teplotě třepejte při 1800 otáčkách za minutu 1 minutu.
|  | 33 | Centrifugujte destičku při 100 x g po dobu 10 sekund, abyste shromáždili veškerý obsah na dně jamky. Vizuálně zkontrolujte, zda jsou kuličky v roztoku stále rovnoměrně rozloženy. Pokud kuličky nejsou v jamce rovnoměrně rozloženy, opakujte protřepávání podle kroku 32.
|  | 34 | Umístěte destičku do termocykleru a spusťte program Indexing PCR.
Použitý termocykler:
Název programu:
Program Indexing PCR:
Reakční objem: 50 µL
Teplota víka: 105 °C
|  |  |  | # | Krok | Teplota | Čas | Počet cyklů
|  |  |  | 1 | Primární extenze | 72°C | 3 minuty | 1
|  |  |  | 2 | Počáteční denaturace | 98°C | 3 minuty | 1
|  |  |  | 3 | Denaturace | 98°C | 20 sekund | 9
|  |  |  | 4 | Annealing | 60°C | 30 sekund
|  |  |  | 5 | Extenze | 72°C | 3 minuty
|  |  |  | 6 | Finální extenze | 72°C | 3 minuty | 1
|  |  |  | 7 | Udržování teploty do vyjmutí destičky | 10°C | Do vyjmutí destičky | 1
POZNÁMKA: Bod bezpečného zastavení. Knihovny lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 týdne.
POZNÁMKA: Pokud okamžitě budete pokračovat kroky size selection a purifikace, zajistěte, aby Purification Beads byly před použitím ohřáty na pokojovou teplotu.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE KROKEM SIZE SELECTION/PURIFIKACE.

## 3.0 SS_Purification
ALLOSEQ TX – SIZE SELECTION/PURIFIKACE
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
VÝPOČTY PRO SIZE SELECTION
Pro optimální výtěžnost a výsledek size selection se objem vzorku, kuliček a supernatantu mění podle počtu analyzovaných vzorků. V následující tabulce jsou uvedeny objemy pro specifikovaný počet vzorků.
|  | Objemy reagencií |  |  |  |  | 6 – 24 vzorků/sdružení | 25 – 48 vzorků/sdružení | 49 – 96 vzorků/sdružení | Objemy pro daný run dle počtu vzorků
|  | Počet vzorků, které mají být sdruženy (max. pro uvedený rozsah) |  |  |  |  | 24 | 48 | 96 | 0
|  | Objem 1 vzorku ve sdružení (μl) |  |  |  |  | 10 | 5 | 2.5 | 0
|  | Objem zředěných kuliček na 1 vzorek (μl) |  |  |  |  | 50 | 25 | 12.5 | 0
|  | Objem přenášeného supernatantu na 1 vzorek (μl) |  |  |  |  | 55 | 27.5 | 13.75 | 0
|  | Objem neředěných kuliček na 1 vzorek (μl) |  |  |  |  | 4.4 | 2.2 | 1.1 | 0
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem na 1 vzorek nebo sdružení (μl) | Objem pro # vzorků (μl): | 0 | Požadovaná příprava
|  |  |  | Purification Beads (Purification Beads – 1 u kitů na 96 vzorků) |  |  | 2 °C až 8 °C KRABIČKA 5 | 24.7 | 0 |  | Dejte do pokojové teploty alespoň na 30 minut. Vraťte do lednice pro další použití.
|  |  |  | Sterilní voda |  |  | 15 °C až 30 °C Zajišťuje uživatel | Různý | 720 |  | Žádná příprava se nepožaduje.
|  |  |  | 100% etanol |  |  | 15 °C až 30 °C Zajišťuje uživatel | 1920 | 2880 |  | Připravte 80% etanol, jak je popsáno níže.
|  |  |  | Resuspension Buffer |  |  | 2 °C až 8 °C KRABIČKA 5 | Různý | 37 |  | Zahřejte na pokojovou teplotu. Vraťte do lednice pro další použití.
|  | 2 | Připravte si zředěné Purification Beads - ujistěte se, že neředěné Purification Beads jsou před použitím důkladně zvortexovány:
|  |  |  | Reagencie |  | Objem pro 1 vzorek (μl) | Objem pro požadovaný počet vzorků + 20% rezerva (μl): | 0
|  |  |  | Purification Beads |  | 0 | 0
|  |  |  | Sterile Water |  | 0 | 0
|  |  |  | Total |  | 0 | 0
|  | 3 | Prepare fresh 80% ethanol (2 washes per pool):
|  |  |  | Reagent |  | Volume per pool (µL) | Volume for Number of pools Required +50% Excess (µL): | 1
|  |  |  | 100% etanol |  | 1920 | 2880
|  |  |  | Sterilní voda |  | 480 | 720
|  |  |  | Celkem |  | 2400 | 3600
POZNÁMKA: Počet sdružení v experimentu je přednastaven na 1 ve žluté buňce výše. Ten lze v případě potřeby ručně přepsat. Úprava této buňky aktualizuje počet sdružení pro další kroky tohoto protokolu/sešitu. Pokud má být v tomto experimentu provedeno více než 1 sdružení se stejným počtem vzorků ve sdružení, měl by se objem Purification Beads odpovídajícím způsobem zvýšit (tj. zdvojnásobit pro dvě sdružení) a níže uvedené pokyny by měly být provedeny v plném rozsahu pro každé sdružení. Pokud je požadováno více sdružení s různým počtem vzorků, měla by být tato karta duplikována a buňky ve žluté barvě aktualizovány tak, aby odrážely počet vzorků. Pak budou správně upraveny objemy požadovaných/přenášených reagencií. Chcete-li kartku duplikovat, klikněte pravým tlačítkem na kteroukoli z karet a vyberte „Přesunout nebo Kopírovat“, ze seznamu vyberte „3.0 SS_Purification“, zaškrtněte políčko „Vytvořit kopii“ a klikněte na „OK“.
|  | 4 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 5 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml zkumavky (volitelné: v případě zvětšení promývacího objemu etanolu v kroku 14a z 1200 μl na 1500 μl lze použít zkumavky o objemu 2 ml).
PROTOKOL PRO SIZE SELECTION A PURIFIKACI
POZNÁMKA: Proces trvá přibližně 1 hodinu.
|  | 1 | Napipetujte 0 μl zředěných Purification Beads do 1,5ml zkumavky.
|  | 2 | Pipetou promíchejte Tagmentation Beads a supernatant v indexing PCR nebo protřepejte destičku s indexing PCR (1 minuta při 1800 ot./min) a poté přidejte 0 μl každého vzorku do zkumavky obsahující zředěné Purification Beads.
|  | 3 | Vortexujte každou zkumavku vysokou rychlostí po dobu 10 sekund, dokud vzorek nebude při vizuální kontrole homogenní.
|  | 4 | Inkubujte při pokojové teplotě 5 minut. Během této inkubace se největší fragmenty vážou na kuličky.
|  | 5 | Krátce stočte.
|  | 6 | Umístěte zkumavku na magnet na 2,5 minuty, aby se kuličky shromáždily ve zkumavce souběžně s magnetem. Pokud supernatant zůstane zakalený, ponechejte na magnetu, dokud nebude čistý.
|  | 7 | Přeneste 0 μl supernatantu do nové zkumavky.
|  | 8 | Přidejte 0 μl Purification Beads (neředěných) do zkumavky obsahující supernatant.
|  | 9 | Vortexujte každou zkumavku vysokou rychlostí po dobu 10 sekund.
|  | 10 | Inkubujte při pokojové teplotě 5 minut. Během této inkubace se největší fragmenty vážou na kuličky.
|  | 11 | Krátce stočte.
|  | 12 | Umístěte zkumavku na magnet na 2,5 minuty, aby se kuličky shromáždily ve zkumavce souběžně s magnetem. Pokud supernatant zůstane zakalený, ponechejte na magnetu, dokud nebude čistý.
|  | 13 | Pomocí pipety odsajte a zlikvidujte supernatant, přičemž kuličky ponechte ve zkumavce souběžně s magnetem,
|  | 14 | Ponechejte zkumavku na magnetu, dvakrát promyjte 80% etanolem:
|  |  | a) | Do každé zkumavky přidejte 1200 μl 80% etanolu,
|  |  | b) | Inkubujte při pokojové teplotě 30 sekund.
|  |  | c) | Pomocí pipety odsajte a zlikvidujte veškerý supernatant,
|  |  | d) | Opakujte kroky a) až c) tak, aby proběhla celkem 2 promytí.
|  | 15 | Odstraňte veškerý zbývající supernatant pipetou P20.
|  | 16 | Zkumavku nechte osychat na vzduchu po dobu 5 minut při pokojové teplotě, aby se zbytkový etanol mohl odpařit.
|  | 17 | Sejměte zkumavku z magnetu a do každé zkumavky přidejte 37 μl pufru Resuspension Buffer pro eluci cílových fragmentů.
|  | 18 | Vortexujte každou zkumavku vysokou rychlostí po dobu 10 sekund.
|  | 19 | Inkubujte při pokojové teplotě 5 minut.
|  | 20 | Krátce stočte.
|  | 21 | Umístěte zkumavku na magnet na 30 sekund, aby se kuličky ve zkumavce shromáždily souběžně s magnetem.
|  | 22 | Přeneste 35 μl supernatantu do nové zkumavky o objemu 1,5 ml a uschovejte nebo můžete pokračovat hybridizací.
POZNÁMKA: Bod bezpečného zastavení. Knihovny lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE HYBRIDIZACÍ. (NEBO VOLITELNĚ KVANTIFIKACÍ/VIZUALIZACÍ NA LISTU 8.0 QUBIT NEBO 9.0 TAPESTATION)

## 4.0 Hybridisation
ALLOSEQ TX - HYBRIDIZACE
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem na 1 sdružení (μl) | Objem pro # sdružení (μl): | 1 | Požadovaná příprava
|  |  |  | AlloSeq Tx Probe Panel |  |  | –15 °C až –25 °C KRABIČKA 4 | 10 | 10 |  | Dejte do pokojové teploty.
|  |  |  | Hybridisation Buffer 1 |  |  | –15 °C až –25 °C KRABIČKA 4 | 50 | 50 |  | Umístěte do Hybexu při teplotě 58 °C na 15 minut. Zvortexujte a vizuálně zkontrolujte, pokud sraženina přetrvává, pak inkubujte při 58 ͦC po dobu dalších 15 minut.
|  |  |  | Hybridisation Buffer 2 |  |  | 2 °C až 8 °C KRABIČKA 5 | 10 | 10 |  | Dejte do pokojové teploty
|  | 2 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 3 | Připravte si požadovaný spotřební materiál uvedený níže:
PCR zkumavky/stripy a uzávěry.
PROTOKOL HYBRIDIZACE
POZNÁMKA: Neexistuje žádný bezpečný bod zastavení – nutné je dokončit zachycování a promývání. Se sdružením je nutné z fáze hybridizace při 62 °C okamžitě přejít do fáze zachycování a promývání.
POZNÁMKA: Příprava trvá přibližně 20 minut, v termocykleru minimálně 1,5 hodiny a maximálně 18 hodin (reakce ponechané přes noc, nebo až 18 hodin, musí být udržovány při teplotě 62 °C).
|  | 1 | Pro každou hybridizační reakci se do PCR zkumavky/stripu zkombinují následující reagencie v níže uvedeném pořadí:
|  |  |  | Reagencie |  | Objem pro 1 sdružení (μl)
|  |  |  | Sdružení vzorků |  | 30
|  |  |  | AlloSeq Tx Probe Panel |  | 10
|  |  |  | Hybridisation Buffer 1 |  | 50
|  |  |  | Hybridisation Buffer 2 |  | 10
|  |  |  | Celkem |  | 100
|  | 2 | S pipetou nastavenou na 70 μl, propipetujte 10krát každou hybridizační reakci, zavíčkujte a poté krátce odstřeďte.
|  | 3 | Pokud roztok zůstane zakalený, promíchejte pipetou ještě 6-8krát, uzavřete a poté krátce odstřeďte.
|  | 4 | Umístěte zkumavku/strip do termocykleru a spusťte program hybridizace.
Použitý termocykler:
Název programu:
Program hybridizace:
Reakční objem: 100 µL
Teplota víka termocykleru: 100 °C
|  |  |  | # | Krok | Teplota | Čas | Počet cyklů
|  |  |  | 1 | Denaturace | 98°C | 5 minut | 1
|  |  |  | 2 | Ramping | 98 °C – 62 °C, pokles o 2 °C/cyklus | 1 minuta | 1
|  |  |  | 3 | Návrat ke kroku 2 - 18x zopakovat (celkem 19 cyklů) s poklesem o 2 ͦC/cyklus.
|  |  |  | 4 | Hybridisation | 62°C | 60 minut | 1
|  |  |  | 5 | Udržování teploty | 62°C | Udržování teploty(nepřekračujte 18 hodin při 62 °C, včetně kroku #4) | 1
|  | 5 | Nechte zkumavku/strip v termocykleru, dokud nebudete připraveni pokračovat v zachycování. Ujistěte se, že capture beads (zachytávací kuličky) dosáhly pokojové teploty a že Capture Wash Buffer a Hybex jsou zahřáté na 58 °C.
Čas dokončení:
OKAMŽITĚ POKRAČUJTE V ZACHYCOVÁNÍ.

## 5.0 Capture
ALLOSEQ TX – ZACHYCOVÁNÍ
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Objem pro # sdružení (μl): | 1 | Požadovaná příprava
|  |  |  | Capture Beads |  |  | 2 °C až 8 °C KRABIČKA 5 | 250 | 250 |  | Dejte do pokojové teploty alespoň na 30 minut.
|  |  |  | Capture Wash Buffer |  |  | –15 °C až –25 °C KRABIČKA 4 | 800 | 900 |  | Před použitím předehřejte na 58 ͦC.
|  |  |  | Capture Elution Buffer 1 |  |  | –15 °C až –25 °C KRABIČKA 4 | 28.5 | 28.5 |  | Dejte do pokojové teploty.
|  |  |  | 2N NaOH |  |  | –15 °C až –25 °C KRABIČKA 4 | 1.5 | 1.5 |  | Dejte do pokojové teploty . Vraťte do mrazničky pro další použití.
|  |  |  | Capture Elution Buffer 2 |  |  | 2 °C až 8 °C KRABIČKA 5 | 4 | 4 |  | Dejte do pokojové teploty.
|  | 2 | Připravte eluční master mix z následujících reagencií:
|  |  |  | Reagencie |  | Objem pro 1 sdružení (μl) | Objem pro # sdružení (μl): | 1
|  |  |  | Capture Elution Buffer 1 |  | 28.5 | 28.5
|  |  |  | 2N NaOH |  | 1.5 | 1.5
POZNÁMKA: Roztok NaOH snadno absorbuje CO2 z ovzduší, mění pH a účinnost. Pokud se zkumavka 2N NaOH nepoužívá, pečlivě ji uzavřete.
|  | 3 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 4 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml mikrocentrifugační zkumavky,
PCR zkumavky/stripy a uzávěry.
PROTOKOL ZACHYCOVÁNÍ
POZNÁMKA: Proces trvá přibližně 1 hodinu.
|  | 1 | Pro každou hybridizační reakci přidejte 250 µl Capture Beads do nové zkumavky o objemu 1,5 ml.
|  | 2 | Přeneste 100 μl z každé hybridizační reakce do příslušné zkumavky obsahující Capture Beads.
|  | 3 | Vortexujte zkumavku při vysoké rychlosti po dobu 10 sekund. Necentrifugujte ani krátce nestáčejte.
|  | 4 | Zkumavku inkubujte 15 minut při teplotě 58 °C v Hybexu.
|  | 5 | Krátce stočte.
|  | 6 | Umístěte zkumavku na magnet na 30 sekund, aby se kuličky shromáždily ve zkumavce souběžně s magnetem.
|  | 7 | Pomocí pipety odsajte a zlikvidujte supernatant, přičemž kuličky ponechte ve zkumavce souběžně s magnetem.
|  | 8 | Promyjte třikrát zahřátým pufrem Capture Wash Buffer (promývací pufr), jak je popsáno níže:
POZNÁMKA: Pokud se Capture Wash Buffer nepoužívá, ponechte jej v Hybexu, aby se udržovala teplota 58 °C. Vyjměte pouze bezprostředně před přidáním do reakce v krocích 8b a 10. Pracujte rychle, abyste minimalizovali dobu, po kterou jsou sdružení vzorků a pufr při pokojové teplotě.
|  |  |  | a) | Vyjměte zkumavku z magnetu.
|  |  |  | b) | Přidejte 200 μl zahřátého promývacího pufru Capture Wash Buffer (58 °C).
|  |  |  | c) | Vortexujte zkumavku při vysoké rychlosti po dobu 10 sekund. Necentrifugujte ani krátce nestáčejte.
|  |  |  | d) | Inkubujte zkumavku při teplotě 58 °C v Hybexu 5 minut.
|  |  |  | e) | Krátce stočte a okamžitěumístěte zkumavku na magnet na 30 sekund, aby se kuličky shromažďovaly ve zkumavce souběžně s magnetem.
|  |  |  | f) | Pomocí pipety odsajte a zlikvidujte supernatant, přičemž kuličky ponechte ve zkumavce souběžně s magnetem.
|  |  |  | g) | Opakujte kroky a) až f) ještě dvakrát, k dosažení celkem 3 promytí.
|  | 9 | Vyjměte zkumavku z magnetu.
|  | 10 | Přidejte 200 μl zahřátého promývacího pufru Capture Wash Buffer (58 °C).
|  | 11 | Zvortexujte zkumavku při vysoké rychlosti po dobu 10 sekund. Necentrifugujte ani krátce nestáčejte.
|  | 12 | Přeneste celý obsah (promývací roztok a kuličky) do nové zkumavky o objemu 1,5 ml.
POZNÁMKA: Tento krok přenosu je rozhodující pro odstranění inhibitorů PCR, které mohou negativně ovlivnit celkový výsledek.
|  | 13 | Inkubujte zkumavku při teplotě 58 °C v Hybexu po dobu 5 minut.
|  | 14 | Okamžitě krátce odstřeďte a pak umístěte zkumavku na magnet na 30 sekund, aby se kuličky shromáždily souběžně s magnetem.
|  | 15 | Pomocí pipety odsajte a zlikvidujte supernatant, přičemž kuličky ponechte ve zkumavce souběžně s magnetem.
|  | 16 | Krátce odstřeďte a pak okamžitě umístěte zkumavku na magnet na 30 sekund, aby se kuličky shromáždily ve zkumavce souběžně s magnetem.
|  | 17 | Pomocí pipety P20 odsajte a zlikvidujte supernatant, přičemž kuličky ponechte ve zkumavce souběžně s magnetem.
|  | 18 | Zvortexujte výše připravený eluční master mix a poté vyjměte zkumavku z magnetu a do každé zkumavky přidejte 23 μl elučního master mixu.
|  | 19 | Vortexujte zkumavku při vysoké rychlosti po dobu 10 sekund.
|  | 20 | Inkubujte při pokojové teplotě 2 minuty.
|  | 21 | Krátce stočte
|  | 22 | Umístěte zkumavku na magnet na 30 sekund, aby se kuličky shromáždily souběžně s magnetem.
|  | 23 | Přeneste 21 μl supernatantu do nové PCR zkumavky/stripu.
|  | 24 | Přidejte 4 μl pufru Capture Elution Buffer 2.
|  | 25 | Propipetujte 6-8krát. Konečný objem je 25 μl.
POZNÁMKA: Bod bezpečného zastavení. Knihovny lze skladovat při teplotě –15 °C až –25 °C po dobu až 24 hodin.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE V POST ENRICHMENT PCR.

## 6.0 Enrich_PCR
ALLOSEQ TX - POST ENRICHMENT PCR
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem pro1 sdružení (μl) | Objem pro # sdružení (μl): | 1 | Požadovaná příprava
|  |  |  | PCR Primers |  |  | –15 °C až –25 °C KRABIČKA 4 | 5 | 5 |  | Rozmrazte na ledu. Promíchejte obracením a poté krátce odstřeďte.
|  |  |  | PCR Mix (or PCR Mix-2 for 96 test kits) |  |  | –15 °C až –25 °C KRABIČKA 4 | 20 | 20 |  | Rozmrazte při pokojové teplotě a poté dejte na led.
|  | 2 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 3 | Připravte si položky uvedené níže:
PCR zkumavky obsahující knihovny z kroku 5.0 Capture.
PROTOKOL PRO POST ENRICHMENT PCR
POZNÁMKA: Příprava trvá přibližně 5 minut, v termocykleru 1:40 hodiny.
|  | 1 | Přidejte 5 µl PCR Primerů do knihoven (z kroku 5.0 Capture).
|  | 2 | Do zkumavky přidejte 20 μl PCR Mixu.
|  | 3 | Zamíchejte pipetou 10krát.
|  | 4 | Krátce stočte.
|  | 5 | Umístěte zkumavku/strip do termocykleru a spusťte program Post enrichment PCR.
Použitý termocykler:
Název programu:
Program Post enrichment PCR:
Reakční objem: 50 µL
Teplota víka: 105 °C
|  |  |  | # | Krok | Teplota | Čas | Počet cyklů
|  |  |  | 1 | Denaturace | 98°C | 30 sekund | 1
|  |  |  | 2 | Denaturace | 98°C | 1 minuta | 17
|  |  |  | 3 | Annealing | 60°C | 30 sekund
|  |  |  | 4 | Extenze | 72°C | 3 minuty
|  |  |  | 5 | Finální extenze | 72°C | 5 minut | 1
|  |  |  | 6 | Udržování teploty do vyjmutí zkumavky | 10°C | Do vyjmutí zkumavky | 1
POZNÁMKA: Bod bezpečného zastavení. Knihovny lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 týdne.
POZNÁMKA: Pokud okamžitě přistoupíte k purifikaci, ujistěte se, že jsou Purification Beads ohřáty na pokojovou teplotu.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE V PURIFIKACI.

## 7.0 Purification
ALLOSEQ TX – PURIFIKACE POST ENRICHMENT PCR
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Objem pro # sdružení (μl): | 1 | Požadovaná příprava
|  |  |  | Purification Beads (Purification Beads – 2 u kitů pro 96 vzorků) |  |  | 2 °C až 8 °C KRABIČKA 5 | 27 | 27 |  | Zahřejte napokojovou teplotu (alespoň 30 minut).
|  |  |  | Sterilní voda |  |  | 15 °C až 30 °C Dodáno uživatelem | 80 | 120 |  | Žádná příprava se nepožaduje.
|  |  |  | 100% etanol |  |  | 15 °C až 30 °C Dodáno uživatelem | 320 | 480 |  | Připravte 80% etanol, jak je popsáno níže.
|  |  |  | Resuspension Buffer |  |  | 2 °C až 8 °C KRABIČKA 5 | 32 | 32 |  | Zahřejte na pokojovou teplotu. Vraťte zpátky do lednice pro další použití.
|  | 2 | Připravte si čerstvý 80% etanol (2 promytí na každé sdružení):
|  |  |  | Reagencie |  | Objem pro 1 sdružení (μl) | Objem pro požadovaný počet sdružení + 50% rezerva (μl): | 1
|  |  |  | 100% etanol |  | 320 | 480
|  |  |  | Sterilní voda |  | 80 | 120
|  |  |  | Celkem |  | 400 | 600
|  | 3 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 4 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml mikrocentrifugační zkumavky.
PROTOKOL PURIFIKACE
POZNÁMKA: Proces trvá přibližně 30 minut.
|  | 1 | Přidejte 27 μl důkladně zvortexovaných Purification Beads do nové 1,5ml zkumavky (pro každé sdružení 1).
|  | 2 | Přeneste 45 μl z každé Post enrichment PCR do příslušné zkumavky obsahující Purification Beads.
|  | 3 | Vortexujte každou zkumavku při vysokých otáčkách10 sekund.
|  | 4 | Krátce stočte.
|  | 5 | Inkubujte při pokojové teplotě 5 minut.
|  | 6 | Umístěte zkumavku na magnet na 30 sekund, aby se kuličky shromáždily souběžně s magnetem.
|  | 7 | Pomocí pipety odsajte a zlikvidujte supernatant, přičemž kuličky ponechte ve zkumavce souběžně s magnetem.
|  | 8 | Ponechejte zkumavku na magnetu, dvakrát promyjte 80% etanolem:
|  |  | a) | Do každé zkumavky přidejte 200 μl 80% etanolu,
|  |  | b) | Inkubujte při pokojové teplotě 30 sekund.
|  |  | c) | Pomocí pipety odsajte a zlikvidujte veškerý supernatant,
|  |  | d) | Opakujte kroky a) až c) ještě jednou (celkem jsou třeba 2 promytí).
|  | 9 | Odstraňte veškerý zbývající supernatant pipetou P20.
|  | 10 | Zkumavku nechte osychat na vzduchu 5 minut při pokojové teplotě, aby se zbytkový etanol mohl odpařit.
|  | 11 | Sejměte zkumavku z magnetu a do každé zkumavky přidejte 32 μl pufru Resuspension Buffer, aby došlo k eluci požadovaných fragmentů.
|  | 12 | Vortexujte každou zkumavku při vysokých otáčkách 10 sekund.
|  | 13 | Inkubujte při pokojové teplotě 5 minut.
|  | 14 | Krátce stočte.
|  | 15 | Umístěte zkumavku na magnet na 30 sekund, aby se kuličky shromáždily ve zkumavce souběžně s magnetem.
|  | 16 | Přeneste 30 μl supernatantu do nové zkumavky o objemu 1,5 ml.
POZNÁMKA: Bod bezpečného zastavení. Knihovny lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
Čas dokončení:
UKONČENÍ PROTOKOLU. PŘEJDĚTE KE KVANTIFIKACI.

## 8.0 Qubit
ALLOSEQ TX – KVANTIFIKACE QUBIT
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Objem pro # sdružení (μl): | 1 | Požadovaná příprava
|  |  |  | Qubit BR buffer |  |  | –25 °C až 30 °C zajišťuje uživatel | 199 | 796 |  | Žádná příprava se nepožaduje.
|  |  |  | Qubit BR dye |  |  | –25 °C až 30 °C zajišťuje uživatel | 1 | 4 |  | Žádná příprava se nepožaduje.
|  |  |  | BR Standard #1 |  |  | 2 °C až 8 °C zajišťuje uživatel | 10 | 10 |  | Zahřejte na pokojovou teplotu.
|  |  |  | BR Standard #2 |  |  | 2 °C až 8 °C zajišťuje uživatel | 10 | 10 |  | Zahřejte na pokojovou teplotu .
POZNÁMKA: Počet sdružení se automaticky zkopíruje dle údajů v listu hybridizace, ale v případě potřeby jej lze ručně přepsat.
|  | 2 | Připravte si požadovaný spotřební materiál uvedený níže:
Zkumavky Qubit,
Zkumavka o objemu 1,5 ml nebo 5 ml pro přípravu pracovního roztoku, v závislosti na výše uvedeném objemu.
KVANTIFIKACE OBOHACENÝCH KNIHOVEN
|  | 1 | Připravte 2 zkumavky pro standardy a jednu pro každé sdružení.
|  | 2 | Připravte pracovní roztok dle následujícího rozpisu:
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | Qubit buffer |  | 796
|  |  |  | Qubit dye |  | 4
|  |  |  | Celkem |  | 800
POZNÁMKA: Tento objem pracovního roztoku zahrnuje rezervu.
|  | 3 | Vortexujte pracovní roztok po dobu 2–3 sekund a poté krátce odstřeďte.
|  | 4 | Dejte 190 μl pracovního roztoku do každé zkumavky pro standardy.
|  | 5 | Dejte 198 μl pracovního roztoku do každé zkumavky pro sdružení.
|  | 6 | Dejte 10 μl zásobního roztoku daného standardu do příslušné zkumavky.
|  | 7 | Dejte 2 μl z každého sdružení do příslušné zkumavky.
|  | 8 | Vortexujte všechny zkumavky po dobu 2–3 sekund a poté krátce odstřeďte.
|  | 9 | Inkubujte při pokojové teplotě 2 minuty.
|  | 10 | Vložte zkumavky do fluorometru Qubit a odečtěte hodnoty (další informace viz protokol Qubit).
|  | 11 | Zaznamenejte hodnoty Qubit do níže uvedené tabulky.
|  |  | Sdružení # |  | Zkumavka # | ID obohaceného sdružení | 1. odečet (ng/μl) | 2. odečet (ng/μl) | Průměr (ng/μl)
1
2
3
4
5
6
7
8
|  | 12 | Zaznamenejte výsledky do woorkbooku a vypočítejte průměr.
Čas dokončení:
UKONČENÍ PROTOKOLU. PROVEĎTE ŘEDĚNÍ A DENATURACI.

## 9.0 TapeStation(Optional)
ALLOSEQ TX - TAPESTATION VISUALISATION
| ID experimentu: |  |  |  | Zadejte
Operátor:
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie |  | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Objem pro # sdružení (μl): | 1 | Požadovaná příprava
|  |  |  | D1000 Ladder |  |  | 2 °C až 8 °C Dodáno uživatelem | 1 | 1 |  | Zahřejte na pokojovou teplotu.
|  |  |  | D1000 Sample Buffer |  |  | 2 °C až 8 °C Dodáno uživatelem | 3 | 6 |  | Zahřejte na pokojovou teplotu.
POZNÁMKA: Počet sdružení se automaticky zkopíruje dle údajů v listu hybridizace, ale v případě potřeby jej lze ručně přepsat.
|  | 2 | Připravte si požadovaný spotřební materiál uvedený níže:
D1000 ScreenTape,
TapeStation Optical tube strips (8x Strip),
TapeStation Optical tube strip caps (8x Strip)
VIZUALIZACE DISTRIBUCE VELIKOSTÍ OBOHACENÝCH KNIHOVEN
|  | 1 | Přeneste 1 μl z každého sdružení do nové zkumavky.
|  | 2 | Přidejte 1 μl D1000 Ladder do referenční zkumavky.
|  | 3 | Přidejte 3 μl pufru D1000 Sample Buffer do každé zkumavky určené pro testování sdružení a do referenční zkumavky.
|  | 4 | Všechny zkumavky utěsněte uzávěry.
|  | 5 | Důkladně vortexujte všechny zkumavky za použití vortexu IKA při 2000 otáčkách za minutu po dobu 1 minuty.
|  | 6 | Krátce odstřeďte, abyste zajistili, že obsah zkumavek bude na dně.
|  | 7 | Odvíčkujte a vložte zkumavky se vzorky do přístroje 2200 TapeStation.
|  | 8 | Vyberte požadované pozice zkumavek v 2200 TapeStation Controller Software a spusťte analýzu.
|  | 9 | Po dokončení testu spusťte TapeStation Analysis Software a zobrazte výsledky.
|  | 10 | Zaznamenejte výsledky do níže uvedené tabulky:
|  |  | Sdružení # |  | Zkumavka # | ID obohaceného sdružení | Minimální velikost fragmentů (bp) | Poloha píku (bp) | Maximální velikost fragmentů (bp)
1
2
3
4
5
6
7
8
Čas dokončení:
UKONČENÍ PROTOKOLU. PROVEĎTE ŘEDĚNÍ A DENATURACI.

## 10.0 PhiX
ALLOSEQ TX - ŘEDĚNÍ A DENATURACE PHIX
| Operátor: |  |  |  |  | POZNÁMKA: ZELENÁ pole musí vyplnit operátor během nastavení.
| Datum: |  |  |  |  | POZNÁMKA: ŽLUTÁ pole je nutno před tiskem vyplnit.
Čas zahájení:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Požadovaná příprava
|  |  |  | 10 nM PhiX |  | –15 °C až –25 °C Dodáno uživatelem | 2 | Rozmrazte a poté uchovávejte na ledu.
|  |  |  | 2N NaOH |  | –15 °C až –25 °C KRABIČKA 4 | 5 | Zahřejte na pokojovou teplotu.
|  |  |  | Sterile Water |  | 15 °C až 30 °C Dodáno uživatelem | 45 | Žádná příprava se nepožaduje.
|  |  |  | Resuspension Buffer |  | 2 °C až 8 °C KRABIČKA 5 | 3 | Zahřejte na pokojovou teplotu.
|  |  |  | HT1 (se sekvenační kártridží) |  | –15 °C až –25 °C Dodáno uživatelem | 990 | Rozmrazte a poté uchovávejte na ledu.
|  | 2 | Připravte pracovní roztok 0,2N NaOH
|  |  |  | Reagencie | Objem (μl)
|  |  |  | Sterilní voda | 45
|  |  |  | 2N NaOH | 5
POZNÁMKA: Roztok NaOH snadno absorbuje CO2 z ovzduší, mění pH a svou účinnost. Pokud se zkumavka 2N NaOH nepoužívá, dobře ji uzavřete.
|  | 3 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 4 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml mikrocentrifugační zkumavky
ŘEDĚNÍ A DENATURACE PHIX PRO MISEQ A MINISEQ, VÝSLEDKEM JE 20pM DENATUROVANÝ PHIX
|  | 1 | Dejte 3 μl pufru Resuspension Buffer do nové zkumavky.
|  | 2 | Do zkumavky přidejte 2 µL 10nM PhiX
|  | 3 | Do zkumavky přidejte 5 μl pracovního roztoku 0,2N NaOH (příprava viz výše).
|  | 4 | Zkumavku zvortexujte a krátce odstřeďte.
|  | 5 | Inkubujte při pokojové teplotě 5 minut.
|  | 6 | Přidejte 990 μl předchlazeného HT1 do zkumavky obsahující denaturovaný PhiX.
|  | 7 | Invertujte zkumavku (nevortexujte) a poté zkumavku krátce odstřeďte.
POZNÁMKA: Výsledkem je 1 ml 20pM PhiX, který je připraven k použití v přístroji MiSeq. Pro přístroj MiniSeq proveďte další ředění (na 5 pM) popsané níže. Denaturovaný PhiX lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
Čas dokončení:
POUZE PRO PŘÍSTROJ MINISEQ: DÁLE ZŘEĎTE DENATUROVANÝ PHIX NA 5 pM
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
Celkový požadovaný objem (μl)
|  | 2 | Zřeďte PhiX na 5 pM smícháním následujících objemů reagencií v mikrocentrifugační zkumavce:
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | 20pM denaturovaný PhiX |  | 0
|  |  |  | Chlazený HT1 |  | 0
|  |  |  | Celkem |  | 0
|  | 3 | Invertujte zkumavku (nevortexujte) a poté zkumavku krátce odstřeďte.
POZNÁMKA: Denaturovaný PhiX lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
Čas dokončení:
POUZE PRO PŘÍSTROJ ISEQ: ZŘEĎTE PHIX NA 20 pM (BEZ DENATURACE)
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
Celkový požadovaný objem (μl)
|  | 2 | Zřeďte PhiX na 20 pM smícháním následujících objemů reagencií v mikrocentrifugační zkumavce:
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | 10 nM PhiX |  | 0
|  |  |  | Resuspension Buffer |  | 0
|  |  |  | Celkem |  | 0
|  | 3 | Invertujte zkumavku (nevortexujte) a poté zkumavku krátce odstřeďte.
POZNÁMKA: Denaturovaný PhiX lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
Čas dokončení:
UKONČENÍ PROTOKOLU.

## 10.1 MiSeq (hidden)
ALLOSEQ TX - ŘEDĚNÍ A DENATURACE PRO MISEQ
| ID flowcelly #: |  |  |  |  | POZNÁMKA: ZELENÁ pole musí vyplnit operátor během nastavení.
| Operátor: |  |  |  |  | POZNÁMKA: ŽLUTÁ pole je nutno před tiskem vyplnit.
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Požadovaná příprava
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | –15 °C až –25 °C Připraveno uživatelem | #DIV/0! | Rozmrazte a poté uchovávejte na ledu.
|  |  |  | 2N NaOH |  | –15 °C až –25 °C KRABIČKA 4 | 5 | Zahřejte na pokojovou teplotu.
|  |  |  | Sterilní voda |  | 15 °C až 30 °C Dodáno uživatelem | 45 | Žádná příprava se nepožaduje.
|  |  |  | HT1 (s kártridží MiSeq) |  | –15 °C až –25 °C Dodáno uživatelem | 1226.4 | Rozmrazte a poté uchovávejte na ledu.
|  |  |  | Resuspension Buffer |  | 2 °C až 8 °C KRABIČKA 5 | #DIV/0! | Zahřejte na pokojovou teplotu.
|  |  |  | 20pM PhiX (dříve zředěné) |  | –15 °C až –25 °C Dodáno uživatelem | 3.6 | Rozmrazte a poté uchovávejte na ledu.
|  | 2 | Připravte pracovní roztok 0,2N NaOH
|  |  |  | Reagencie | Objem (μl)
|  |  |  | Sterilní voda | 45
|  |  |  | 2N NaOH | 5
POZNÁMKA: Roztok NaOH snadno absorbuje CO2 z ovzduší, mění pH a svou účinnost. Pokud se zkumavka 2N NaOH nepoužívá, pečlivě ji zavřete.
|  | 3 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 4 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml mikrocentrifugační zkumavky
INFORMACE POTŘEBNÉ PRO VÝPOČTY ŘEDĚNÍ na 4 nM
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
Koncentrace obohaceného sdružení (ng/μl)
|  |  |  | Velikost fragmentů obohaceného sdružení (bp) |  | 800
|  |  |  | Koncentrace obohaceného sdružení (nM) |  | 0
|  |  |  | Celkový požadovaný objem (μl) |  | 30
NAŘEDĚNÍ OBOHACENÉHO SDRUŽENÍ NA 4 nM
|  | 1 | Zřeďte vzorek na 4 nM smícháním následujících objemů reagencií v mikrocentrifugační zkumavce:
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Celkem |  | #DIV/0!
|  | 2 | Před dalším použitím naředěné sdružení zvortexujte a krátce odstřeďte.
DENATURACE 4nM OBOHACENÉHO SDRUŽENÍ, VÝSLEDKEM JE 20pM DENATUROVANÉ SDRUŽENÍ.
|  | 1 | Dejte 5 μl z 4nM obohaceného sdružení do nové zkumavky.
|  | 2 | Do zkumavky přidejte 5 μl pracovního roztoku 0,2N NaOH (naředění viz výše).
|  | 3 | Zkumavku zvortexujte a krátce odstřeďte.
|  | 4 | Inkubujte při pokojové teplotě 5 minut.
|  | 5 | Přidejte 990 μl předchlazeného HT1 do zkumavky obsahující denaturované obohacené sdružení.
|  | 6 | Invertujte zkumavku (nevortexujte) a poté zkumavku krátce odstřeďte.
POZNÁMKA: Výsledkem je 1 ml 20pM sdružení. Sdružení lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
INFORMACE POTŘEBNÉ PRO VÝPOČET FINÁLNÍHO ŘEDĚNÍ
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
|  |  |  | Počet sdružení, které mají být pro sekvenování sloučeny |  | 1
|  |  |  | Nanášecí koncentrace (pM) |  | 12
|  |  |  | Podíl PhiX(%) |  | 1%
|  |  |  | Objem ředění, minimálně 600 (μl) |  | 600
ŘEDĚNÍ OBOHACENÉHO SDRUŽENÍ NA NANÁŠECÍ KONCENTRACI S OHLEDEM NA PODÍL PHIX
|  | 1 | Nařeďte sdružení na nanášecí koncentraci smícháním objemů, které jsou uvedeny v následující tabulce (použijte mikrocentrifugační zkumavku):
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | 20pM sdružení # 1 |  | 360
|  |  |  | 20pM sdružení # 2 |  | 0
|  |  |  | 20pM sdružení # 3 |  | 0
|  |  |  | 20pM sdružení # 4 |  | 0
|  |  |  | 20pM denaturovaný PhiX |  | 3.6
|  |  |  | Chlazený HT1 |  | 236.4
|  |  |  | Celkem |  | 600
|  | 2 | Invertujte zkumavku (nevortexujte) a poté ji krátce odstřeďte.
|  | 3 | Dejte stranou na led, dokud nebude připravena reagenční kazeta MiSeq. Po té do kazety napipetujte 600 μl zředěné knihovny.
POZNÁMKA: Sdružení lze skladovat při teplotě –15 °C až –25 °C až 48 hodin před sekvenováním.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE NA PŘÍSTROJI MISEQ PODLE PROTOKOLU PŘÍSTROJE.

## 10.2 MiniSeq (hidden)
ALLOSEQ TX - ŘEDĚNÍ A DENATURACE PRO MINISEQ
| ID flowcelly #: |  |  |  |  | POZNÁMKA: ZELENÁ pole musí vyplnit operátor během nastavení.
| Operátor: |  |  |  |  | POZNÁMKA: ŽLUTÁ pole je nutno před tiskem vyplnit.
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Požadovaná příprava
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | –15 °C až –25 °C Připraveno uživatelem | #DIV/0! | Rozmrazte a poté uchovávejte na ledu.
|  |  |  | 2N NaOH |  | –15 °C až –25 °C KRABIČKA 4 | 5 | Zahřejte na pokojovou teplotu.
|  |  |  | Sterilní voda |  | 15 °C až 30 °C Dodáno uživatelem | 95 | Žádná příprava se nepožaduje.
|  |  |  | HT1 (with sequencing cartridge) |  | –15 °C až –25 °C Dodáno uživatelem | 1323.4 | Rozmrazte a poté uchovávejte na ledu.
|  |  |  | Resuspension Buffer |  | 2 °C až 8 °C KRABIČKA 5 | #DIV/0! | Zahřejte na pokojovou teplotu.
|  |  |  | 200 mM Tris-HCl, pH 7.0 |  | 15 °C až 30 °C Dodáno uživatelem | 5 | Žádná příprava se nepožaduje.
|  |  |  | 5 pM PhiX (dříve zředěné) |  | –15 °C až –25 °C Dodáno uživatelem | 1.6 | Rozmrazte a poté uchovávejte na ledu.
|  | 2 | Připravte pracovní roztok 0,1N NaOH
|  |  |  | Reagencie | Objem (μl)
|  |  |  | Sterilní voda | 95
|  |  |  | 2N NaOH | 5
POZNÁMKA: Roztok NaOH snadno absorbuje CO2 z ovzduší, mění pH a svou účinnost. Pokud se zkumavka 2N NaOH nepoužívá, pečlivě ji zavřete.
|  | 3 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 4 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml mikrocentrifugační zkumavky
INFORMACE POTŘEBNÉ PRO VÝPOČTY ŘEDĚNÍ na 4 nM
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
Koncentrace obohaceného sdružení (ng/μl)
|  |  |  | Velikost fragmentů obohaceného sdružení (bp) |  | 800
|  |  |  | Koncentrace obohaceného sdružení (nM) |  | 0
|  |  |  | Celkový požadovaný objem (μl) |  | 100
NAŘEDĚNÍ OBOHACENÉHO SDRUŽENÍ NA 1 nM
|  | 1 | Zřeďte vzorek na 1 nM smícháním následujících objemů reagencií v mikrocentrifugační zkumavce:
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Před dalším použitím naředěné sdružení zvortexujte a krátce odstřeďte.
DENATURACE 1nM OBOHACENÉHO SDRUŽENÍ, VÝSLEDKEM JE 5pM DENATUROVANÉ SDRUŽENÍ.
|  | 1 | Dejte 5 μl z 1nM obohaceného sdružení do nové zkumavky.
|  | 2 | Do zkumavky přidejte 5 μl pracovního roztoku 0,1N NaOH (naředění viz výše).
|  | 3 | Zkumavku zvortexujte a krátce odstřeďte.
|  | 4 | Inkubujte při pokojové teplotě 5 minut.
|  | 5 | Přidejte 5 µL 200mM Tris-HCl (pH 7.0).
|  | 6 | Zkumavku zvortexujte a krátce odstřeďte.
|  | 7 | Přidejte 985 μl předchlazeného HT1 do zkumavky obsahující denaturované obohacené sdružení.
|  | 8 | Invertujte zkumavku (nevortexujte) a poté zkumavku krátce odstřeďte.
POZNÁMKA: Výsledkem je 1 ml 5pM sdružení. Sdružení lze skladovat při teplotě –15 °C až –25 °C po dobu až 1 měsíce.
INFORMACE POTŘEBNÉ PRO VÝPOČET FINÁLNÍHO ŘEDĚNÍ
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
|  |  |  | Počet sdružení, které mají být pro sekvenování sloučeny |  | 1
|  |  |  | Nanášecí koncentrace (pM) |  | 1.6
|  |  |  | Podíl PhiX(%) |  | 1%
|  |  |  | Objem ředění, minimálně 500 (μl) |  | 500
ŘEDĚNÍ OBOHACENÉHO SDRUŽENÍ NA NANÁŠECÍ KONCENTRACI S OHLEDEM NA PODÍL PHIX
|  | 1 | Nařeďte sdružení na nanášecí koncentraci smícháním objemů, které jsou uvedeny v následující tabulce (použijte mikrocentrifugační zkumavku):
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | 5pM sdružení # 1 |  | 160
|  |  |  | 5pM sdružení # 2 |  | 0
|  |  |  | 5pM sdružení # 3 |  | 0
|  |  |  | 5pM sdružení # 4 |  | 0
|  |  |  | 5pM denaturovaný PhiX |  | 1.6
|  |  |  | Chlazený HT1 |  | 338.4
|  |  |  | Celkem |  | 500
|  | 2 | Invertujte zkumavku (nevortexujte) a poté ji krátce odstřeďte.
|  | 3 | Dejte stranou na led, dokud nebude připravena reagenční kazeta MiniSeq. Po té do kazety napipetujte 500 μl zředěné knihovny.
POZNÁMKA: Sdružení lze skladovat při teplotě –15 °C až –25 °C až 48 hodin před sekvenováním.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE NA PŘÍSTROJI MINSEQ PODLE PROTOKOLU PŘÍSTROJE.

## 10.3 iSeq (hidden)
ALLOSEQ TX - ŘEDĚNÍ PRO ISEQ
| ID flowcelly #: |  |  |  |  | POZNÁMKA: ZELENÁ pole musí vyplnit operátor během nastavení.
| Operátor: |  |  |  |  | POZNÁMKA: ŽLUTÁ pole je nutno před tiskem vyplnit.
Datum:
Čas zahájení:
PŘÍPRAVA REAGENCIÍ A EVIDENCE ŠARŽÍ
|  | 1 | Shromážděte požadované reagencie (objem je uveden v tabulce níže) a připravte podle tabulky:
|  |  |  | Reagencie | Šarže # a datum exspirace | Podmínky skladování | Objem pro 1 sdružení (μl) | Požadovaná příprava
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | –15 °C až –25 °C Připraveno uživatelem | #DIV/0! | Rozmrazte a poté uchovávejte na ledu.
|  |  |  | Resuspension Buffer |  | 2 °C až 8 °C KRABIČKA 5 | #DIV/0! | Zahřejte na pokojovou teplotu.
|  |  |  | 20pM PhiX (dříve zředěné) |  | –15 °C až –25 °C Dodáno uživatelem | 10 | Rozmrazte a poté uchovávejte na ledu.
|  | 2 | Před použitím zvortexujte a krátce odstřeďte všechny reagencie.
|  | 3 | Připravte si požadovaný spotřební materiál uvedený níže:
1,5ml mikrocentrifugační zkumavky
INFORMACE POTŘEBNÉ PRO VÝPOČTY ŘEDĚNÍ na 1 nM
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
Koncentrace obohaceného sdružení (ng/μl)
|  |  |  | Velikost fragmentů obohaceného sdružení (bp) |  | 800
|  |  |  | Koncentrace obohaceného sdružení (nM) |  | 0
|  |  |  | Celkový požadovaný objem (μl) |  | 100
NAŘEDĚNÍ OBOHACENÉHO SDRUŽENÍ NA 1 nM
|  | 1 | Zřeďte vzorek na 1 nM smícháním následujících objemů reagencií v mikrocentrifugační zkumavce:
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Celkem |  | #DIV/0!
|  | 2 | Před dalším použitím naředěné sdružení zvortexujte a krátce odstřeďte.
INFORMACE POTŘEBNÉ PRO VÝPOČET FINÁLNÍHO ŘEDĚNÍ
|  | 1 | Zadejte níže požadované informace, abyste mohli vypočítat vhodné ředění knihovny:
|  |  |  | Požadované informace |  | Hodnota
|  |  |  | Počet sdružení, které mají být pro sekvenování sloučeny |  | 1
|  |  |  | Nanášecí koncentrace (pM) |  | 200
|  |  |  | Podíl PhiX(%) |  | 1%
|  |  |  | Objem ředění, minimálně 20 (μl) |  | 100
ŘEDĚNÍ OBOHACENÉHO SDRUŽENÍ NA NANÁŠECÍ KONCENTRACI S OHLEDEM NA PODÍL PHIX
|  | 1 | Nařeďte sdružení na nanášecí koncentraci smícháním objemů, které jsou uvedeny v následující tabulce (použijte mikrocentrifugační zkumavku):
|  |  |  | Reagencie |  | Objem (μl)
|  |  |  | 1nM sdružení # 1 |  | 20
|  |  |  | 1nM sdružení # 2 |  | 0
|  |  |  | 1nM sdružení # 3 |  | 0
|  |  |  | 1nM sdružení # 4 |  | 0
|  |  |  | 20pM nedenaturovaný PhiX |  | 10
|  |  |  | Resuspension Buffer |  | 70
|  |  |  | Total |  | 100
|  | 2 | Invertujte zkumavku (nevortexujte) a poté ji krátce odstřeďte.
|  | 3 | Dejte stranou na led, dokud nebude připravena reagenční kazeta iSeq. Po té do kazety napipetujte 20 μl zředěné knihovny.
POZNÁMKA: Sdružení lze skladovat při teplotě –15 °C až –25 °C až 48 hodin před sekvenováním.
Čas dokončení:
UKONČENÍ PROTOKOLU. POKRAČUJTE NA PŘÍSTROJI ISEQ PODLE PROTOKOLU PŘÍSTROJE.

## VersionHistory
VERSION | REVISION DATE (DDMMMYY ) | REVISION DESCRIPTION (IFU095-5-CS v1,0 is translated from English master IFU095-5 v4,1)
IFU095-5 v1,0 | 03MAY2021 | Initial release of the CE-IVD AlloSeq Tx assay Early Pooling workbook. This workbook was built from IFU083-5_AlloSeq Tx Assay Workbook RUO, v6.0 Subsequent workbook updates: Deactivation of the 96 Set B/C/D Indexes options. Typographical error corrections. Return to storage after use for subsequent steps statements added in reagent tables where relevant. Added extra guidance around thoroughly mixing Purification Beads to step 1, tab 3.0 and and step 1, tab 7.0 following feedback from field team. Harmonised all instances of pulse spin to pulse-spin. Corrected all instances of 'hybex' to 'Hybex' and 'QuBit' to 'Qubit'. Step 3, tab 3.0 changed from 'Vortex each tube at high speed for 10 seconds' to 'Vortex each tube at high speed until the sample appears homogenous upon visual inspection'. Included clarification comment, "Hold (do not exceed 18 hours at 62°C, inclusive of step #4)" in hybridisation program table. Added 'If beads aren't evenly distributed, repeat shaking as per step 32' in step 33, Library_Prep tab. Added Qubit Broad Range (BR) clarification in reagent table, tab 8.0. Corrected all references to plates/wells to tubes/caps in tab 9.0. Workbook updated to include CE cover page and added CE0197 logo. Updated Volume per Sample (µL) for indexes from 20 to 5 in 2.0 Library_Prep to reflect the actual volume required according to step 29&30. Reissued by L.Langley 20 Oct 21
IFU095-5 v2,0 | 28Jan22 | Updated the Library Preparation to include instructions for using the index plate and separated from the index tube instructions. Added clarification to "Confirm the correct orientation of the plate and correct index set. Do not peel the foil seal." Reissued by E. Naughton on 27Apr22
IFU095-5 v3,0 | 28JUL2022 | Workbook updates: 1) Added ASTX9.1(96)-A-IVD and ASTX9.1(96)-B-IVD product codes. 2) Aligned phrasing between workbook and PDF IFU. Reissued by H. Meraj 29-Jul-22
IFU095-5 v 3,1 | 11Aug23 | Corrections to header and footers of hidden tabs. Reissued by N.Carrucan 11Aug23
IFU095-5 v 4,0 | 19Sep23 | Explanded import functionality to utilize 24 or 96 sample kits. Includes well drop down to select the starting position and macro to copy sample names to 1.0 Sample_Prep tab. Removed references to HistoTrac on 0.1 Import tab; can import an excel file, not just from HT. Added export functionality to export sample sheet as .csv file and remove any empty rows. Updated Qubit average.
IFU095-5 v4,1 | 28 Sep 23 | No changes to document itself. Password protection changed on Sample sheet tab to stop error message.
IFU095-5-CS v1.0 | 21Mar24 | First version of the workbook translated to Czech created

## WorkbookVerification (hidden)
EXCEL WORKBOOK VERIFICATION
WORKBOOK FILE |  | IFU095-5_AlloSeq Tx Assay Workbook |  | WORKBOOK LOCATION |  | S:\DocumentLibrary2\Issued\PDFs\Instructions for Use
INTENT OF WORKBOOK, INCLUDING HIGH LEVEL FUNCTIONALITY
This laboratory workbook is intended to be used to perform the AlloSeq Tx Assay Protocol. The Sample Preparation tab in used to define the experiment number, the probe set, sequencer, number of samples, and list the samples to be processed in a plate layout, which is converted to a table on the linear view tab, that can be used to generate a sample sheet for the sequencer on the sample sheet tab. The calculated the total number of samples is used throughout the workbook to perform volume calculations. The library preparation, size selection/purification, hybridisation, capture, enrichment PCR, purification, QuBit and TapStation tabs describe the required protocol. The VersionHistory tab is to record each issue/re-issue of the workbook, as well as the changes made to that version. This Workbook Verification tab is intended to detail the intended functionality and formulas used within the workbook, as well as give clear instructions on how to verify the functionality of the workbook. This tab is to be removed from copies shared externally.
INTENDED USERS OF WORKBOOK
RESEARCH AND DEVELOPMENT TEAM |  |  | Development team to create new workbooks as needed during product development projects.
OPERATIONS TEAM |  |  | Operations team to utilise workbooks during QC of new batches.
QA TEAM |  |  | QA team to review workbook/verification to issue workbook.
EXTERNAL CUSTOMERS |  |  | External customers to utilise workbooks to run samples.
INSTRUCTIONS FOR VERIFICATION
1) Print WorkbookVerification tab.
2) Enter the information as listed in the 'VERIFICATION METHOD' column of this tab (shown in red text).
3) Visually inspect workbook tabs and verify the information as listed below.
4) If the workbook functions as expected record Y in the pass column.
5) If the workbook does not function as expected, record N in the pass column.
6) If all functions pass, the workbook may be issued. If some functions fail, consult workbook owner for further action.
Hidden TAB's | DropDowns IndexLayout 10.1 MiSeq 10.2 MiniSeq 10.3 iSeq Workbook Verification
WORKBOOK TAB: Cover Pages
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells A2, B2 and C2 of each cover page and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
WORKBOOK TAB: DropDowns
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | HIDDEN | The entire tab should be hidden from users. |  | N/A |  | Verify that the tab is not visible when opening the workbook. Verify that the tab is present by right clicking on the 1.0 Sample_Prep tab and selecting 'unhide', the DropDowns tab should be avaliable to unhide. Unhide the tab to confirm locking below.
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells A2, B2 and F2 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
WORKBOOK TAB: IndexLayout
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | HIDDEN | The entire tab should be hidden from users. |  | N/A |  | Verify that the tab is not visible when opening the workbook. Verify that the tab is present by right clicking on the 1.0 Sample_Prep tab and selecting 'unhide', the IndexLayout tab should be available to unhide. Unhide the tab to confirm locking below.
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells A2, B2 and C2 of each cover page and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
V3-V8 | INDEX LIST | These cells are used to create an index list that can be applied to the reagents preparation table of the 2.0 Library_Prep. tab. |  | ="Indexes: "&B8&", "&C8&", "&D8&", "&E8&", "&F8&", "&G8&", "&H8&", "&I8&", "&J8&", "&K8&", "&L8&", "&M8&", "&N8&", "&O8&", "&P8&", "&Q8&", "&R8&", "&S8&", "&T8&", "&U8 |  | Performance of these cells will be assessed on the 2.0 Library_Prep tab.
WORKBOOK TAB: 0.1 Import
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells A2, B2 and C2 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Button | IMPORT MARCO | Macro has been created to import sample lists exported from an excel file. Clicking the button will import the information onto the Sample List column. |  | N/A |  | Follow the below intructions: - open new excel file and create a list in row A of the workbook with no header, save to desktop, - go to 0.1 Import tab and click the 'Import Sample List' button (enable macro if required), select the file and click open. Verify that the samples are imported into the Sample List Column (column U) of 0.1 Import tab
Button | BUILD PLATE MAP MACRO | Macro has been created to build the plate map on 1.0 Sample_Prep using the sample list imported from IMPORT MACRO. Clicking the button will copy the sample list to appropriate wells of the plate map, based on kit type selected and starting well. |  | N/A |  | Follow the below instructions: -ensure that there are samples present in the Sample List Column (column U) of 0.1 Import tab. (Import Macro should be run first) -select the type of kit (24 or 96) -select the starting well Verify that the samples have been copied into 1.0 Sample_Prep into appropriate wells. Verify that the error box pops up if the sample list is longer than the number of wells available.
WORKBOOK TAB: 1.0 Sample_Prep
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL (except E3, E4, E5, E6, E8, E9, G23-R31, and F23-F31 and green cells) | LOCK | At the time of issue, the worksheet shall be locked to prevent editing. The password to be used is the standard laboratory use password for controlled documents and workbooks. All cells should be locked except those listed below. |  | N/A |  | Click on cells F7, F32, G22 and E24 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
E3, E4, E5, E6, E8, E9, G23-R31, and F23-F31 and green cells | UNLOCKED | The listed cells are to be unlocked to allow operators to record pertinent information. |  | N/A |  | Enter the values listed below into the cells as listed. Confirm that the cells are unlocked and the values can be entered. E3 = Enter date and initials (ie 20190730HH) E4 = Enter initials G28-L31 = Enter date and well number into each cell M24-R31 = Enter date and well number into each cell
E5 | PROBE SET DROPDOWN | This cell is used to allow users to select between a defined number of entry options. |  | Dropdown Option |  | Verify that the cell allows user to select from: - Select - AlloSeqTx17.1 - AlloSeqTx9.1 Select AlloSeqTx17.1
E6 | SEQUENCER DROPDOWN | This cell is used to allow users to select between a defined number of entry options. |  | Dropdown Option |  | Verify that the cell allows user to select from: - Select - MiSeq - MiniSeq - iSeq Select MiSeq.
E6 | SEQUENCER MACRO | The sequencer drop down is used to control which sequencing instruction is present in the workbook. For example, if MiSeq is selected then the 10.1 MiSeq tab will appear. |  | Dropdown Option |  | Verify that the cell allows user to select from: - Select - MiSeq - MiniSeq - iSeq Select MiSeq.
E7 | # SAMPLES | This cell uses the plate layout in cells G24-R31 to calculate the number of samples to be processed on the run. This Cell should count any samples entered into the plate layout in either text or number format. Macro command: Private Sub Worksheet_Change(ByVal Target As Range) If [E6] = "MiSeq" Then Sheets("10.1 MiSeq").Visible = True Else Sheets("10.1 MiSeq").Visible = False End If If [E6] = "MiniSeq" Then Sheets("10.2 MiniSeq").Visible = True Else Sheets("10.2 MiniSeq").Visible = False End If If [E6] = "iSeq" Then Sheets("10.3 iSeq").Visible = True Else Sheets("10.3 iSeq").Visible = False End If End Sub |  | MACRO |  | Verify that the corresponding sequencing tab appears when selected from the dropdown.
E23 | INDEX SET DROPDOWN | This cell is used to allow users to select between a defined number of entry options. |  | Dropdown Option |  | Verify that the cell allows user to select from: - Select - 24 Set A Indexes - 24 Set B Indexes - 96 Set A Indexes - 96 Set B Indexes Select 96 Set A Indexes.
G23-R23 | INDEX DROPDOWN | This cell is used to allow users to select between a defined number of entry options. |  | Dropdown Option |  | Verify that the cell allows user to select from: - H705 - H701 - H716 - H723 - H706 - H702 - H718 - H724 - H707 - H703 - H719 - H726 - H710 - H704 - H720 - H727 - H711 - H712 - H721 - H728 - H714 - H715 - H722 - H729 Confirm that the indexes can be manually overridden, and then return to default.
F24-F31 | INDEX DROPDOWN | This cell is used to allow users to select between a defined number of entry options. |  | Dropdown Option |  | Verify that the cell allows user to select from: - H503 - H510 - H505 - H511 - H506 - H513 - H517 - H522 - H502 - H515 - H507 - H516 - H508 - H518 - H521 - H520 Confirm that the indexes can be manually overridden, and then return to default.
WORKBOOK TAB: 1.1 Linear View
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells A2, B2 and F2 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
D8-D103 | SAMPLE ID | These cells create the sample number in the following format for each sample entered into the plate layout: SampleID-HLA-ExperimentID# |  | =IF('1.0 Sample_Prep'!G24="","",('1.0 Sample_Prep'!G24&"-HLA-"&'1.0 Sample_Prep'!$E$3)) |  | Verify that the contents in D8 matches the values entered into G24 of 1.0 Sample_Prep, -HLA-, plus experiment number entered into E3of 1.0 Sample_Prep. Verify that the sample ID is generated and corresponds to expected format for all samples entered.
E8-E103 | SAMPLE NAME | When a sample is present on the plate layout, the sample name shall correspond to the sample ID. |  | =IF(D8="","",(D8)) |  | Verify that the sample name is generated and corresponds to the sample ID for all samples entered.
F8-F103 | SAMPLE PLATE | When a sample is present on the plate layout, the sample plate cell will correspond to the experiment number. |  | =IF(D8="","",'1.0 Sample_Prep'!$E$3) |  | Verify that the sample plate is generated and corresponds to the experiment ID number entered into E3 of 1.0 Sample_Prep.
G8-G103 | SAMPLE WELL | When a sample is present on the plate layout, the sample well will correspond to the Well ID in column C. |  | =IF(D8="","",(C8)) |  | Verify that the sample well is generated and corresponds to the expected well number from row C.
H8-H103 | I7 INDEX ID | When a sample is present on the plate layout, the I7 Index ID will correspond to the index selected from the i7 index dropdown for the sample column on the plate layout. |  | =IF('1.1 LinearView'!D8="","",('1.0 Sample_Prep'!$G$23)) |  | Verify that the indexes listed correspond to those selected in the plate layout for each sample.
I8-I103 | INDEX | When a sample is present on the plate layout, the index sequence will correspond to the sequence of the index selected from the i7 index dropdown for the sample column on the plate layout. |  | =IFERROR(VLOOKUP(H8,DropDowns!$C$2:$D$26,2,FALSE),"") |  | For all cells that correspond to wells beginning with 'A', verify that the sequence listed corresponds to the expected sequence for the index in column I, by consulting the table in rows C-D of the dropdowns tab.
J8-J103 | I5 INDEX ID | When a sample is present on the plate layout, the I5 Index ID will correspond to the index selected from the I5 index dropdown for the sample row on the plate layout. |  | =IF(D8="","",('1.0 Sample_Prep'!$F$24)) |  | Verify that the indexes listed correspond to those selected in the plate layout for each sample.
K8-K103 | INDEX2 | When a sample is present on the plate layout, the index sequence will correspond to the sequence of the index selected from the i5 index dropdown for the sample row on the plate layout. |  | =IFERROR(VLOOKUP(J8,DropDowns!$E$2:$F$18,2,FALSE),"") |  | For cells corresponding to row 1, verify that the sequence listed corresponds to the expected sequence for the index in column K, by consulting the table in rows E-F of the dropdowns tab.
L8-L103 | SAMPLE PROJECT | When a sample is present on the plate layout, the sample project will correspond to the experiment ID entered into E3 of 1.0 Sample_Prep. |  | =IF(D8="","",('1.0 Sample_Prep'!$E$3)) |  | Verify that the cell shows the experiment ID entered into cell E3 of the 1.0 Sample_Prep tab, for each row containing sample.
WORKBOOK TAB: 1.2 SampleSheet
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells B3, A22 and J22 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
B3 | INVESTIGATOR NAME | Information copied from the 1.0 Sample_Prep tab to populate the sample sheet. |  | ='1.0 Sample_Prep'!E8 |  | Verify that the contents of the cell matches the values entered into E8 of 1.0 Sample_Prep tab. If cell includes "Enter" or "Select" the text wil appear in red.
B4 | EXPERIMENT ID NAME | Information copied from the 1.0 Sample_Prep tab to populate the sample sheet. |  | ='1.0 Sample_Prep'!E3 |  | Verify that the contents of the cell matches the values entered into E3 of 1.0 Sample_Prep tab. If cell includes "Enter" or "Select" the text wil appear in red.
B5 | DATE | Information copied from the 1.0 Sample_Prep tab to populate the sample sheet. |  | ='1.0 Sample_Prep'!E9 |  | Verify that the contents of the cell matches the values entered into E9 of 1.0 Sample_Prep tab. If cell includes "Enter" or "Select" the text wil appear in red.
B8 | INSTRUMENT TYPE | Information copied from the 1.0 Sample_Prep tab to populate the sample sheet. |  | ='1.0 Sample_Prep'!E6 |  | Verify that the contents of the cell matches the values entered into E6 of 1.0 Sample_Prep tab. If cell includes "Enter" or "Select" the text wil appear in red.
B9 | ASSAY | Information copied from the 1.0 Sample_Prep tab to populate the sample sheet. |  | ='1.0 Sample_Prep'!E5 |  | Verify that the contents of the cell matches the values entered into E5 of 1.0 Sample_Prep tab. If cell includes "Enter" or "Select" the text wil appear in red.
B10 | DESCRIPTION | Information copied from the 1.0 Sample_Prep tab to populate the sample sheet. |  | ='1.0 Sample_Prep'!E4 |  | Verify that the contents of the cell matches the values entered into E4 of 1.0 Sample_Prep tab. If cell includes "Enter" or "Select" the text wil appear in red.
A22-J117 | SAMPLE SHEET DATA | Information copied from the 1.1 LinearView tab to populate the sample sheet. |  | =IF('1.1 LinearView'!D8=" ","",('1.1 LinearView'!D8)) |  | Verify that the contents of the cells match the values in the sample list table on the 1.1 LinearView tab. If cell includes "Enter" or "Select" the text wil appear in red.
BUTTON | EXPORT MACRO | A macro has been added to export the sample sheet as a csv and remove the blank rows. |  | N/a |  | Verify that the correct .csv file is generated at the location selected during the macro run. Selecting no location will result in an error box.
WORKBOOK TAB: 2.0 Library_Prep
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, K11 and H33 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!E3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
K11 | # OF SAMPLES REQUIRED | Total number of samples to be processed, as calculated on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!E7 |  | Verify that the contents in K11 matches the values entered into E7 of 1.0 Sample_Prep tab.
J12 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I12 |  | Verify that the cell shows 960
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I13 |  | Verify that the cell shows 2,880
J14 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I14 |  | Verify that the cell shows 960
J15 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I15 |  | Verify that the cell shows 960
J16 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I16 |  | Verify that the cell shows 28,800
J17 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I17 |  | Verify that the cell shows 1,920
J18 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I18 |  | Verify that the cell shows 960
I33 | # OF SAMPLES REQUIRED | Total number of samples to be processed, as calculated on the Sample_Prep tab. |  | =K11 |  | Verify that the contents in I33 matches the values entered into E7 of 1.0 Sample_Prep tab.
H34 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 10% extra. |  | =$I$33*G34*1.1 |  | Verify that the cell shows 1,056
H35 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 10% extra. |  | =$I$33*G35*1.1 |  | Verify that the cell shows 2,112
H36 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 10% extra. |  | =$I$33*G36*1.1 |  | Verify that the cell shows 1,056
G37 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(G34:G36) |  | Verify that the cell shows 40
H37 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H34:I36) |  | Verify that the cell shows 4,224
I77 | # OF SAMPLES REQUIRED | Total number of samples to be processed, as calculated on the Sample_Prep tab. |  | =K11 |  | Verify that the contents in I77 matches the values entered into E7 of 1.0 Sample_Prep tab.
H78 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 10% extra. |  | =$I$77*G78*1.1 |  | Verify that the cell shows 2,112
H79 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 10% extra. |  | =$I$77*G79*1.1 |  | Verify that the cell shows 2,112
G80 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(G78:G79) |  | Verify that the cell shows 40
H80 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H78:I79) |  | Verify that the cell shows 4,224
WORKBOOK TAB: 3.0 SS_Purification
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells, K13, K14, and I35 | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, K22 and I29 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G23 to confirm that green cells are unlocked.
K13, K14, I35 | UNLOCKED | The cell shall be unlocked to allow operators to manually override the automated calculations. |  | N/A |  | Verify that the default contents of the cells is: K13 = 96 K14 = 2.5 I35 = 1 Verify the cells are unlocked and can be edited.
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!E3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
K13 | # OF SAMPLES | This cell displays the number of samples to be pooled, based on the plate layout on 1.0 Sample_Prep tab. This value can be manually overridden if users wish to vary the number of samples. |  | ='1.0 Sample_Prep'!E7 |  | Verify that the cell shows 96
K14 | VOLUME PER SAMPLE | This cell calculates the volume to pool per sample depending on the number of samples to be pooled and which range this falls within. This value can be manually overridden if users wish to pool higher volume for poor quality samples. |  | =IF(AND(K13>=6,K13<=24),H14,IF(AND(K13>=25,K13<=48),I14,IF(AND(K13>=49,K13<=96),J14,0))) |  | Verify that the cell shows 2.5
K15 | VOLUME OF DILUTED BEADS | This cell calculates the amount of diluted beads that is required, based on the number of samples and volume being pooled. |  | =K14*5 |  | Verify that the cell shows 12.5
K16 | TRANSFER VOLUME | This cell calculates the transfer volume that is required, based on the number of samples and volume being pooled. |  | =K14*5.5 |  | Verify that the cell shows 13.75
K17 | VOLUME OF NEAT BEADS | This cell calculates the amount of neat beads that is required, based on the number of samples and volume being pooled. |  | =K14*0.44 |  | Verify that the cell shows 1.1
K22 | # OF SAMPLES REQUIRED | Total number of samples to be processed, copied from K13 originally calculated on the Sample_Prep tab, but can be manually adjusted. |  | =K13 |  | Verify that the contents in K22 matches the values entered into K13.
J23 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of samples (K13) by the volume per reagent needed neat (K17), and adding to the volume required for diluted beads (H30). |  | =H30+(K17*K13) |  | Verify that the cell shows 681.6
J24 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by combining the values in the bead dilution and 80% ethanol preparation tables. |  | =H31+H37 |  | Verify that the cell shows 1,584
J25 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, copied from the 80% ethanol preparation table. |  | =H36 |  | Verify that the cell shows 2,880
J26 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent. |  | =I26*I35 |  | Verify that the cell shows 37
I29 | # OF SAMPLES REQUIRED | Total number of samples to be processed, copied from K13 originally calculated on the Sample_Prep tab, but can be manually adjusted. |  | =K13 |  | Verify that the contents in I29 matches the values entered into K13.
G30 | TOTAL VOLUME OF REAGENT REQUIRED | The volume of reagent required to dilute the purification beads according to the size selection calculations table and number of samples on the run. |  | =($K$15/5)*2 |  | Verify that the cell shows 5
G31 | TOTAL VOLUME OF REAGENT REQUIRED | The volume of reagent required to dilute the purification beads according to the size selection calculations table and number of samples on the run. |  | =($K$15/5)*3 |  | Verify that the cell shows 7.5
H30 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 20% extra. |  | =$I$29*G30*1.2 |  | Verify that the cell shows 576
H31 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 20% extra. |  | =$I$29*G31*1.2 |  | Verify that the cell shows 864
G32 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(G30:G31) |  | Verify that the cell shows 12.5
H32 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H30:I31) |  | Verify that the cell shows 1,440
H36 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 50% extra. |  | =$I$35*G36*1.5 |  | Verify that the cell shows 2,880
H37 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 50% extra. |  | =$I$35*G37*1.5 |  | Verify that the cell shows 720
G38 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(G36:G37) |  | Verify that the cell shows 2,400
H38 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H36:I37) |  | Verify that the cell shows 3,600
D49 | INSTRUCTION | This cell contians formula to incorportate the volume required into the instructions depending on the number of samples on the run. |  | ="Aliquot "&(G32*I29)&" µL of diluted Purification Beads into the 1.5mL tube." |  | Verify that the cell shows 1,200
D50 | INSTRUCTION | This cell contians formula to incorportate the volume required into the instructions depending on the number of samples on the run. |  | ="Pipette mix the tagmentation beads and supernatant from indexing PCR, and then add "&K14&" µL of each sample to tube containing diluted Purification Beads." |  | Verify that the cell shows 2.5
D55 | INSTRUCTION | This cell contians formula to incorportate the volume required into the instructions depending on the number of samples on the run. |  | ="Transfer "&(K13*K16)&" µL of the supernatant to a new tube." |  | Verify that the cell shows 1,320
D56 | INSTRUCTION | This cell contians formula to incorportate the volume required into the instructions depending on the number of samples on the run. |  | ="Add "&(K13*K17)&" µL of Purification Beads (undiluted) to the tube containing the supernatant." |  | Verify that the cell shows 105.6
WORKBOOK TAB: 4.0 Hybridisation
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, K11 and G27 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the 1.0 Sample_Prep tab. |  | ='1.0 Sample_Prep'!$E$3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
K11 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab. |  | ='3.0 SS_Purification'!I35 |  | Verify that the contents in K11 matches the values entered into I35 of 3.0 SS_Purification tab.
J12 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I12 |  | Verify that the cell shows 10
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I13 |  | Verify that the cell shows 50
J14 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I14 |  | Verify that the cell shows 10
G31 | TOTAL VOLUME PER POOL | The sum of the volumes to be added together. |  | =SUM(G27:G30) |  | Verify that the cell shows 100
WORKBOOK TAB: 5.0 Capture
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, K11 and I19 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!$E$3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
K11 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | ='4.0 Hybridisation'!K11 |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab.
J12 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =K11*I12 |  | Verify that the cell shows 250
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent, plus 100uL for excess. |  | =(K11*I13)+(K11*100) |  | Verify that the cell shows 900
J14 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =K11*I14 |  | Verify that the cell shows 28.5
J15 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =K11*I15 |  | Verify that the cell shows 1.5
J16 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =K11*I16 |  | Verify that the cell shows 4
I19 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | =K11 |  | Verify that the contents in I19 matches the values entered into K11 of 4.0 Hybridisation tab.
H20 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =I19*G20 |  | Verify that the cell shows 28.5
H21 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =I19*G21 |  | Verify that the cell shows 1.5
WORKBOOK TAB: 6.0 Enrich_PCR
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, K11 and H19 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!$E$3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
K11 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | ='4.0 Hybridisation'!K11 |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab.
J12 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I12 |  | Verify that the cell shows 5
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I13 |  | Verify that the cell shows 20
WORKBOOK TAB: 7.0 Purification
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, K11 and H19 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!$E$3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
K11 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | ='4.0 Hybridisation'!K11 |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab.
J12 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I12 |  | Verify that the cell shows 27
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, refer to calculations for H20 below. |  | =H20 |  | Verify that the cell shows 120
J14 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, refer to calculations for H19 below. |  | =H19 |  | Verify that the cell shows 480
J15 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools by the volume per reagent. |  | =$K$11*I15 |  | Verify that the cell shows 32
I18 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | ='4.0 Hybridisation'!K11 |  | Verify that the contents in I18 matches the values entered into K11 of 4.0 Hybridisation tab.
H19 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 50% extra. |  | =$I$18*G19*1.5 |  | Verify that the cell shows 480
H20 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of pools by the volume per reagent, plus 50% extra. |  | =$I$18*G20*1.5 |  | Verify that the cell shows 120
G21 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(G19:G20) |  | Verify that the cell shows 400
H21 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H19:I20) |  | Verify that the cell shows 600
WORKBOOK TAB: 8.0 QuBit
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells and K11 | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, L11 and H19 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
K11 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | ='4.0 Hybridisation'!K11 |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab.
K11 | UNLOCKED | The cell shall be unlocked to allow operators to manually override the automated calculations. |  | N/A |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab. Modify such that K11 = 2
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!$E$3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
J12 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools plus 3, by the volume per reagent. |  | =(($K$11+3)*I12) |  | Verify that the cell shows 995
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools plus 3, by the volume per reagent. |  | =(($K$11+3)*I13) |  | Verify that the cell shows 5
G27 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools plus 3, by the volume per reagent. |  | =J12 |  | Verify that the cell shows 995
G28 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools plus 3, by the volume per reagent. |  | =J13 |  | Verify that the cell shows 5
G29 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =G28+G27 |  | Verify that the cell shows 1000
J43-J50 | AVERAGE YIELD | These cells are used to calculate the average yield, based on values entered into column H and I. |  | =IFERROR(AVERAGE(H##:I##),"") |  | Enter the following values: H43 = 14.8 I43 = 15.0 H44 = 10.4 I44 = 12.2 Verify that the cell J43 shows 14.9 Verify that the cell J44 shows 11.3
J43-J50 | AVERAGE YIELD | These cells are used to calculate the average yield, based on values entered into column H and I. If the user enters only one value, the calculation will use that value only in the calculation. |  | =IFERROR(AVERAGE(H##:I##),"") |  | Enter the following values: H42 = 14 I43 = 10.2 Verify that the cell J42 shows 14 Verify that the cell J43 shows 10.2
WORKBOOK TAB: 9.0 TapeStation(Optional)
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL except green cells and K11 | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells F3, L11 and H19 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
K11 | # OF POOLS REQUIRED | Total number of pools to be processed, as entered by operator on the 3.0 SS_Purification tab (replicated on 4.0 Hybridisation). |  | ='4.0 Hybridisation'!K11 |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab.
K11 | UNLOCKED | The cell shall be unlocked to allow operators to manually override the automated calculations. |  | N/A |  | Verify that the contents in K11 matches the values entered into K11 of 4.0 Hybridisation tab. Modify such that K11 = 2
F3 | EXPERIMENT ID # | The unique identifier given to the experiment, as entered by the operator on the Sample_Prep tab. |  | ='1.0 Sample_Prep'!$E$3 |  | Verify that the contents in F3 matches the values entered into E3 of 1.0 Sample_Prep tab.
J13 | TOTAL VOLUME OF REAGENT REQUIRED | The total volume of reagent required, calculated by multiplying the total number of tests/pools plus 1, by the volume per reagent. |  | =($K$11+1)*I13 |  | Verify that the cell shows 9
WORKBOOK TAB: 10.0 PhiX
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL (except H45, H63 and green cells) | LOCK | At the time of issue, the workbook shall be locked to prevent editing. The password to be used is the standard laboratory use password for controlled documents and workbooks. All cells should be locked except those listed below. |  | N/A |  | Click on cells F2, I15 and H48 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
Green cells | UNLOCKED | Green cells should remain unlocked for users who wish to enter values electroincally. |  | N/A |  | Enter values in cells F4 and G12 to confirm that green cells are unlocked.
H45, H63 | UNLOCKED | The listed cells are to be unlocked to allow operators to record pertinent information. |  | N/A |  | Enter the values listed below into the cells as listed. Confirm that the cells are unlocked and the values can be entered. H45 = 100 H63= 100
I15 | VOLUME PER DILUTION | The total volume of HT1 reagent required, calculated by adding 990uL used at step D34 and the value in H50. |  | =990+H50 |  | Verify that the cell shows 1065
H49 | VOLUME REQUIRED | The volume of PhiX required,to make the desired volume (in H45) at 5pM, from the 20pM vials. Equivalent to 1 in 4 dilution. |  | =(5*H45)/20 |  | Verify that the cell shows 25
H50 | VOLUME REQUIRED | The volume of HT1 required, calculated by subtracting the PhiX volume in H49 from the total required volume in H45. |  | =H45-H49 |  | Verify that the cell shows 75
H51 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H49:H50) |  | Verify that the cell shows 100. Verify that value matches that entered into H45.
H67 | VOLUME REQUIRED | The volume of PhiX required,to make the desired volume (in H63) at 20pM, from the 10nM vials. |  | =(0.02*H63)/10 |  | Verify that the cell shows 0.2
H68 | VOLUME REQUIRED | The volume of Resuspension Buffer required, calculated by subtracting the PhiX volume in H67 from the total required volume in H63. |  | =H63-H67 |  | Verify that the cell shows 99.8
H69 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H67:H68) |  | Verify that the cell shows 100. Verify that value matches that entered into H63.
WORKBOOK TAB: 10.1 MiSeq
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL (except F3, H33, H34, H36, H63, H64, H65, H66 and all green cells) | LOCK | At the time of issue, the workbook shall be locked to prevent editing. The password to be used is the standard laboratory use password for controlled documents and workbooks. All cells should be locked except those listed below. |  | N/A |  | Click on cells B3, I12 and H39 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
F3, H33, H34, H36, H63, H64, H65, H66 and all green cells | UNLOCKED | The listed cells are to be unlocked to allow operators to record pertinent information. |  | N/A |  | Enter the values listed below into the cells as listed. Confirm that the cells are unlocked and the values can be entered. F3 = Date H33 = 12 H34 = 800 H36 = 30 H63 = 1 (enter 2, 3, and 4 as described below) H64 = 12 H65 = 1% H66 = 600 Default values will be entered for cells H34, H36, H63, H64, H65 and H66 for ease of customer use.
I12 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H43 |  | Verify that the cell shows 5.3
I13 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =G22 |  | Verify that the cell shows 5
I14 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =G21 |  | Verify that the cell shows 45
I15 | VOLUME REQUIRED | The total volume of HT1 reagent required, calculated by adding 990uL used at step D33 and the value in H77. |  | =990+H77 |  | Verify that the cell shows 1226.4
I16 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H42 |  | Verify that the cell shows 24.7
I17 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H76 |  | Verify that the cell shows 3.6
H35 | nM CONC. | This cell converts the pool concentration from ng/uL to nM. |  | =((H33)/(660*H34))*10^6 |  | Verify that the cell shows 22.73
H42 | VOLUME REQUIRED | The volume of Resuspension Buffer required, calculated by subtracting the pool volume in H43 from the total required volume in H36. |  | =H36-H43 |  | Verify that the cell shows 24.7
H43 | VOLUME REQUIRED | The volume of pool required ,to make the desired volume (in H36) at 4nM, from the concentration calculated in cell H35. |  | =(4*H36)/H35 |  | Verify that the cell shows 5.3
H44 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H42:H43) |  | Verify that the cell shows 30.0. Verify that value matches that entered into H36.
H72 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H63), loading concentration (H64), and the desired total volume (H66). |  | =((H64*H66)/20)/H63 |  | Verify that when value on left is entered into H63, the value and colour of cell matches those on the right: 1 = 360.0, white 2 = 180.0, white 3 = 120.0, white 4 = 90.0, white
H73 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H63), loading concentration (H64), and the desired total volume (H66). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$63=1,0,IF($H$63>1,$H$72,"")) |  | Verify that when value on left is entered into H63, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 180.0, white 3 = 120.0, white 4 = 90.0, white
H74 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H63), loading concentration (H64), and the desired total volume (H66). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$63=2,0,IF($H$63>2,$H$72,0)) |  | Verify that when value on left is entered into H63, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 0.0, grey 3 = 120.0, white 4 = 90.0, white
H75 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H63), loading concentration (H64), and the desired total volume (H66). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$63=3,0,IF($H$63>3,$H$72,0)) |  | Verify that when value on left is entered into H63, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 0.0, grey 3 =0.0, grey 4 = 90.0, white
H76 | VOLUME REQUIRED | The volume of phiX required, based on the loading concentration (H64), spike in concentration (H65) and the desired total volume (H66). |  | =((H64*H66)/20)*H65 |  | Verify that the cell shows 3.6
H77 | VOLUME REQUIRED | The volume of HT1 required to reach the desired total volume (H66), minus the volume of reagents in H72-H76. |  | =H66-(H72+H73+H74+H75+H76) |  | Verify that the cell shows 236.4
H78 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H72:H77) |  | Verify that the cell shows 600. Verify that value matches that entered into H66.
WORKBOOK TAB: 10.2 MiniSeq
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL (except F3, H34, H35, H37, H66, H67, H68 and H69 and all green cells) | LOCK | At the time of issue, the workbook shall be locked to prevent editing. The password to be used is the standard laboratory use password for controlled documents and workbooks. All cells should be locked except those listed below. |  | N/A |  | Click on cells B3, I12 and H39 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
F3, H34, H35, H37, H66, H67, H68 and H69 and all green cells | UNLOCKED | The listed cells are to be unlocked to allow operators to record pertinent information. |  | N/A |  | Enter the values listed below into the cells as listed. Confirm that the cells are unlocked and the values can be entered. F3 = Date H34 = 12 H35 = 800 H37 = 100 H66 = 1 (enter 2, 3, and 4 as described below) H67 = 1.6 H68 = 1% H69 = 500 Default values will be entered for cells H35, H37, H66, H67, H68 and H69 for ease of customer use.
I12 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H44 |  | Verify that the cell shows 4.4
I13 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =G23 |  | Verify that the cell shows 5
I14 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =G22 |  | Verify that the cell shows 95
I15 | VOLUME REQUIRED | The total volume of HT1 reagent required, calculated by adding 985uL used at step D57 and the value in H80. |  | =985+H80 |  | Verify that the cell shows 1323.4
I16 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H43 |  | Verify that the cell shows 95.6
I18 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H79 |  | Verify that the cell shows 1.6
H36 | nM CONC. | This cell converts the pool concentration from ng/uL to nM. |  | =((H34)/(660*H35))*10^6 |  | Verify that the cell shows 22.73
H43 | VOLUME REQUIRED | The volume of Resuspension Buffer required, calculated by subtracting the pool volume in H44 from the total required volume in H37. |  | =H37-H44 |  | Verify that the cell shows 95.6
H44 | VOLUME REQUIRED | The volume of pool required ,to make the desired volume (in H37) at 1nM, from the concentration calculated in cell H36. |  | =(1*H37)/H36 |  | Verify that the cell shows 4.4
H45 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H43:H44) |  | Verify that the cell shows 100.0. Verify that value matches that entered into H37.
H75 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H66), loading concentration (H67), and the desired total volume (H69). |  | =((H67*H69)/5)/H66 |  | Verify that when value on left is entered into H66, the value and colour of cell matches those on the right: 1 = 160.0, white 2 = 80.0, white 3 = 53.3, white 4 = 40.0, white
H76 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H66), loading concentration (H67), and the desired total volume (H69). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$66=1,0,IF($H$66>1,$H$75,"")) |  | Verify that when value on left is entered into H66, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 80.0, white 3 = 53.3, white 4 = 40.0, white
H77 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H66), loading concentration (H67), and the desired total volume (H69). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$66=2,0,IF($H$66>2,$H$75,0)) |  | Verify that when value on left is entered into H66, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 0.0, grey 3 = 53.3, white 4 = 40.0, white
H78 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H66), loading concentration (H67), and the desired total volume (H69). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$66=3,0,IF($H$66>3,$H$75,0)) |  | Verify that when value on left is entered into H66, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 0.0, grey 3 = 0.0, grey 4 = 40.0, white
H79 | VOLUME REQUIRED | The volume of phiX required, based on the loading concentration (H67), spike in concentration (H68) and the desired total volume (H69). |  | =((H67*H69)/5)*H68 |  | Verify that the cell shows 1.6
H80 | VOLUME REQUIRED | The volume of HT1 required to reach the desired total volume (H69), minus the volume of reagents in H75-H79. |  | =H69-(H75+H76+H77+H78+H79) |  | Verify that the cell shows 338.4
H81 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H75:H80) |  | Verify that the cell shows 500. Verify that value matches that entered into H69.
WORKBOOK TAB: 10.3 iSeq
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL (except F3, H24, H25, H27, H43, H44, H45 and H46 and all green cells) | LOCK | At the time of issue, the workbook shall be locked to prevent editing. The password to be used is the standard laboratory use password for controlled documents and workbooks. All cells should be locked except those listed below. |  | N/A |  | Click on cells B3, H26 and H48 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
F3, H24, H25, H27, H43, H44, H45 and H46 and all green cells | UNLOCKED | The listed cells are to be unlocked to allow operators to record pertinent information. |  | N/A |  | Enter the values listed below into the cells as listed. Confirm that the cells are unlocked and the values can be entered. F3 = Date H24 = 12 H25 = 800 H27 = 100 H43 = 1 (enter 2, 3, and 4 as described below) H44 = 200 H45 = 1% H46 = 100 Default values will be entered for cells H25, H27, H43, H44, H45 and H46 for ease of customer use.
I12 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H34 |  | Verify that the cell shows 4.4
I13 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H33+H57 |  | Verify that the cell shows 165.6
I14 | VOLUME REQUIRED | The volume of reagent required, referencing calculation cell in tab. |  | =H56 |  | Verify that the cell shows 10.0
H26 | nM CONC. | This cell converts the pool concentration from ng/uL to nM. |  | =((H24)/(660*H25))*10^6 |  | Verify that the cell shows 22.73
H33 | VOLUME REQUIRED | The volume of Resuspension Buffer required, calculated by subtracting the pool volume in H34 from the total required volume in H27. |  | =H27-H34 |  | Verify that the cell shows 95.6
H34 | VOLUME REQUIRED | The volume of pool required ,to make the desired volume (in H27) at 1nM, from the concentration calculated in cell H26. |  | =(1*H27)/H26 |  | Verify that the cell shows 4.4
H35 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H33:H34) |  | Verify that the cell shows 100.0. Verify that value matches that entered into H27.
H52 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H43), loading concentration (H44), and the desired total volume (H46). |  | =((H44*H46)/1000)/H43 |  | Verify that when value on left is entered into H43, the value and colour of cell matches those on the right: 1 = 20.0, white 2 = 10.0, white 3 = 6.7, white 4 = 5.0, white
H53 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H43), loading concentration (H44), and the desired total volume (H46). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$43=1,0,IF($H$43>1,$H$52,"")) |  | Verify that when value on left is entered into H43, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 10.0, white 3 = 6.7, white 4 = 5.0, white
H54 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H43), loading concentration (H44), and the desired total volume (H46). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$43=2,0,IF($H$43>2,$H$52,0)) |  | Verify that when value on left is entered into H43, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 0.0, grey 3 = 6.7, white 4 = 5.0, white
H55 | VOLUME REQUIRED | The volume of pool required, based on the number of pools to be loaded together on same sequencing run (H43), loading concentration (H44), and the desired total volume (H46). If the cell is not required (ie additonal pools not to be added) the cell will be greyed out and show a volume of zero. |  | =IF($H$43=3,0,IF($H$43>3,$H$52,0)) |  | Verify that when value on left is entered into H43, the value and colour of cell matches those on the right: 1 = 0.0, grey 2 = 0.0, grey 3 = 0.0, grey 4 = 5.0, white
H56 | VOLUME REQUIRED | The volume of phiX required, based on the loading concentration (H44), spike in concentration (H45) and the desired total volume (H46). |  | =((H44*H46)/20)*H45 |  | Verify that the cell shows 10.0
H57 | VOLUME REQUIRED | The volume of HT1 required to reach the desired total volume (H46), minus the volume of reagents in H52-H56. |  | =H46-(H52+H53+H54+H55+H56) |  | Verify that the cell shows 70.0
H58 | TOTAL VOLUME | The sum of the volumes to be added together. |  | =SUM(H52:H57) |  | Verify that the cell shows 100. Verify that value matches that entered into H46.
WORKBOOK TAB: VersionHistory
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells A2, B2 and C2 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
WORKBOOK TAB: WorkbookVerification
CELLS FORMATTED | TITLE | INTENDED FUNCTIONALITY |  | FORMULA |  | VERIFICATION METHOD | PASS (Y/N)
ALL | LOCK | Worksheet shall be locked to prevent editing. The password is the standard laboratory use password for controlled documents and workbooks. |  | N/A |  | Click on cells C2, A5 and H21 and confirm that the following warning message appears: "The cell or chart you're trying to change is on a protected sheet. To make a change, unprotect the sheet. You might be requested to enter a password."
WORKBOOK VERIFICATION ACCEPTANCE AND RELEASE
ACCEPTANCE AND RELEASE CRITERIA |  |  |  | Y/N | REVIEW
All required fields of verification report are complete. |  |  |  |  | AUTHORISATION FOR RELEASE (INITIALS) |  | DATE (DDMMMYY)
All listed functionality passes verification testing.
Proceed to issue the workbook.