## CE Cover
AlloSeq Tx
Assay Workbook
IFU095-5
Version Number: 2.0
Issue Date: Apr 2022
ASTX17.1(24)-IVD
ASTX17.1(24)-B-IVD
ASTX17.1(96)-A-IVD
ASTX17.1(96)-B-IVD
| CareDx Pty Ltd |  | Qarad BV,
| 20 Collie St. |  | Cipalstraat 3
| Fremantle WA 6160 |  | 2440 Geel
| Australia |  | Belgium
© 2022 CareDx, Inc. All service marks or trademarks are owned or licensed by CareDx, Inc. or its affiliates. All rights reserved.

## DropDowns (hidden)
Probe_Set | Sequencer | i7_IndexNames | i7_IndexSequences | i5_IndexNames | i5_IndexSequences
Select | Select
AlloSeqTx17.1 | MiSeq | H705 | GGACTCCT | H503 | TATCCTCT
| MiniSeq | H706 | TAGGCATG | H505 | GTAAGGAG
| iSeq | H707 | CTCTCTAC | H506 | ACTGCATA
|  | H710 | CGAGGCTG | H517 | GCGTAAGA
|  | H711 | AAGAGGCA | H502 | CTCTCTAT
|  | H714 | GCTCATGA | H507 | AAGGAGTA
|  | H701 | TAAGGCGA | H508 | CTAAGCCT
AlloSeqTx17.2 |  | H702 | CGTACTAG | H521 | GAGCCTTA
AlloSeqTx8.1 |  | H703 | AGGCAGAA | H510 | CGTCTAAT
AlloSeqTxKIR |  | H704 | TCCTGAGC | H511 | TCTCTCCG
|  | H712 | GTAGAGGA | H513 | TCGACTAG
|  | H715 | ATCTCAGG | H522 | TTATGCGA
|  | H716 | ACTCGCTA | H515 | TTCTAGCT
|  | H718 | GGAGCTAC | H516 | CCTAGAGT
|  | H719 | GCGTAGTA | H518 | CTATTAAG
|  | H720 | CGGAGCCT | H520 | AAGGCTAT
|  | H721 | TACGCTGC
|  | H722 | ATGCGCAG
|  | H723 | TAGCGCTC
|  | H724 | ACTGAGCG
|  | H726 | CCTAAGAC
|  | H727 | CGATCAGT
|  | H728 | TGCAGCTA
|  | H729 | TCGACGTC

## IndexLayout (hidden)
Product Code | A | B | C | D | E | F | G | H | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 2.0 Library_Prep Lookup
Select | [20 empty] | Select index set on 1.0 Sample_Prep tab.
24 Set A Indexes | H503 | H505 | H506 | H517 |  |  |  |  | H705 | H706 | H707 | H710 | H711 | H714 | [6 empty] | Indexes: H503, H505, H506, H517, H705, H706, H707, H710, H711, H714
24 Set B Indexes | H502 | H507 | H508 | H521 |  |  |  |  | H701 | H702 | H703 | H704 | H712 | H715 | [6 empty] | Indexes: H502, H507, H508, H521, H701, H702, H703, H704, H712, H715
96 Set A Indexes | H503 | H505 | H506 | H517 | H502 | H507 | H508 | H521 | H705 | H706 | H707 | H710 | H711 | H714 | H701 | H702 | H703 | H704 | H712 | H715 | Indexes: H503, H505, H506, H517, H502, H507, H508, H521, H705, H706, H707, H710, H711, H714, H701, H702, H703, H704, H712, H715
96 Set B Indexes | H510 | H511 | H513 | H522 | H515 | H516 | H518 | H520 | H716 | H718 | H719 | H720 | H721 | H722 | H723 | H724 | H726 | H727 | H728 | H729 | Indexes: H510, H511, H513, H522, H515, H516, H518, H520, H716, H718, H719, H720, H721, H722, H723, H724, H726, H727, H728, H729
| H503 | H505 | H506 | H517 | H502 | H507 | H508 | H521 | H716 | H718 | H719 | H720 | H721 | H722 | H723 | H724 | H726 | H727 | H728 | H729 | Indexes: H503, H505, H506, H517, H502, H507, H508, H521, H716, H718, H719, H720, H721, H722, H723, H724, H726, H727, H728, H729
| H510 | H511 | H513 | H522 | H515 | H516 | H518 | H520 | H705 | H706 | H707 | H710 | H711 | H714 | H701 | H702 | H703 | H704 | H712 | H715 | Indexes: H510, H511, H513, H522, H515, H516, H518, H520, H705, H706, H707, H710, H711, H714, H701, H702, H703, H704, H712, H715
96 Set C Indexes
96 Set D Indexes

## 0.1 HT_Import (hidden)
IMPORT SAMPLE LIST FROM HISTOTRAC
The below macro can be used to import sample lists, of 24 samples or less, from a HistoTrac exported list.
NOTE: The sample list generated from HistoTrac must be down the first row (Column A) beginning with cell A1. Do not include a column header.
| 1 | Click the "Enable Macros" button in the yellow ribbon above if necessary.
| 2 | Generate the HistoTrac sample list and save to a local or network location.
| 3 | Click the "Import HistoTrac Sample List" button to the right.
| 4 | Locate the list and click Open.
| 5 | Review the list in the "ImportData" sheet, and make any modifications as desired.
| 6 | Click on the "1.0 Sample_Prep" tab, and review the plate layout.
END OF SAMPLE IMPORT. PROCEED WITH PROTOCOL.
Import HistoTrac Sample List

## 1.0 Sample_Prep
ALLOSEQ TX - SAMPLE PREPARATION
| Experiment ID #: |  |  | Enter |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Description: |  |  | Enter |  |  |  | NOTE: ORANGE fields must be selected from drop down before printing.
| Probe Set: |  |  | Select |  |  |  | NOTE: YELLOW fields must be completed before printing.
| Sequencer: |  |  | Select
| # Samples: |  |  | 0
| Operator: |  |  | Enter
| Date: |  |  | Enter
EXPERIMENT SAMPLE LAYOUT
|  | 1 | Enter an experiment ID, description, operator and date into the yellow fields above.
|  | 2 | Select the AlloSeq probe set to be used for the experiment from the drop down above.
|  | 3 | Select the sequencer type to be used from the drop down above. Once the sequencer type has been selected, the sequencing instructions will be displayed in the workbook.
|  | 4 | Enter the ID of samples to be tested in the yellow section of the plate layout below, according to the desired configuration. NOTE: Only alphanumeric characters can be entered. NOTE: Duplicate sample IDs will be flagged in red to prompt user to correct. It is a requirement of the sequencer software to have only unique sample IDs on a run. NOTE: Do not enter any information for wells which are not intended to contain any sample.
|  | 5 | Select the desired index set from the orange dropdown under Plate Layout.
|  | 6 | If required the index order can be modified, select the alternative i7 index to be used from the drop down options for the required column. NOTE: If duplicate i7 indexes are selected then the cells will be flagged in red to prompt the user to correct.
|  | 7 | If required the index order can be modified, select the alternative i5 index to be used from the drop down options for the required row. NOTE: If duplicate i5 indexes are selected then the cells will be flagged in red to prompt the user to correct.
|  | 8 | Once all the actions above have been completed, click on the 1.2 SampleSheet tab and review. Red text will be used to flag where information is still required. If no red text is evident then the SampleSheet tab can be saved as a CSV (Comma delimited) (*.csv). Save the file as 'SampleSheet.csv'. Saving in a csv format will only save the active tab. Open the saved SampleSheet.csv file in excel and delete any empty rows in the [Data] table, that is any rows between 22 to 117 which do not contain sample information. Once this has been done, save the file. The sample sheet is then ready to be imported for sequencing on an Illumina sequencer.
|  |  |  | PLATE LAYOUT |  | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12
|  |  |  | Select |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | A | 0
|  |  |  | B | 0
|  |  |  | C | 0
|  |  |  | D | 0
|  |  |  | E | 0
|  |  |  | F | 0
|  |  |  | G | 0
|  |  |  | H | 0
PROCEED TO LIBRARY PREPARATION.
ImportData!A1
ImportData!A2
ImportData!A3
ImportData!A4
ImportData!A5
ImportData!A6
ImportData!A7
ImportData!A8
ImportData!A9
ImportData!A10
ImportData!A11
ImportData!A12
ImportData!A13
ImportData!A14
ImportData!A15
ImportData!A16
ImportData!A17
ImportData!A18
ImportData!A19
ImportData!A20
ImportData!A21
ImportData!A22
ImportData!A23
ImportData!A24

## 1.1 LinearView
SAMPLE LAYOUT - LINEAR VIEW
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
| Sample # | Well # | Sample_ID | Sample_Name | Sample_Plate | Sample_Well | I7_Index_ID | index | I5_Index_ID | index2 | Sample_Project
| 1 | A01 | -HLA-Enter | -HLA-Enter | Enter | A01 | 0 |  | 0 |  | Enter
| 2 | B01 | -HLA-Enter | -HLA-Enter | Enter | B01 | 0 |  | 0 |  | Enter
| 3 | C01 | -HLA-Enter | -HLA-Enter | Enter | C01 | 0 |  | 0 |  | Enter
| 4 | D01 | -HLA-Enter | -HLA-Enter | Enter | D01 | 0 |  | 0 |  | Enter
| 5 | E01
| 6 | F01
| 7 | G01
| 8 | H01
| 9 | A02 | -HLA-Enter | -HLA-Enter | Enter | A02 | 0 |  | 0 |  | Enter
| 10 | B02 | -HLA-Enter | -HLA-Enter | Enter | B02 | 0 |  | 0 |  | Enter
| 11 | C02 | -HLA-Enter | -HLA-Enter | Enter | C02 | 0 |  | 0 |  | Enter
| 12 | D02 | -HLA-Enter | -HLA-Enter | Enter | D02 | 0 |  | 0 |  | Enter
| 13 | E02
| 14 | F02
| 15 | G02
| 16 | H02
| 17 | A03 | -HLA-Enter | -HLA-Enter | Enter | A03 | 0 |  | 0 |  | Enter
| 18 | B03 | -HLA-Enter | -HLA-Enter | Enter | B03 | 0 |  | 0 |  | Enter
| 19 | C03 | -HLA-Enter | -HLA-Enter | Enter | C03 | 0 |  | 0 |  | Enter
| 20 | D03 | -HLA-Enter | -HLA-Enter | Enter | D03 | 0 |  | 0 |  | Enter
| 21 | E03
| 22 | F03
| 23 | G03
| 24 | H03
| 25 | A04 | -HLA-Enter | -HLA-Enter | Enter | A04 | 0 |  | 0 |  | Enter
| 26 | B04 | -HLA-Enter | -HLA-Enter | Enter | B04 | 0 |  | 0 |  | Enter
| 27 | C04 | -HLA-Enter | -HLA-Enter | Enter | C04 | 0 |  | 0 |  | Enter
| 28 | D04 | -HLA-Enter | -HLA-Enter | Enter | D04 | 0 |  | 0 |  | Enter
| 29 | E04
| 30 | F04
| 31 | G04
| 32 | H04
| 33 | A05 | -HLA-Enter | -HLA-Enter | Enter | A05 | 0 |  | 0 |  | Enter
| 34 | B05 | -HLA-Enter | -HLA-Enter | Enter | B05 | 0 |  | 0 |  | Enter
| 35 | C05 | -HLA-Enter | -HLA-Enter | Enter | C05 | 0 |  | 0 |  | Enter
| 36 | D05 | -HLA-Enter | -HLA-Enter | Enter | D05 | 0 |  | 0 |  | Enter
| 37 | E05
| 38 | F05
| 39 | G05
| 40 | H05
| 41 | A06 | -HLA-Enter | -HLA-Enter | Enter | A06 | 0 |  | 0 |  | Enter
| 42 | B06 | -HLA-Enter | -HLA-Enter | Enter | B06 | 0 |  | 0 |  | Enter
| 43 | C06 | -HLA-Enter | -HLA-Enter | Enter | C06 | 0 |  | 0 |  | Enter
| 44 | D06 | -HLA-Enter | -HLA-Enter | Enter | D06 | 0 |  | 0 |  | Enter
| 45 | E06
| 46 | F06
| 47 | G06
| 48 | H06
| 49 | A07
| 50 | B07
| 51 | C07
| 52 | D07
| 53 | E07
| 54 | F07
| 55 | G07
| 56 | H07
| 57 | A08
| 58 | B08
| 59 | C08
| 60 | D08
| 61 | E08
| 62 | F08
| 63 | G08
| 64 | H08
| 65 | A09
| 66 | B09
| 67 | C09
| 68 | D09
| 69 | E09
| 70 | F09
| 71 | G09
| 72 | H09
| 73 | A10
| 74 | B10
| 75 | C10
| 76 | D10
| 77 | E10
| 78 | F10
| 79 | G10
| 80 | H10
| 81 | A11
| 82 | B11
| 83 | C11
| 84 | D11
| 85 | E11
| 86 | F11
| 87 | G11
| 88 | H11
| 89 | A12
| 90 | B12
| 91 | C12
| 92 | D12
| 93 | E12
| 94 | F12
| 95 | G12
| 96 | H12
END OF SAMPLE LAYOUT

## 1.2 SampleSheet
[Header]
IEMFileVersion | 5
Investigator Name | Enter
Experiment Name | Enter
Date | Enter
Workflow | GenerateFASTQ
Application | FASTQ Only
Instrument Type | Select
Assay | Select
Description | Enter
Chemistry | Amplicon
[Settings]
[Reads]
151
151
[Data]
Sample_ID | Sample_Name | Sample_Plate | Sample_Well | I7_Index_ID | index | I5_Index_ID | index2 | Sample_Project | Autolauncher_Project
-HLA-Enter | -HLA-Enter | Enter | A01 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | B01 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | C01 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | D01 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | A02 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | B02 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | C02 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | D02 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | A03 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | B03 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | C03 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | D03 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | A04 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | B04 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | C04 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | D04 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | A05 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | B05 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | C05 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | D05 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | A06 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | B06 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | C06 | 0 |  | 0 |  | Enter | Enter
-HLA-Enter | -HLA-Enter | Enter | D06 | 0 |  | 0 |  | Enter | Enter

## 2.0 Library_Prep
ALLOSEQ TX - LIBRARY PREPARATION
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Sample (µL) | Volume for # of Samples Required (µL): | 0 | Preparation Required
|  |  |  | Tagmentation Buffer |  |  | -15°C to -25°C BOX 1 | 10 | 0 |  | Bring to room temperature.
|  |  |  | Sterile Water |  |  | 15°C to 30°C User supplied | 30 | 0 |  | No preparation required.
|  |  |  | Tagmentation Beads |  |  | -15°C to -25°C BOX 1 | 10 | 0 |  | Bring to room temperature, at least 30 mins.
|  |  |  | Stop Buffer |  |  | 15°C to 30°C BOX 2 | 10 | 0 |  | No preparation required.
|  |  |  | Tagmentation Wash Buffer |  |  | 15°C to 30°C BOX 2 | 300 | 0 |  | No preparation required.
|  |  |  | PCR Mix |  |  | -15°C to -25°C BOX 4 | 20 | 0 |  | Thaw and keep on ice. Return to storage after use for subsequent steps.
|  |  |  | Select index set on 1.0 Sample_Prep tab. |  |  | -15°C to -25°C BOX 3 | 5 | 0 |  | Thaw and keep on ice.
|  | 2 | Vortex and pulse-spin all reagents before use.
|  | 3 | Gather required consumables listed below:
96-Well Half-Skirted Plate,
Microseal B,
1.5 mL microcentrifuge tube or 5mL tubes.
LIBRARY PREPARATION PROTOCOL
NOTE: There is no safe stopping point until the indexing PCR. Process takes approximately 1 hours to set up, 50 minutes in thermocycler.
|  | 1 | For each DNA sample, aliquot 10 µL of 10-100 ng/µL sample into the appropriate well of a PCR plate according to plate layout on 1.0 Sample_Prep sheet.
|  | 2 | Prepare a tagmentation master mix of the following reagents:
|  |  |  | Reagent |  | Volume per Sample (µL) | Volume for Number of Samples Required +10% Excess (µL): | 0
|  |  |  | Tagmentation Buffer |  | 10 | 0
|  |  |  | Sterile Water |  | 20 | 0
|  |  |  | Tagmentation Beads |  | 10 | 0
|  |  |  | Total |  | 40 | 0
|  | 3 | Mix the above master mix by vortexing and then pulse-spin.
|  | 4 | Pipette 40 µL of the tagmentation master mix into each well that contains a DNA sample.
|  | 5 | Seal the plate with Microseal B film.
|  | 6 | Centrifuge the plate at 100 x g for 10 seconds to collect all reagents at the bottom of the well.
|  | 7 | Use plate shaker to mix at 1800 rpm for 1 minute.
|  | 8 | Visually inspect the plate:
|  |  | a) | If there are beads that are not evenly distributed in the well, repeat shaking as per step 7,
|  |  | b) | If material is not in the bottom of the well or has splashed onto the Microseal B film, pulse-spin and repeat shaking step 7.
|  | 9 | Place plate in thermocycler and run the Tagmentation program.
Thermocycler Used:
Program Name:
Tagmentation program:
Reaction Volume: 50 µL
Lid temperature: 105°C
|  |  |  | Temperature |  | Time
|  |  |  | 55°C |  | 5 minutes
|  |  |  | 10°C |  | 2 minutes
|  | 10 | Once the program has finished, immediately remove plate from thermocycler and leave at room temperature for 2 minutes.
|  | 11 | Remove Microseal B film.
|  | 12 | Add 10 µL of Stop Buffer to each reaction well.
|  | 13 | Re-seal the plate with new Microseal B film.
|  | 14 | Use plate shaker to mix at 1800 rpm for 1 minute.
|  | 15 | Incubate the plate for an additional 5 minutes at room temperaure.
|  | 16 | During incubation remove PCR Mix and Indexes from the freezer to thaw, and then place on ice/cold rack.
|  | 17 | Visually inspect the plate, if material is not at the bottom of the well or has splashed on to the Microseal B film, pulse-spin the plate.
|  | 18 | Wash three times with Tagmentation Wash Buffer as described below:
|  |  |  | a) | Remove Microseal B and place the plate on magnet for 30 seconds, allowing beads to collect in wells alongside magnet,
|  |  |  | b) | Using a pipette, aspirate and discard supernatant, leaving the beads in the wells alongside the magnet,
|  |  |  | c) | Slowly add 100 µL of Tagmentation Wash Buffer to each well,
|  |  |  | d) | Re-seal the plate with new Microseal B film, ensure that the film is properly attached,
|  |  |  | e) | Use plate shaker to mix at 1800 rpm for 2 minutes at room temperature,
|  |  |  | f) | CAUTION: If samples have splashed onto the seal, centrifuge at 100 x g for 10 secs before removing the seal,
|  |  |  | g) | Repeat steps a) through f) two more times for a total of 3 washes.
NOTE: Before discarding the supernatant of the third wash, prepare the PCR master mix as described below.
|  | 19 | Prepare PCR master mix according to the following volumes:
|  |  |  | Reagent |  | Volume per Sample (µL) | Volume for Number of Samples Required +10% Excess (µL): | 0
|  |  |  | Sterile Water |  | 20 | 0
|  |  |  | PCR Mix |  | 20 | 0
|  |  |  | Total |  | 40 | 0
|  | 20 | Remove Microseal B and place the plate on magnet for 30 seconds, allowing beads to collect in wells alongside magnet.
|  | 21 | Using a pipette set to 100 µL, aspirate and discard the supernatant from the final Tagmentation Wash.
|  | 22 | Remove the plate from the magnetic stand.
|  | 23 | Add 40 µL of the PCR master mix to each well.
|  | 24 | Seal the plate with Microseal B film.
|  | 25 | Use plate shaker to mix at 1800 rpm for 2 minutes at room temperature.
|  | 26 | Centrifuge the plate at 100 x g for 10 seconds to ensure all beads are suspended within the PCR master mix.
For index tubes (T),
27 (T) Vortex and pulse-spin index tubes to ensure all volume is at the bottom of the tube. 28 (T) Remove the Microseal B film from the sample plate. 29 (T) Add 5 µL of the i7 index to each well, according to the plate layout on 1.0 Sample_Prep sheet. 30 (T) Add 5 µL of the i5 index to each well, according to the plate layout on 1.0 Sample_Prep sheet. Resume to step 31.
For index plate (P),
27 (P) Use a plate shaker to mix the index plate at 1800 rpm for 1 minute. 28 (P) Centrifuge the index plate at 100 x g for 10 seconds to ensure all volume is at the bottom of the well. 29 (P) Remove the Microseal B film from the sample plate. 30 (P) Confirm the correct orientation of the plate and correct index set. Do not peel the foil seal. Pierce the foil seal of the index plate with a tip. With a new tip, transfer 10 µL of the combined indices from the index plate to each sample well, according to the plate layout on 1.0 Sample_Prep sheet. Resume to step 31.
|  | 31 | Seal with new Microseal B film.
|  | 32 | Use plate shaker to mix at 1800 rpm for 1 minute at room temperature.
|  | 33 | Centrifuge the plate at 100 x g for 10 seconds to collect all reagents at the bottom of the well. Visually inspect to ensure beads are still evenly distributed in solution. If beads aren't evenly distributed, repeat shaking as per step 32.
|  | 34 | Place plate in thermocycler and run the Indexing PCR program.
Thermocycler Used:
Program Name:
Indexing PCR program:
Reaction Volume: 50 µL
Lid temperature: 105°C
|  |  |  | # | Step | Temperature | Time | Number of Cycles
|  |  |  | 1 | Gap fill | 72°C | 3 minutes | 1
|  |  |  | 2 | Initial denature | 98°C | 3 minutes | 1
|  |  |  | 3 | Denaturation | 98°C | 20 seconds | 9
|  |  |  | 4 | Annealing | 60°C | 30 seconds
|  |  |  | 5 | Extension | 72°C | 3 minutes
|  |  |  | 6 | Final extension | 72°C | 3 minutes | 1
|  |  |  | 7 | Final hold | 10°C | Hold | 1
NOTE: Safe stopping point. Libraries may be stored at -15°C to -25°C for up to 1 week.
NOTE: If proceeding immediately to size selection and purification, ensure Purification Beads are brought to room temperature before use.
Finish time:
PROTOCOL END. PROCEED TO SIZE SELECTION /PURIFICATION STEP.

## 3.0 SS_Purification
ALLOSEQ TX - SIZE SELECTION AND PURIFICATION
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
SIZE SELECTION CALCULATIONS
For optimal yield and size selection, the sample, bead and supernatant volume are varied according to run size. The table below outlines the volumes tested for the specified range of samples per run.
|  | Reagent Volumes For Calculations |  |  |  |  | 6 - 24 Samples/Pool | 25 - 48 Samples/Pool | 49 - 96 Samples/Pool | Volume to be Used for Run Based on Number of Samples
|  | Number of samples to be pooled (max. listed for ranges) |  |  |  |  | 24 | 48 | 96 | 0
|  | Volume of library to be pooled per sample (µL) |  |  |  |  | 10 | 5 | 2.5 | 0
|  | Volume of diluted beads per sample (µL) |  |  |  |  | 50 | 25 | 12.5 | 0
|  | Transfer volume of supernatant (µL) |  |  |  |  | 55 | 27.5 | 13.75 | 0
|  | Volume of neat beads per sample (µL) |  |  |  |  | 4.4 | 2.2 | 1.1 | 0
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Sample or Pool (µL) | Volume for # of Samples Required (µL): | 0 | Preparation Required
|  |  |  | Purification Beads |  |  | 2°C to 8°C BOX 5 | 24.7 | 0 |  | Bring to room temperature, at least 30 mins. Return to storage after use for subsequent steps.
|  |  |  | Sterile Water |  |  | 15°C to 30°C User supplied | Varied | 720 |  | No preparation required.
|  |  |  | 100% Ethanol |  |  | 15°C to 30°C User supplied | 1920 | 2880 |  | Prepare 80% ethanol, as described below.
|  |  |  | Resuspension Buffer |  |  | 2°C to 8°C BOX 5 | 37 | 37 |  | Bring to room temperature. Return to storage after use for subsequent steps
|  | 2 | Prepare diluted Purification Beads - ensure neat Purification Beads are thoroughly vortexed before use:
|  |  |  | Reagent |  | Volume per sample (µL) | Volume for Number of Samples Required +20% Excess (µL): | 0
|  |  |  | Purification Beads |  | 0 | 0
|  |  |  | Sterile Water |  | 0 | 0
|  |  |  | Total |  | 0 | 0
|  | 3 | Prepare fresh 80% ethanol (2 washes per pool):
|  |  |  | Reagent |  | Volume per pool (µL) | Volume for Number of pools Required +50% Excess (µL): | 1
|  |  |  | 100% Ethanol |  | 1920 | 2880
|  |  |  | Sterile Water |  | 480 | 720
|  |  |  | Total |  | 2400 | 3600
NOTE: The number of pools per experiment is preset to 1 in the yellow cell above. This can be manually overridden if required. Modification to this cell will update the number of pools for subsequent steps of this protocol/workbook. If more than 1 pool are to be performed on this experiement with the same number of samples per pool, the volume of purification beads should be increased accordingly (i.e. doubled for two pools), and the instructions below should be performed in full for each pool. If multiple pools with different number of samples are required, this tab should be duplicated and the cells in yellow updated to reflect the number of samples, so that the volumes of reagent required/transferred are correctly adjusted. In order to duplicate the tab, right click on any of the tabs and select 'Move or Copy', select '3.0 SS_Purification' from the list, check the 'Create Copy' box and click 'OK'.
|  | 4 | Vortex and pulse-spin all reagents before use.
|  | 5 | Gather required consumables listed below:
1.5 mL tubes (Optional: 2 mL tubes can be used if preferred, with increase of ethanol wash volume from 1200 µL to 1500 µL in step 14a).
SIZE SELECTION AND PURIFICATION PROTOCOL
NOTE: Process takes approximately 1 hour.
|  | 1 | Aliquot 0 µL of diluted Purification Beads into the 1.5mL tube.
|  | 2 | Pipette mix the Tagmentation beads and supernatant from indexing PCR or shake the indexing PCR plate for 1 minute at 1800 rpm, and then add 0 µL of each sample to tube containing diluted Purification Beads.
|  | 3 | Vortex each tube at high speed until the sample appears homogenous upon visual inspection.
|  | 4 | Incubate at room temperature for 5 minutes. During this incubation the largest fragments are binding to the beads.
|  | 5 | Quickly pulse-spin tube.
|  | 6 | Place tube on magnet for 2.5 minutes, allowing beads to collect in tube alongside magnet.
|  | 7 | Transfer 0 µL of the supernatant to a new tube.
|  | 8 | Add 0 µL of Purification Beads (undiluted) to the tube containing the supernatant.
|  | 9 | Vortex each tube at high speed for 10 seconds.
|  | 10 | Incubate at room temperature for 5 minutes. During this incubation the target sized fragments are binding to the beads.
|  | 11 | Quickly pulse-spin tube.
|  | 12 | Place tube on magnet for 2.5 minutes, allowing beads to collect in tube alongside magnet.
|  | 13 | Using a pipette, aspirate and discard supernatant, leaving the beads in the tube alongside the magnet.
|  | 14 | Keeping the tube on the magnet, wash twice with 80% ethanol:
|  |  | a) | Add 1200 µL of 80% ethanol to each tube,
|  |  | b) | Incubate at room temperature for 30 seconds,
|  |  | c) | Using a pipette, aspirate and discard all supernatant,
|  |  | d) | Repeat steps a) through c) for a total of 2 washes.
|  | 15 | Remove all remaining supernatant with P20 pipette.
|  | 16 | Air dry the tube for 5 minutes at room temperature, to allow residual ethanol to evaporate.
|  | 17 | Take the tube off the magnet and add 37 µL of Resuspension Buffer to each tube to elute target fragments.
|  | 18 | Vortex each tube at high speed for 10 seconds.
|  | 19 | Incubate at room temperature for 5 minutes.
|  | 20 | Quickly pulse-spin tube.
|  | 21 | Place tube on magnet for 30 seconds, allowing beads to collect in tube alongside magnet.
|  | 22 | Transfer 35 µL of supernatant to a new 1.5 mL tube for storage.
NOTE: Safe stopping point. Libraries may be stored at -15°C to -25°C for up to 1 month.
Finish time:
PROTOCOL END. PROCEED TO HYBRIDISATION. (OR OPTIONAL QUANTIFICATION/VISUALISATION ACCORDING TO 8.0 QUBIT OR 9.0 TAPESTATION)

## 4.0 Hybridisation
ALLOSEQ TX - PROBE HYBRIDISATION
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1 | Preparation Required
|  |  |  | AlloSeq Tx Probe Panel |  |  | -15°C to -25°C BOX 4 | 10 | 10 |  | Bring to room temperature.
|  |  |  | Hybridisation Buffer 1 |  |  | -15°C to -25°C BOX 4 | 50 | 50 |  | Place in Hybex at 58 ͦC for 15 minutes. Vortex and visually inspect, if precipitate remains then incubate at 58 ͦC for another 15 minutes.
|  |  |  | Hybridisation Buffer 2 |  |  | 2°C to 8°C BOX 5 | 10 | 10 |  | Bring to room temperature.
|  | 2 | Vortex and pulse-spin all reagents before use.
|  | 3 | Gather required consumables listed below:
PCR tubes/strips and caps.
PROBE HYBRIDISATION PROTOCOL
NOTE: There is no safe stopping point until after the capture protocol.
NOTE: Process takes approximately 20 minutes to set up, 1.5 hours in thermocycler.
|  | 1 | For each hybridisation reaction, combine the following reagents in the order listed below into a PCR tube/strip:
|  |  |  | Reagent |  | Volume per Pool (µL)
|  |  |  | Pool of Sample Libraries |  | 30
|  |  |  | AlloSeq Tx Probe Panel |  | 10
|  |  |  | Hybridisation Buffer 1 |  | 50
|  |  |  | Hybridisation Buffer 2 |  | 10
|  |  |  | Total |  | 100
|  | 2 | With a pipette set to 70 µL, pipette mix each hybridisation reaction well 10 times, cap and then pulse-spin.
|  | 3 | If the solution remains cloudy, pipette mix a further 6-8 times, cap and then pulse-spin.
|  | 4 | Place tube/strip in thermocycler and run the Hybridisation program.
Thermocycler Used:
Program Name:
Hybridisation program:
Reaction Volume: 100 µL
Lid temperature: 100°C
|  |  |  | # | Step | Temperature | Time | Number of Cycles
|  |  |  | 1 | Denaturation | 98°C | 5 minutes | 1
|  |  |  | 2 | Ramp Down | 98°C - 62°C, decreasing 2°C/cycle | 1 minute | 1
|  |  |  | 3 | Go to step 2 for 18 more cycles (total of 19 cycles) decreasing 2 ͦC/cycle.
|  |  |  | 4 | Hybridisation | 62°C | 60 minutes | 1
|  |  |  | 5 | Final Hold | 62°C | Hold (do not exceed 18 hours at 62°C, inclusive of step #4) | 1
|  | 5 | Leave the tube/strip in the thermocycler until ready to proceed with the capture. Ensure capture beads have reached room temperature, and that Capture Wash Buffer and Hybex are heated to 58°C.
Finish time:
PROCEED IMMEDIATELY TO CAPTURE.

## 5.0 Capture
ALLOSEQ TX - CAPTURE
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1 | Preparation Required
|  |  |  | Capture Beads |  |  | 2°C to 8°C BOX 5 | 250 | 250 |  | Bring to room temperature, at least 30 mins.
|  |  |  | Capture Wash Buffer |  |  | -15°C to -25°C BOX 4 | 800 | 900 |  | Pre-warm to 58 ͦC before use.
|  |  |  | Capture Elution Buffer 1 |  |  | -15°C to -25°C BOX 4 | 28.5 | 28.5 |  | Bring to room temperature.
|  |  |  | 2N NaOH |  |  | -15°C to -25°C BOX 4 | 1.5 | 1.5 |  | Bring to room temperature. Return to storage after use for subsequent steps.
|  |  |  | Capture Elution Buffer 2 |  |  | 2°C to 8°C BOX 5 | 4 | 4 |  | Bring to room temperature
|  | 2 | Prepare an elution master mix of the following reagents:
|  |  |  | Reagent |  | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1
|  |  |  | Capture Elution Buffer 1 |  | 28.5 | 28.5
|  |  |  | 2N NaOH |  | 1.5 | 1.5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes,
PCR tubes/strips and caps.
CAPTURE PROTOCOL
NOTE: Process takes approximately 1 hour.
|  | 1 | For each hybridisation reaction, add 250 µL of Capture Beads to a fresh 1.5 mL tube.
|  | 2 | Transfer 100 µL of each hybridisation reaction to the respective tube containing Capture Beads.
|  | 3 | Vortex tube at high speed for 10 seconds. Do not centrifuge or pulse-spin.
|  | 4 | Incubate tube at 58°C in Hybex for 15 minutes.
|  | 5 | Quickly pulse-spin tube.
|  | 6 | Immediately place tube on magnet for 30 seconds, allowing beads to collect alongside magnet.
|  | 7 | Using a pipette, aspirate and discard supernatant, leaving the beads in the tube alongside the magnet.
|  | 8 | Wash three times with heated Capture Wash Buffer as described below:
NOTE: When not in use, keep the Capture Wash Buffer in the Hybex to maintain 58°C temperature. Only remove from the Hybex immediately prior to use when adding to reaction in step 8b and 10, and return to the Hybex while the library tube is incubating in the Hybex or on the magnet.
|  |  |  | a) | Remove tube from the magnet.
|  |  |  | b) | Add 200 µL of heated Capture Wash Buffer (58°C).
|  |  |  | c) | Vortex tube at high speed for 10 seconds. Do not centrifuge or pulse-spin.
|  |  |  | d) | Incubate tube at 58°C in Hybex for 5 minutes.
|  |  |  | e) | Pulse-spin and then immediately place tube on magnet for 30 seconds, allowing beads to collect alongside magnet.
|  |  |  | f) | Using a pipette, aspirate and discard supernatant, leaving the beads in the tube alongside the magnet.
|  |  |  | g) | Repeat steps a) through f) two more times for a total of 3 washes.
|  | 9 | Remove tube from the magnet.
|  | 10 | Add 200 µL of heated Capture Wash Buffer (58°C).
|  | 11 | Vortex tube at high speed for 10 seconds. Do not centrifuge or pulse-spin.
|  | 12 | Transfer the entire contents (wash solution and beads) to a new 1.5 mL tube.
NOTE: This transfer step is critical to remove PCR inhibitors which may impact downstream performance.
|  | 13 | Incubate tube at 58°C in Hybex for 5 minutes.
|  | 14 | Immediately pulse-spin and then place tube on magnet for 30 seconds, allowing beads to collect alongside magnet.
|  | 15 | Using a pipette, aspirate and discard supernatant, leaving the beads in the tube alongside the magnet.
|  | 16 | Quickly pulse-spin tube, and then immediately place tube on magnet for 30 seconds, allowing beads to collect alongside magnet.
|  | 17 | Using a P20 pipette, aspirate and discard any remaining supernatant, leaving the beads in the tube alongside the magnet.
|  | 18 | Vortex the elution master mix prepared above, and then take the reaction tube off the magnet and add 23 µL of elution master mix to each tube.
|  | 19 | Vortex tube at high speed for 10 seconds.
|  | 20 | Incubate at room temperature for 2 minutes.
|  | 21 | Quickly pulse-spin tube.
|  | 22 | Place tube on magnet for 30 seconds, allowing beads to collect alongside magnet.
|  | 23 | Transfer 21 µL of supernatant to a new PCR tube/strip.
|  | 24 | Add 4 µL of Capture Elution Buffer 2.
|  | 25 | Pipette mix 6-8 times. Final volume is 25 µL.
NOTE: Safe stopping point. Libraries may be stored at -15°C to -25°C for up to 24 hours.
Finish time:
PROTOCOL END. PROCEED TO POST-ENRICHMENT PCR.

## 6.0 Enrich_PCR
ALLOSEQ TX - POST-ENRICHMENT PCR
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1 | Preparation Required
|  |  |  | PCR Primers |  |  | -15°C to -25°C BOX 4 | 5 | 5 |  | Thaw on ice. Invert to mix, then centrifuge briefly.
|  |  |  | PCR Mix |  |  | -15°C to -25°C BOX 4 | 20 | 20 |  | Thaw at room temperature and then place on ice.
|  | 2 | Vortex and pulse-spin all reagents before use.
|  | 3 | Gather required consumables listed below:
PCR Tubes containing capture libraries from step 3.0 Capture.
POST-ENRICHMENT PCR PROTOCOL
NOTE: Process takes approximately 5 minutes to set up, 1:40 hours in thermocycler.
|  | 1 | Add 5 µL of PCR Primers to the captured libraries in the PCR tube.
|  | 2 | Add 20 µL of PCR Mix to the tube.
|  | 3 | Pipette mix 10 times.
|  | 4 | Quickly pulse-spin tube.
|  | 5 | Place tube/strip in thermocycler and run the Post-Enrichment PCR program.
Thermocycler Used:
Program Name:
Post-Enrichment PCR program:
Reaction Volume: 50 µL
Lid temperature: 105°C
|  |  |  | # | Step | Temperature | Time | Number of Cycles
|  |  |  | 1 | Denaturation | 98°C | 30 seconds | 1
|  |  |  | 2 | Denaturation | 98°C | 1 minute | 17
|  |  |  | 3 | Annealing | 60°C | 30 seconds
|  |  |  | 4 | Extension | 72°C | 3 minutes
|  |  |  | 5 | Final extension | 72°C | 5 minutes | 1
|  |  |  | 6 | Final hold | 10°C | Hold | 1
NOTE: Safe stopping point. Libraries may be stored at -15°C to -25°C for up to 1 week.
NOTE: If proceeding immediately to purification, ensure Purification Beads are brought to room temperature before use.
Finish time:
PROTOCOL END. PROCEED TO PURIFICATION.

## 7.0 Purification
ALLOSEQ TX - POST-ENRICHMENT PCR PURIFICATION
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1 | Preparation Required
|  |  |  | Purification Beads |  |  | 2°C to 8°C BOX 5 | 27 | 27 |  | Bring to room temperature, at least 30 mins.
|  |  |  | Sterile Water |  |  | 15°C to 30°C User supplied | 80 | 120 |  | No preparation required.
|  |  |  | 100% Ethanol |  |  | 15°C to 30°C User supplied | 320 | 480 |  | Prepare 80% ethanol, as described below.
|  |  |  | Resuspension Buffer |  |  | 2°C to 8°C BOX 5 | 32 | 32 |  | Bring to room temperature. Return to storage after use for subsequent steps.
|  | 2 | Prepare fresh 80% ethanol (2 washes per pool):
|  |  |  | Reagent |  | Volume per Pool (µL) | Volume for Number of Pools Required +50% Excess (µL): | 1
|  |  |  | 100% Ethanol |  | 320 | 480
|  |  |  | Sterile Water |  | 80 | 120
|  |  |  | Total |  | 400 | 600
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes.
PURIFICATION PROTOCOL
NOTE: Process takes approximately 30 minutes.
|  | 1 | For each purification reaction, add 27 µL of thoroughly vortexed Purification Beads to a fresh 1.5 mL tube.
|  | 2 | Transfer 45 µL of each post-enrichment PCR reaction to the respective tube containing Purification Beads.
|  | 3 | Vortex each tube at high speed for 10 seconds.
|  | 4 | Quickly pulse-spin tube.
|  | 5 | Incubate at room temperature for 5 minutes.
|  | 6 | Place tube on magnet for 30 seconds, allowing beads to collect alongside magnet.
|  | 7 | Using a pipette, aspirate and discard supernatant, leaving the beads in the tube alongside the magnet.
|  | 8 | Keeping the tube on the magnet, wash twice with 80% ethanol:
|  |  | a) | Add 200 µL of 80% ethanol to each tube,
|  |  | b) | Incubate at room temperature for 30 seconds,
|  |  | c) | Using a pipette, aspirate and discard all superanatant,
|  |  | d) | Repeat steps a) through c) for a total of 2 washes.
|  | 9 | Remove all remaining supernatant with P20 pipette.
|  | 10 | Air dry the tube for 5 minutes at room temperature, to allow residual ethanol to evaporate.
|  | 11 | Take the tube off the magnet and add 32 µL of Resuspension Buffer to each tube to elute target fragments.
|  | 12 | Vortex each tube at high speed for 10 seconds.
|  | 13 | Incubate at room temperature for 5 minutes.
|  | 14 | Quickly pulse-spin tube.
|  | 15 | Place tube on magnet for 30 seconds, allowing beads to collect in tube alongside magnet.
|  | 16 | Transfer 30 µL of supernatant to a new 1.5 mL tube for storage.
NOTE: Safe stopping point. Libraries may be stored at -15°C to -25°C for up to 1 month.
Finish time:
PROTOCOL END. PROCEED TO QUANTIFICATION.

## 8.0 Qubit
ALLOSEQ TX - QUBIT QUANTIFICATION
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1 | Preparation Required
|  |  |  | Qubit BR buffer |  |  | -25°C to 30°C User supplied | 199 | 796 |  | No preparation required.
|  |  |  | Qubit BR dye |  |  | -25°C to 30°C User supplied | 1 | 4 |  | No preparation required.
|  |  |  | BR Standard #1 |  |  | 2°C to 8°C User Supplied | 10 | 10 |  | Bring to room temperature.
|  |  |  | BR Standard #2 |  |  | 2°C to 8°C User Supplied | 10 | 10 |  | Bring to room temperature.
NOTE: The number of pools is automatically copied over based on the hybridisation tab, but can be manually overridden if required.
|  | 2 | Gather required consumables listed below:
Qubit tubes,
1.5 mL or 5 mL tube for working solution preparation, dependent on volume above.
QUANTIFICATION OF ENRICHED LIBRARIES
|  | 1 | Set up two assay tubes for the standards and one for each pool.
|  | 2 | Prepare the Qubit working solution using the following volumes:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Qubit buffer |  | 796
|  |  |  | Qubit dye |  | 4
|  |  |  | Total |  | 800
NOTE: This volume of Qubit working solution includes excess.
|  | 3 | Vortex working solution for 2-3 seconds, and then pulse-spin.
|  | 4 | Aliquot 190 µL of working solution into each of the Standard tubes.
|  | 5 | Aliquot 198 µL of working solution into each of the Pool tubes.
|  | 6 | Aliquot 10 µL of Standard solution into each of the respective Standard tubes.
|  | 7 | Aliquot 2 µL of each pool into the respective tube.
|  | 8 | Vortex all tubes for 2-3 seconds, and then pulse-spin.
|  | 9 | Incubate the tubes for 2 minutes at room temperature.
|  | 10 | Insert the tubes in the Qubit Fluorometer and take readings (see Qubit protocol for further information).
|  | 11 | Record Qubit readings in the table below.
|  |  | Pool # |  | Tube # | Enrichment Pool ID | 1st reading (ng/µL) | 2nd reading (ng/µL) | Average (ng/µL)
|  |  | 1 | [5 empty] | 0
|  |  | 2 | [5 empty] | 0
|  |  | 3 | [5 empty] | 0
|  |  | 4 | [5 empty] | 0
|  |  | 5 | [5 empty] | 0
|  |  | 6 | [5 empty] | 0
|  |  | 7 | [5 empty] | 0
|  |  | 8 | [5 empty] | 0
|  | 12 | Record results in the electronic workbook to calculate the average.
Finish time:
PROTOCOL END. PROCEED TO DILUTE AND DENATURE.

## 9.0 TapeStation(Optional)
ALLOSEQ TX - TAPESTATION VISUALISATION
| Experiment ID #: |  |  |  | Enter
Operator:
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent |  | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Volume for # of Pools Required (µL): | 1 | Preparation Required
|  |  |  | D1000 Ladder |  |  | 2°C to 8°C User supplied | 1 | 1 |  | Bring to room temperature.
|  |  |  | D1000 Sample Buffer |  |  | 2°C to 8°C User supplied | 3 | 6 |  | Bring to room temperature.
NOTE: The number of pools is automatically copied over based on the hybridisation tab, but can be manually overridden if required.
|  | 2 | Gather required consumables listed below:
D1000 ScreenTape,
TapeStation Optical tube strips (8x Strip),
TapeStation Optical tube strip caps (8x Strip)
SIZE DISTRIBUTION VISUALISATION OF ENRICHED LIBRARIES
|  | 1 | Transfer 1 µL of each enriched library to a new tube.
|  | 2 | Add 1 µL of D1000 Ladder into a reference tube.
|  | 3 | Add 3 µL of D1000 sample buffer to each enriched library tube and reference tube.
|  | 4 | Seal all tubes with caps.
|  | 5 | Vortex all tubes thoroughly using IKA vortex at 2000 rpm for 1 minute.
|  | 6 | Centrifuge briefly to ensure that all samples are at the bottom of the tubes.
|  | 7 | Uncap and load sample tubes into the 2200 TapeStation instrument.
|  | 8 | Select the required tubes on the 2200 TapeStation Controller Software and run the samples.
|  | 9 | Once the run is complete, start TapeStation Analysis Software to view results.
|  | 10 | Record results in a table below:
|  |  | Pool # |  | Tube # | Enrichment Pool ID | Minimum Fragment Size (bp) | Peak Fragment Size (bp) | Maximum Fragment Size (bp)
1
2
3
4
5
6
7
8
Finish time:
PROTOCOL END. PROCEED TO DILUTE AND DENATURE.

## 10.0 PhiX
ALLOSEQ TX - DILUTE AND DENATURE PHIX
| Operator: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Date: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | 10 nM PhiX |  | -15°C to -25°C User supplied | 2 | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 4 | 5 | Bring to room temperature.
|  |  |  | Sterile Water |  | 15°C to 30°C User supplied | 45 | No preparation required.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | 3 | Bring to room temperature.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | 990 | Thaw and then keep on ice.
|  | 2 | Prepare 0.2N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | Sterile Water | 45
|  |  |  | 2N NaOH | 5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
DILUTE AND DENATURE PHIX FOR MISEQ AND MINISEQ RUNS, RESULTING IN 20 pM DENATURED PHIX
|  | 1 | Add 3 µL of Resuspension Buffer to a fresh tube.
|  | 2 | Add 2 µL of 10 nM PhiX to the tube.
|  | 3 | Add 5 µL of 0.2N NaOH working solution (as diluted above) to the tube.
|  | 4 | Vortex and pulse-spin the tube.
|  | 5 | Incubate at room temperature for 5 minutes.
|  | 6 | Add 990 µL prechilled HT1 to the tube containing denatured PhiX.
|  | 7 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: This results in 1 ml of 20 pM PhiX. Denatured PhiX may be stored at -15°C to -25°C for up to 1 month.
Finish time:
FOR MINISEQ RUNS ONLY: FURTHER DILUTE DENATURED PHIX TO 5 pM
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Total Volume Required (µL)
|  | 2 | Dilute PhiX to 5 pM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 20 pM Denatured PhiX |  | 0
|  |  |  | Chilled HT1 |  | 0
|  |  |  | Total |  | 0
|  | 3 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: Denatured PhiX may be stored at -15°C to -25°C for up to 1 month.
Finish time:
FOR ISEQ RUNS ONLY: DILUTE PHIX TO 20 pM (WITHOUT DENATURING)
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Total Volume Required (µL)
|  | 2 | Dilute PhiX to 20 pM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 10 nM PhiX |  | 0
|  |  |  | Resuspension Buffer |  | 0
|  |  |  | Total |  | 0
|  | 3 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: Denatured PhiX may be stored at -15°C to -25°C for up to 1 month.
Finish time:
PROTOCOL END.

## 10.1 MiSeq (hidden)
ALLOSEQ TX - DILUTE AND DENATURE FOR MISEQ
| Flow Cell ID #: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Operator: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | -15°C to -25°C User prepared | #DIV/0! | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 4 | 5 | Bring to room temperature.
|  |  |  | Sterile Water |  | 15°C to 30°C User supplied | 45 | No preparation required.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | 1226.4 | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | #DIV/0! | Bring to room temperature.
|  |  |  | 20 pM PhiX (previously diluted) |  | -15°C to -25°C User supplied | 3.6 | Thaw and then keep on ice.
|  | 2 | Prepare 0.2N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | Sterile Water | 45
|  |  |  | 2N NaOH | 5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
INFORMATION REQUIRED FOR 4 nM DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Enriched Sample Pool Concentration (ng/µL)
|  |  |  | Enriched Sample Pool Fragment Size (bp) |  | 800
|  |  |  | Enriched Sample Pool Concentration (nM) |  | 0
|  |  |  | Total Volume Required (µL) |  | 30
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO 4 nM
|  | 1 | Dilute the sample pool to 4 nM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex and pulse-spin diluted sample pool before further use.
DENATURE 4 nM ALLOSEQ TX ENRICHMENT SAMPLE POOL, RESULTING IN 20 pM DENATURED POOL
|  | 1 | Add 5 µL of 4 nM enriched sample pool to a fresh tube.
|  | 2 | Add 5 µL of 0.2N NaOH working solution (as diluted above) to the tube.
|  | 3 | Vortex and pulse-spin the tube.
|  | 4 | Incubate at room temperature for 5 minutes.
|  | 5 | Add 990 µL prechilled HT1 to the tube containing denatured enriched sample pool.
|  | 6 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: This results in 1 ml of 20 pM sample pool. Pools may be stored at -15°C to -25°C for up to 1 month.
INFORMATION REQUIRED FOR FINAL DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
|  |  |  | Number of pools to be combined for sequencing |  | 1
|  |  |  | Loading Concentration (pM) |  | 12
|  |  |  | PhiX Spike in Concentration (%) |  | 1%
|  |  |  | Dilution Volume, minimum 600 (µL) |  | 600
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO LOADING CONCENTRATION, INCLUDING PHIX SPIKE-IN
|  | 1 | Dilute the sample pool to loading concentration by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 20 pM Sample Pool # 1 |  | 360
|  |  |  | 20 pM Sample Pool # 2 |  | 0
|  |  |  | 20 pM Sample Pool # 3 |  | 0
|  |  |  | 20 pM Sample Pool # 4 |  | 0
|  |  |  | 20 pM Denatured PhiX |  | 3.6
|  |  |  | Chilled HT1 |  | 236.4
|  |  |  | Total |  | 600
|  | 2 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
|  | 3 | Set aside on ice until ready to be loaded onto the MiSeq reagent cartridge for sequencing. Load 600µL of diluted library as per MiSeq protocol.
NOTE: Pools may be stored at -15°C to -25°C for up to 48 hours prior to sequencing.
Finish time:
PROTOCOL END. PROCEED TO LOAD MISEQ ACCORDING TO INSTRUMENT PROTOCOL.

## 10.2 MiniSeq (hidden)
ALLOSEQ TX - DILUTE AND DENATURE FOR MINISEQ
| Flow Cell ID #: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Operator: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | -15°C to -25°C User prepared | #DIV/0! | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 4 | 5 | Bring to room temperature.
|  |  |  | Sterile Water |  | 15°C to 30°C User supplied | 95 | No preparation required.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | 1323.4 | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | #DIV/0! | Bring to room temperature.
|  |  |  | 200 mM Tris-HCl, pH 7.0 |  | 15°C to 30°C User supplied | 5 | No preparation required.
|  |  |  | 5 pM PhiX (previously diluted) |  | -15°C to -25°C User supplied | 1.6 | Thaw and then keep on ice.
|  | 2 | Prepare 0.1N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | Sterile Water | 95
|  |  |  | 2N NaOH | 5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
INFORMATION REQUIRED FOR 1 nM DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Enriched Sample Pool Concentration (ng/µL)
|  |  |  | Enriched Sample Pool Fragment Size (bp) |  | 800
|  |  |  | Enriched Sample Pool Concentration (nM) |  | 0
|  |  |  | Total Volume Required (µL) |  | 100
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO 1 nM
|  | 1 | Dilute the sample pool to 1 nM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex and pulse-spin diluted sample pool before further use.
DENATURE 1 nM ALLOSEQ TX ENRICHMENT SAMPLE POOL, RESULTING IN 5 pM DENATURED POOL
|  | 1 | Add 5 µL of 1 nM enriched sample pool to a fresh tube.
|  | 2 | Add 5 µL of 0.1N NaOH working solution (as diluted above) to the tube.
|  | 3 | Vortex and pulse-spin the tube.
|  | 4 | Incubate at room temperature for 5 minutes.
|  | 5 | Add 5 µL of 200 mM Tris-HCl, at pH 7.0.
|  | 6 | Vortex and pulse-spin the tube.
|  | 7 | Add 985 µL prechilled HT1 to the tube containing denatured enriched sample pool.
|  | 8 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: This results in 1 ml of 5 pM sample pool. Pools may be stored at -15°C to -25°C for up to 1 month.
INFORMATION REQUIRED FOR FINAL DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
|  |  |  | Number of pools to be combined for sequencing |  | 1
|  |  |  | Loading Concentration (pM) |  | 1.6
|  |  |  | PhiX Spike in Concentration (%) |  | 1%
|  |  |  | Dilution Volume, minimum 500 (µL) |  | 500
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO LOADING CONCENTRATION, INCLUDING PHIX SPIKE-IN
|  | 1 | Dilute the sample pool to loading concentration by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 5 pM Sample Pool # 1 |  | 160
|  |  |  | 5 pM Sample Pool # 2 |  | 0
|  |  |  | 5 pM Sample Pool # 3 |  | 0
|  |  |  | 5 pM Sample Pool # 4 |  | 0
|  |  |  | 5 pM Denatured PhiX |  | 1.6
|  |  |  | Chilled HT1 |  | 338.4
|  |  |  | Total |  | 500
|  | 2 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
|  | 3 | Set aside on ice until ready to be loaded onto the MiniSeq reagent cartridge for sequencing. Load 500µL of diluted library as per MiniSeq protocol.
NOTE: Pools may be stored at -15°C to -25°C for up to 48 hours prior to sequencing.
Finish time:
PROTOCOL END. PROCEED TO LOAD MINISEQ ACCORDING TO INSTRUMENT PROTOCOL.

## 10.3 iSeq (hidden)
ALLOSEQ TX - DILUTE FOR ISEQ
| Flow Cell ID #: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Operator: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | -15°C to -25°C User prepared | #DIV/0! | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | #DIV/0! | Bring to room temperature.
|  |  |  | 20 pM PhiX (previously diluted) |  | -15°C to -25°C User supplied | 10 | Thaw and then keep on ice.
|  | 2 | Vortex and pulse-spin all reagents before use.
|  | 3 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
INFORMATION REQUIRED FOR 1 nM DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Enriched Sample Pool Concentration (ng/µL)
|  |  |  | Enriched Sample Pool Fragment Size (bp) |  | 800
|  |  |  | Enriched Sample Pool Concentration (nM) |  | 0
|  |  |  | Total Volume Required (µL) |  | 100
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO 1 nM
|  | 1 | Dilute the sample pool to 1 nM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex and pulse-spin diluted sample pool before further use.
INFORMATION REQUIRED FOR FINAL DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
|  |  |  | Number of pools to be combined for sequencing |  | 1
|  |  |  | Loading Concentration (pM) |  | 200
|  |  |  | PhiX Spike in Concentration (%) |  | 1%
|  |  |  | Dilution Volume, minimum 20 (µL) |  | 100
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO LOADING CONCENTRATION, INCLUDING PHIX SPIKE-IN
|  | 1 | Dilute the sample pool to loading concentration by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 1 nM Sample Pool # 1 |  | 20
|  |  |  | 1 nM Sample Pool # 2 |  | 0
|  |  |  | 1 nM Sample Pool # 3 |  | 0
|  |  |  | 1 nM Sample Pool # 4 |  | 0
|  |  |  | 20 pM Non-denatured PhiX |  | 10
|  |  |  | Resuspension Buffer |  | 70
|  |  |  | Total |  | 100
|  | 2 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
|  | 3 | Set aside on ice until ready to be loaded onto the iSeq reagent cartridge for sequencing. Load 20µL of diluted library as per iSeq protocol.
NOTE: Pools may be stored at -15°C to -25°C for up to 48 hours prior to sequencing.
Finish time:
PROTOCOL END. PROCEED TO LOAD ISEQ ACCORDING TO INSTRUMENT PROTOCOL.

## VersionHistory
VERSION | REVISION DATE (DDMMMYY ) | REVISION DESCRIPTION
1 | 03MAY2021 | Initial release of the CE-IVD AlloSeq Tx assay Early Pooling workbook. This workbook was built from IFU083-5_AlloSeq Tx Assay Workbook RUO, v6.0 Subsequent workbook updates: Deactivation of the 96 Set B/C/D Indexes options. Typographical error corrections. Return to storage after use for subsequent steps statements added in reagent tables where relevant. Added extra guidance around thoroughly mixing Purification Beads to step 1, tab 3.0 and and step 1, tab 7.0 following feedback from field team. Harmonised all instances of pulse spin to pulse-spin. Corrected all instances of 'hybex' to 'Hybex' and 'QuBit' to 'Qubit'. Step 3, tab 3.0 changed from 'Vortex each tube at high speed for 10 seconds' to 'Vortex each tube at high speed until the sample appears homogenous upon visual inspection'. Included clarification comment, "Hold (do not exceed 18 hours at 62°C, inclusive of step #4)" in hybridisation program table. Added 'If beads aren't evenly distributed, repeat shaking as per step 32' in step 33, Library_Prep tab. Added Qubit Broad Range (BR) clarification in reagent table, tab 8.0. Corrected all references to plates/wells to tubes/caps in tab 9.0. Workbook updated to include CE cover page and added CE0197 logo. Updated Volume per Sample (µL) for indexes from 20 to 5 in 2.0 Library_Prep to reflect the actual volume required according to step 29&30. Reissued 20 Oct 21
2 | 28Jan22 | Updated the Library Preparation to include instructions for using the index plate and separated from the index tube instructions. Added clarification to "Confirm the correct orientation of the plate and correct index set. Do not peel the foil seal." Reissued 27Apr22