## CE Cover
AlloSeq Tx
Sequencing Workbook
IFU095-3
Version Number: 4.0
Issue Date: Apr 2022
ASTX17.1(24)-IVD
ASTX17.1(24)-B-IVD
ASTX17.1(96)-A-IVD
ASTX17.1(96)-B-IVD
| CareDx Pty Ltd |  | Qarad BV,
| 20 Collie St. |  | Cipalstraat 3
| Fremantle WA 6160 |  | 2440 Geel
| Australia |  | Belgium
© 2022 CareDx, Inc. All service marks or trademarks are owned or licensed by CareDx, Inc. or its affiliates. All rights reserved.

## 1.0 PhiX
ALLOSEQ TX - DILUTE AND DENATURE PHIX
| Operator: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Date: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | 10 nM PhiX |  | -15°C to -25°C User supplied | 2 | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 4 | 5 | Bring to room temperature.
|  |  |  | Sterile Water |  | 15°C to 30°C User supplied | 45 | No preparation required.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | 3 | Bring to room temperature.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | 990 | Thaw and then keep on ice.
|  | 2 | Prepare 0.2N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | Sterile Water | 45
|  |  |  | 2N NaOH | 5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
DILUTE AND DENATURE PHIX FOR MISEQ AND MINISEQ RUNS, RESULTING IN 20 pM DENATURED PHIX
|  | 1 | Add 3 µL of Resuspension Buffer to a fresh tube.
|  | 2 | Add 2 µL of 10 nM PhiX to the tube.
|  | 3 | Add 5 µL of 0.2N NaOH working solution (as diluted above) to the tube.
|  | 4 | Vortex and pulse-spin the tube.
|  | 5 | Incubate at room temperature for 5 minutes.
|  | 6 | Add 990 µL prechilled HT1 to the tube containing denatured PhiX.
|  | 7 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: This results in 1 ml of 20 pM PhiX. Denatured PhiX may be stored at -15°C to -25°C for up to 1 month.
Finish time:
FOR MINISEQ RUNS ONLY: FURTHER DILUTE DENATURED PHIX TO 5 pM
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Total Volume Required (µL)
|  | 2 | Dilute PhiX to 5 pM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 20 pM Denatured PhiX |  | 0
|  |  |  | Chilled HT1 |  | 0
|  |  |  | Total |  | 0
|  | 3 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: Denatured PhiX may be stored at -15°C to -25°C for up to 1 month.
Finish time:
FOR ISEQ RUNS ONLY: DILUTE PHIX TO 20 pM (WITHOUT DENATURING)
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Total Volume Required (µL)
|  | 2 | Dilute PhiX to 20 pM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 10 nM PhiX |  | 0
|  |  |  | Resuspension Buffer |  | 0
|  |  |  | Total |  | 0
|  | 3 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: Denatured PhiX may be stored at -15°C to -25°C for up to 1 month.
Finish time:
PROTOCOL END.

## 1.1 MiSeq
ALLOSEQ TX - DILUTE AND DENATURE FOR MISEQ
| Flow Cell ID #: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Operator: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | -15°C to -25°C User prepared | #DIV/0! | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 4 | 5 | Bring to room temperature.
|  |  |  | Sterile Water |  | 15°C to 30°C User supplied | 45 | No preparation required.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | 1226.4 | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | #DIV/0! | Bring to room temperature.
|  |  |  | 20 pM PhiX (previously diluted) |  | -15°C to -25°C User supplied | 3.6 | Thaw and then keep on ice.
|  | 2 | Prepare 0.2N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | Sterile Water | 45
|  |  |  | 2N NaOH | 5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
INFORMATION REQUIRED FOR 4 nM DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Enriched Sample Pool Concentration (ng/µL)
|  |  |  | Enriched Sample Pool Fragment Size (bp) |  | 800
|  |  |  | Enriched Sample Pool Concentration (nM) |  | 0
|  |  |  | Total Volume Required (µL) |  | 30
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO 4 nM
|  | 1 | Dilute the sample pool to 4 nM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex and pulse-spin diluted sample pool before further use.
DENATURE 4 nM ALLOSEQ TX ENRICHMENT SAMPLE POOL, RESULTING IN 20 pM DENATURED POOL
|  | 1 | Add 5 µL of 4 nM enriched sample pool to a fresh tube.
|  | 2 | Add 5 µL of 0.2N NaOH working solution (as diluted above) to the tube.
|  | 3 | Vortex and pulse-spin the tube.
|  | 4 | Incubate at room temperature for 5 minutes.
|  | 5 | Add 990 µL prechilled HT1 to the tube containing denatured enriched sample pool.
|  | 6 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: This results in 1 ml of 20 pM sample pool. Pools may be stored at -15°C to -25°C for up to 1 month.
INFORMATION REQUIRED FOR FINAL DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
|  |  |  | Number of pools to be combined for sequencing |  | 1
|  |  |  | Loading Concentration (pM) |  | 12
|  |  |  | PhiX Spike in Concentration (%) |  | 1%
|  |  |  | Loading Volume |  | 600
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO LOADING CONCENTRATION, INCLUDING PHIX SPIKE-IN
|  | 1 | Dilute the sample pool to loading concentration by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 20 pM Sample Pool # 1 |  | 360
|  |  |  | 20 pM Sample Pool # 2 |  | 0
|  |  |  | 20 pM Sample Pool # 3 |  | 0
|  |  |  | 20 pM Sample Pool # 4 |  | 0
|  |  |  | 20 pM Denatured PhiX |  | 3.6
|  |  |  | Chilled HT1 |  | 236.4
|  |  |  | Total |  | 600
|  | 2 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
|  | 3 | Set aside on ice until ready to be loaded onto the MiSeq reagent cartridge for sequencing.
NOTE: Pools may be stored at -15°C to -25°C for up to 48 hours prior to sequencing.
Finish time:
PROTOCOL END. PROCEED TO LOAD MISEQ ACCORDING TO INSTRUMENT PROTOCOL.

## 1.2 MiniSeq
ALLOSEQ TX - DILUTE AND DENATURE FOR MINISEQ
| Flow Cell ID #: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Operator: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | -15°C to -25°C User prepared | #DIV/0! | Thaw and then keep on ice.
|  |  |  | 2N NaOH |  | -15°C to -25°C BOX 4 | 5 | Bring to room temperature.
|  |  |  | Sterile Water |  | 15°C to 30°C User supplied | 95 | No preparation required.
|  |  |  | HT1 (with MiSeq cartridge) |  | -15°C to -25°C User supplied | 1323.4 | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | #DIV/0! | Bring to room temperature.
|  |  |  | 200 mM Tris-HCl, pH 7.0 |  | 15°C to 30°C User supplied | 5 | No preparation required.
|  |  |  | 5 pM PhiX (previously diluted) |  | -15°C to -25°C User supplied | 1.6 | Thaw and then keep on ice.
|  | 2 | Prepare 0.1N NaOH working solution
|  |  |  | Reagent | Volume (µL)
|  |  |  | Sterile Water | 95
|  |  |  | 2N NaOH | 5
NOTE: NaOH solution will readily absorb CO2 from the atmosphere, altering the pH and performance of the reagent. Ensure the 2N NaOH tube is sealed when not in use.
|  | 3 | Vortex and pulse-spin all reagents before use.
|  | 4 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
INFORMATION REQUIRED FOR 1 nM DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Enriched Sample Pool Concentration (ng/µL)
|  |  |  | Enriched Sample Pool Fragment Size (bp) |  | 800
|  |  |  | Enriched Sample Pool Concentration (nM) |  | 0
|  |  |  | Total Volume Required (µL) |  | 100
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO 1 nM
|  | 1 | Dilute the sample pool to 1 nM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex and pulse-spin diluted sample pool before further use.
DENATURE 1 nM ALLOSEQ TX ENRICHMENT SAMPLE POOL, RESULTING IN 5 pM DENATURED POOL
|  | 1 | Add 5 µL of 1 nM enriched sample pool to a fresh tube.
|  | 2 | Add 5 µL of 0.1N NaOH working solution (as diluted above) to the tube.
|  | 3 | Vortex and pulse-spin the tube.
|  | 4 | Incubate at room temperature for 5 minutes.
|  | 5 | Add 5 µL of 200 mM Tris-HCl, at pH 7.0.
|  | 6 | Vortex and pulse-spin the tube.
|  | 7 | Add 985 µL prechilled HT1 to the tube containing denatured enriched sample pool.
|  | 8 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
NOTE: This results in 1 ml of 5 pM sample pool. Pools may be stored at -15°C to -25°C for up to 1 month.
INFORMATION REQUIRED FOR FINAL DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
|  |  |  | Number of pools to be combined for sequencing |  | 1
|  |  |  | Loading Concentration (pM) |  | 1.6
|  |  |  | PhiX Spike in Concentration (%) |  | 1%
|  |  |  | Loading Volume |  | 500
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO LOADING CONCENTRATION, INCLUDING PHIX SPIKE-IN
|  | 1 | Dilute the sample pool to loading concentration by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 5 pM Sample Pool # 1 |  | 160
|  |  |  | 5 pM Sample Pool # 2 |  | 0
|  |  |  | 5 pM Sample Pool # 3 |  | 0
|  |  |  | 5 pM Sample Pool # 4 |  | 0
|  |  |  | 5 pM Denatured PhiX |  | 1.6
|  |  |  | Chilled HT1 |  | 338.4
|  |  |  | Total |  | 500
|  | 2 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
|  | 3 | Set aside on ice until ready to be loaded onto the MiniSeq reagent cartridge for sequencing.
NOTE: Pools may be stored at -15°C to -25°C for up to 48 hours prior to sequencing.
Finish time:
PROTOCOL END. PROCEED TO LOAD MINISEQ ACCORDING TO INSTRUMENT PROTOCOL.

## 1.3 iSeq
ALLOSEQ TX - DILUTE FOR ISEQ
| Flow Cell ID #: |  |  |  |  | NOTE: GREEN fields are to be completed by operator during set up.
| Operator: |  |  |  |  | NOTE: YELLOW fields must be completed before printing.
Date:
Start time:
REAGENT PREPARATION AND BATCH RECORDS
|  | 1 | Gather the required reagents (volume specified in table below) and prepare according to table:
|  |  |  | Reagent | Lot # and Expiry Date | Storage Conditions | Volume per Pool (µL) | Preparation Required
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | -15°C to -25°C User prepared | #DIV/0! | Thaw and then keep on ice.
|  |  |  | Resuspension Buffer |  | 2°C to 8°C BOX 5 | #DIV/0! | Bring to room temperature.
|  |  |  | 20 pM PhiX (previously diluted) |  | -15°C to -25°C User supplied | 10 | Thaw and then keep on ice.
|  | 2 | Vortex and pulse-spin all reagents before use.
|  | 3 | Gather required consumables listed below:
1.5 mL microcentrifuge tubes
INFORMATION REQUIRED FOR 1 nM DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
Enriched Sample Pool Concentration (ng/µL)
|  |  |  | Enriched Sample Pool Fragment Size (bp) |  | 800
|  |  |  | Enriched Sample Pool Concentration (nM) |  | 0
|  |  |  | Total Volume Required (µL) |  | 100
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO 1 nM
|  | 1 | Dilute the sample pool to 1 nM by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | Resuspension Buffer |  | #DIV/0!
|  |  |  | AlloSeq Tx Enriched Sample Pool |  | #DIV/0!
|  |  |  | Total |  | #DIV/0!
|  | 2 | Vortex and pulse-spin diluted sample pool before further use.
INFORMATION REQUIRED FOR FINAL DILUTION CALCULATIONS
|  | 1 | Enter the information required below electronically, in order to calculate the appropriate dilution for the library:
|  |  |  | Information Required |  | Value
|  |  |  | Number of pools to be combined for sequencing |  | 1
|  |  |  | Loading Concentration (pM) |  | 200
|  |  |  | PhiX Spike in Concentration (%) |  | 1%
|  |  |  | Loading Volume |  | 100
DILUTE ALLOSEQ TX ENRICHMENT SAMPLE POOL TO LOADING CONCENTRATION, INCLUDING PHIX SPIKE-IN
|  | 1 | Dilute the sample pool to loading concentration by combining the following volume of reagent in a microcentrifuge tube:
|  |  |  | Reagent |  | Volume (µL)
|  |  |  | 1 nM Sample Pool # 1 |  | 20
|  |  |  | 1 nM Sample Pool # 2 |  | 0
|  |  |  | 1 nM Sample Pool # 3 |  | 0
|  |  |  | 1 nM Sample Pool # 4 |  | 0
|  |  |  | 20 pM Non-denatured PhiX |  | 10
|  |  |  | Resuspension Buffer |  | 70
|  |  |  | Total |  | 100
|  | 2 | Invert to mix (do NOT vortex) and then pulse-spin the tube.
|  | 3 | Set aside on ice until ready to be loaded onto the iSeq reagent cartridge for sequencing.
NOTE: Pools may be stored at -15°C to -25°C for up to 48 hours prior to sequencing.
|  | 4 | Proceed to load 20µL of diluted sample pool into the iSeq cartridge according to instrument protocol.
Finish time:
PROTOCOL END. PROCEED TO LOAD ISEQ ACCORDING TO INSTRUMENT PROTOCOL.

## VersionHistory
VERSION | REVISION DATE (DDMMMYY ) | REVISION DESCRIPTION
1 | 13AUG2019 | Initial release.
1.1 | 03OCT2019 | Workbook has been update with the following changes: 1) Added cover sheet to the workbook, 2) Added storage box number to the reagent preparation tab, 3) Corrected resuspension buffer preparation to bring to room temperature not thaw, 4) Minor format changes to cells (increase decimal places displayed, converted to %), 5) Unlocked data lot number and other data entry cells so operators can enter into workbook if required, 6) Minor clarifications throughout the protocol to separate values for calculations from dilution steps, and working solutions.
1.2 | 17OCT2019 | Workbook updated to include the document version number, date issued and product code on the cover sheet. Added instructions to dilute PhiX to 20pM for iSeq runs.
1.3 | 17FEB2020 | Workbook updated to add CE cover page.
2 | 18FEB2020 | Workbook updated to add CE 0197 logo.
2.1 | 30JUL2020 | Workbook updated to include all RUO AlloSeq Tx products.
2.2 | 21OCT2020 | Workbook updated to include warning message not to leave NaOH open when not in use.
| 21OCT2020 | Minor corrections included in Workbook verification tab - refer to attached worbook verification in DCR 2020-303. SC 21Oct20
3 | 18MAY2021 | Workbook updated to include "Proceed to load 20µL of diluted sample pool into the iSeq cartridge according to instrument protocol." comment in "1.3 iSeq" tab in response to Zendesk case. Added ASTX17.1(24)-B-IVD to CE Cover tab. Reissued 19Oct21
4 | 27Apr22 | Added ASTX17.1(96)-A-IVD and ASTX17.1(96)-B-IVD product codes. Issued 27Apr22.