Olerup SSP® DQB1*02

Product number:	101.213-24 – including Taq polymerase

101.213-24u – without Taq polymerase

Lot number:	98Y

Expiry date:	2018-April-01

Number of tests:	24

Number of wells per test:	23+1

Changes compared to the previous DQB1*02 Lot (55X):

| Well | 5’-primer | 3’-primer | rationale |
| --- | --- | --- | --- |
| 1 | - | Added | 3’-primer added for the DQB1*02:01:21 allele. |
| 2 | Modified | - | 5’-primer modified for improved HLA-specific amplification. |
| 14 | Added | - | 5’-primer added for the DQB1*02:54 allele. |
| 15 | - | - | Exchanged positive control primer pair for improved HLA-specific amplification. |
| 16 | Added | - | 5’-primer added for the DQB1*02:54 allele. |
| 17 | Added | Added | Primer pair added for the DQB1*02:48 allele. |
| 18 | Added | - | 5’-primer added for the DQB1*02:58N allele. |
| 19 | Added | Added | Primer pair added for the DQB1*02:59 allele. |
| 20 | Added |  | 5’-primer added for the DQB1*02:14:02 allele. |
| 21 | Added | - | 5’-primer added for the DQB1*02:58N allele. |
| 22 | Added | Added | Primer pair added for the DQB1*02:59 allele. |
| 23 | Added | Exchanged | 5’-primer added and 3’-primer exchanged for the DQB1*02:53Q allele. |
| 24 | - | - | Updated negative control. |

The number of wells is unchanged.

Allele coverage:

DQB1*02:01 to DQB1*02:61, i.e. all the currently recognized DQB1*02 alleles, will be amplified by the primers in the DQB1*021 subtyping kit; www.ebi.ac.uk/imgt/hla, 2015-August-11, release 3.21.1.

The DQB1*02 kit enables separation of the confirmed DQB1*02 alleles as listed in the IMGT/HLA database. An HLA allele is listed as confirmed by IMGT/HLA if it has been sequenced by more than a single laboratory or from multiple sources.

The DQB1*02 kit also enables identification of polymorphisms in exons outside of the region encoding the peptide binding domain and of null and alternatively expressed alleles.

The DQB1*02 kit cannot distinguish the silent mutation in the DQB1*02:01:01-02:01:23 alleles, the DQB1*02:02:01-02:02:02, the DQB1*02:07:01-02:07:02 or the DQB1*02:14:01-02:14:02 alleles.

The following DQB1*02 alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| DQB1*02:07:01-02:07:02, 02:16 | 9 | DQB1*02:20N, 02:22 | 16 |
| DQB1*02:09, 02:24 | 11 | DQB1*02:21, 02:35 | 15 |
| DQB1*02:15, 02:29 | 18 | DQB1*02:27, 02:28 | 22 |
| DQB1*02:18N, 02:34 | 14 | DQB1*02:41, 02:53Q | 23 |

1Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

Resolution in DQB1*02 homo- and heterozygotes:

Very good.

Influence on the interpretation of DQB1*02 subtypings by non-DQB1*02 alleles:

None of importance.

Modifications made due to comments from customers:

In primer mix 2, the 5’-primer was modified for improved HLA-specific amplification.