PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A-B-C Combi Tray

Prod. No:	101.702-24/24u

101.702-06/06u

Lot No:    	17K

Expiry Date:    	2012-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-A low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-A serology3 | Amplified HLA-A alleles4-6 |
| --- | --- | --- | --- | --- |
| 17 | 120 bp, 140 bp | 800 bp | A1, A36 | *01:01:01:01-01:04N, 01:06-01:55, 36:01-36:05 |
| 28 | 135 bp, 200 bp, 255 bp | 1070 bp | A2, A203, A210, A19, A28 | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:22:02, 02:24-02:35:01, 02:35:03-02:47, 02:49-02:97:02, 02:99, 02:101-02:121, 02:123-02:128, 02:130-02:199 |
| 39 | 205 bp, 235 bp | 1070 bp | A1, A3, A11, A32, A34, A36 | *01:12, 01:19, 01:21, 03:01:01:01-03:17, 03:19-03:74, 03:76, 11:25, 24:92, 32:04, 34:02-34:04, 34:07-34:08, 36:02 |
| 4 | 190 bp | 800 bp | A1, A3, A11, A30, A36 | *01:01:01:01-01:04N, 01:06-01:07, 01:09-01:11N, 01:13, 01:16N-01:18N, 01:20-01:29, 01:31N-01:55, 02:78, 02:169, 03:12, 03:18, 11:01:01-11:27, 11:29-11:59, 26:19, 30:08, 36:04 |
| 5 | 160 bp, 535 bp | 800 bp | A3, A9, A23, A24, A2403, A31, A32 | *03:30, 23:01:01-23:25, 24:02:01:01-24:11N, 24:13:01-24:15, 24:17-24:64, 24:66-24:88, 24:90N-24:99, 31:08, 32:05, 32:13 |
| 610 | 135 bp, 175 bp, 210 bp | 800 bp | A2, A9, A23, A24, A29, A80 | *23:01:01-23:25, 24:05, 24:13:02, 24:24, 29:07, 31:29, 80:01-80:02, B*18:27 |
| 7 | 175 bp, 200 bp | 1070 bp | A2, A3, A9, A23, A24, A2403, A26 | *02:17:01w-02:17:02w, 23:14, 24:02:01:01-24:11N, 24:13:01-24:13:02, 24:17-24:50, 24:54-24:56, 24:58-24:63, 24:66-24:91, 24:93, 24:95-24:99, 26:16, 33:19, 68:45 |
| 8 | 165 bp, 200 bp | 800 bp | A2, A3, A10, A11, A25, A26, A28, A32, A34, A66, A68, A69 | *01:51, 02:55, 03:24, 03:50, 11:10, 25:01:01-25:13, 26:01:01-26:06, 26:08-26:15, 26:17-26:18, 26:20-26:45, 32:15, 34:01:01-34:08, 66:01-66:13, 68:01:01-68:49N, 69:01 |
| 97 | 75 bp | 800 bp | A3, A25, A32 | *25:01:01-25:13, 32:01:01-32:02, 32:04, 32:06-32:22 |
| 107,10 | 85 bp | 1070 bp | A10, A26 | *01:51, 02:146, 26:01:01-26:02, 26:04, 26:07:01-26:18, 26:20, 26:22-26:29, 26:32-26:45, 33:13 |
| 117,8,10 | 80 bp, 175 bp, 500 bp | 1070 bp | A1, A9, A10, A11, A26, A31, A34, A66 | *01:13, 01:28, 03:63, 11:01:01-11:27, 11:29-11:59, 24:19, 24:44, 26:03:01-26:03:02, 26:06, 26:21, 31:03, 34:01:01-34:08, 66:01, 66:04-66:10, 66:13, 80:02 |
| 12 | 185 bp | 800 bp | A10, A25, A26, A31, A34, A43, A66 | *11:11, 25:06, 26:09, 31:03-31:04, 34:01:01-34:08, 43:01, 66:02-66:03 |
| 13 | 175 bp, 225 bp | 1070 bp | A1, A2, A3, A10, A25, A26, A34, A43, A66 | *01:13, 02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:103, 02:135, 03:01:01:01-03:07, 03:09-03:11N, 03:13-03:31, 03:33-03:35, 03:37-03:40, 03:42-03:56, 03:58, 03:60-03:71, 03:73-03:76, 25:01:01-25:05, 25:07-25:13, 26:01:01-26:01:13, 26:02w, 26:03:01-26:03:02, 26:05-26:08, 26:10-26:33, 26:35-26:45, 34:08, 43:01, 66:01, 66:04-66:13, 74:13 |
| 14 | 200 bp, 240 bp | 1070 bp | A2, A29 | *02:122, 29:01:01:01-29:22, 33:08-33:09, 33:13 |
| 157 | 90 bp, 135 bp, 205 bp | 1070 bp | A1, A30 | *01:13, 01:28, 03:43, 30:01:01-30:04, 30:06-30:20, 30:22-30:37 |
| 16 | 190 bp, 370 bp, 395 bp | 1070 bp | A2, A24, A31, A32 | *02:122, 29:14, 31:01:02-31:31, 32:05, 33:08-33:09, 74:04 |
| 17 | 150 bp, 180 bp | 1070 bp | A32 | *03:43, 29:13, 32:01:01-32:03, 32:05-32:22 |
| 18 | 200 bp | 1070 bp | A33, A68 | *33:01:01-33:01:03, 33:03:01-33:29, 68:29 |
| 19 | 160 bp, 200 bp | 800 bp | A74 | *29:19, 74:01-74:14N |
| 20 | 155 bp, 200 bp, 240 bp | 800 bp | A2, A210, A25, A68 | *02:10, 02:34-02:35:03, 02:39, 02:46, 02:48, 02:50, 02:52, 02:56:01-02:56:02, 02:62, 02:70, 02:73, 02:78, 02:93, 02:95, 02:103, 02:110, 02:114, 02:117, 02:122, 02:129, 02:156, 11:57, 24:94, 25:05, 68:01:01-68:49N |
| 21 | 240 bp, 375 bp | 800 bp | A2, A26, A28, A68, A69 | *02:55, 24:82, 26:22, 33:22, 66:09, 68:29, 69:01 |
| 227,11 | 85 bp, 240 bp | 800 bp | A2, A36 | *02:34-02:35:03, 02:46, 02:48, 02:56:01-02:56:02, 02:62, 02:70, 02:78, 02:103, 02:129, 36:01-36:05 |
| 237,10,12 | 70 bp, 160 bp, 240 bp, 495 bp | 800 bp | A2, A26, A28, A36, A68, A80 | *02:55, 03:41, 03:63, 03:75, 24:18, 26:03:01-26:03:02, 26:05-26:06, 26:21, 26:30, 33:24, 36:02, 68:05, 68:15, 68:20, 80:01 |
| 24 | 360 bp | 1070 bp | A1, A11, A36, A80 | *01:01:01:01-01:04N, 01:06, 01:08-01:29, 01:31N-01:35, 01:37-01:49, 01:51-01:55, 03:41, 03:63, 11:01:01-11:01:14, 11:01:16-11:11, 11:13-11:17, 11:19-11:27, 11:29-11:40, 11:42-11:59, 33:13, 36:01-36:05, 74:10, 80:01-80:02 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A low resolution typing.

In addition, wells 4, 5, 6, 8, 9, 12 and 19 to 23 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-A alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

4For several HLA-A alleles only partial 1st exon nucleotide sequences are available. We assume that unknown sequences are conserved within allelic groups.

Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

5The A*03:18 and 26:19 alleles will give rise to identical amplification patterns. These two alleles can be separated by the respective high resolution SSP primer sets.

The A*23:14, 24:05, 24:13:02 and 24:24 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

The A*26:31 and the A*66:11 and 66:12 alleles will give rise to identical amplification patterns. These three alleles can be separated by the respective high resolution SSP primer sets.

6Primer mix 6 will amplify the B*18:27 allele.

7Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

8The primer pairs in vial 2 and 11 will in most samples give rise to two to three specific PCR fragments.

9Primer mix 3 may faintly amplify the A*30:04, 30:06, 30:17 and 30:29 alleles.

10Primer mixes 6, 10, 11 and 23 may yield less specific PCR product than the other HLA-A low primer mixes.

11Primer mix 22 might faintly amplify most A*11 alleles.

12Primer mix 23 may have a tendency of primer dimer formation.

‘w’, might be weakly amplified.

1The A*03:18 and 26:19 alleles will give rise to identical amplification patterns. These two alleles can be separated by the respective high resolution SSP primer sets.

The A*23:14, 24:05, 24:13:02 and 24:24 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

The A*26:31 and the A*66:11 and 66:12 alleles will give rise to identical amplification patterns. These three alleles can be separated by the respective high resolution SSP primer sets.

The sequence of the A*01:05N has been shown to be identical to A*01:04N. The A*02:01:16 allele has been renamed to A*02:134. The A*02:01:20 has been renamed to A*02:01:18.The sequence of the A*02:23 allele has been shown to be identical to A*02:22. The sequence of the A*02:98 allele has been shown to be identical to A*02:96. The A*11:28 allele has been renamed to A*11:15:02. The sequence of the A*24:01 allele has been shown to be in error. The sequence of the A*24:12 allele has been shown to be identical to A*24:08. The A*24:16 allele has been renamed to A*31:08. The A*24:65 allele has been renamed to A*24:13:02. The sequence of the A*30:05 allele has been shown to be identical to A*30:04.

The A*30:21 allele has been renamed to A*30:11:02. The sequence of the A*31:011 allele has been shown to be identical to A*31:01:02. The sequence of the A*33:02 allele has been shown to be identical to A*33:03.

2The serological reactivity of all HLA-A alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

3The primer pairs in vials 2 and 11 will in most samples give rise to multiple specific PCR fragments.

‘w’, may be weakly amplified.

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-B low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-B serol.3 | Amplified HLA-B alleles4,5 |
| --- | --- | --- | --- | --- |
| 256,10 | 110 bp | 800 bp | 7, 703, 40, 41, 42, 48, 61 | *07:02:01-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:47, 07:49N-07:50, 07:52, 07:54-07:59, 07:61-07:99, 15:138, 37:07, 40:15-40:16, 40:23, 40:32, 40:98, 41:08, 42:05:01-42:05:02, 48:05, 48:08, 48:15, C*02:23 |
| 26 | 215 bp | 1070 bp | 8, 44 | *08:01:01-08:25, 08:27-08:49, 08:51-08:61, 15:142, 15:180, 44:49, 51:68 |
| 27 | 140 bp, 235 bp | 1070 bp | 7, 8, 13, 15, 35, 4005, 44, 49, 61, 62, 77 | *07:20, 07:24, 07:60, 08:21, 08:25, 13:01:01-13:04, 13:06-13:08Q, 13:10-13:23, 13:25-13:37, 15:07, 15:36w, 15:45, 15:55, 15:68, 15:89w, 15:126, 35:05:01-35:05:02, 35:16-35:17, 35:22, 35:30, 35:51, 35:58, 35:72, 35:89, 35:97, 40:05, 40:71, 44:08w, 44:54, 44:57w, 44:60, 46:12, 46:20, 51:64, 53:14, 58:18, C*03:05, C*03:25, C*03:27 |
| 287 | 130 bp, 265 bp | 800 bp | 12, 13, 14, 17, 21, 35, 40, 41, 44, 45, 47, 49, 50, 60, 61, 64, 65 | *13:01:01-13:04, 13:06-13:13, 13:15-13:23, 13:25-13:37, 14:01-14:04, 14:07N, 14:09, 14:11-14:12, 14:14-14:18, 15:46, 15:53, 15:106, 18:44, 35:46-35:47, 35:63, 40:01:01-40:01:10, 40:07, 40:10, 40:14:01-40:16, 40:22N-40:23, 40:25, 40:30-40:34, 40:36, 40:38, 40:42-40:43, 40:45, 40:47-40:49, 40:51-40:55, 40:58-40:63, 40:65-40:67, 40:69, 40:72:01-40:73, 40:76-40:77, 40:79-40:81, 40:84, 40:87-40:88, 40:92, 41:01-41:12, 44:02:01:01-44:03:06, 44:03:08-44:05:02, 44:09-44:39, 44:41-44:43, 44:45-44:56N, 44:58N-44:80, 44:82, 44:84-44:99, 45:01-45:11, 47:01:01:01-47:07, 49:01:01-49:02, 49:04-49:10, 50:01:01-50:02, 50:04-50:06, 50:08-50:09, C*01:30 |
| 298 | 185 bp, 235 bp | 800 bp | 7, 16, 17, 27, 2708, 37, 38, 3901, 3902, 39, 64, 67 | *14:01, 14:07N-14:08, 14:10, 14:12, 14:14, 27:01-27:06, 27:08-27:10, 27:12-27:13, 27:16-27:18, 27:20, 27:23, 27:26-27:27, 27:29, 27:31, 27:35-27:37, 27:39-27:42, 27:44-27:46, 27:48-27:61, 27:64N-27:65N, 37:02, 38:01:01-38:03, 38:07-38:22, 39:01:01:01-39:01:01:02L, 39:01:03-39:01:08, 39:01:10-39:02:01, 39:03, 39:05:01-39:09, 39:11, 39:14-39:15, 39:18, 39:19:02, 39:22, 39:24-39:44, 39:46-39:48, 39:50-39:57, 47:04-47:05, 48:21, 67:01:01 |
| 30 | 190 bp | 800 bp | 14, 35, 38, 39, 65 | *07:28, 14:02:01-14:02:02, 14:02:04-14:06:02, 14:09, 14:11, 14:13, 14:15-14:18, 15:77, 15:189, 35:26, 38:05, 39:04, 44:16, 44:37, 44:64:01-44:64:02, 44:91, 57:04, C*16:10 |
| 31 | 290 bp | 1070 bp | 15, 22, 62, 63, 71, 72, 75, 76, 77, 46 | *15:01:01:01-15:01:03, 15:01:06-15:02:04, 15:03:03-15:08, 15:10:02-15:11:03, 15:12-15:15, 15:17:01:01-15:17:02, 15:19-15:21, 15:24-15:28, 15:30-15:36, 15:38:01-15:40, 15:42-15:46, 15:48, 15:50, 15:55-15:58, 15:60, 15:63, 15:65-15:66, 15:70-15:71, 15:73, 15:75-15:79N, 15:81-15:89, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:107, 15:109-15:113, 15:116-15:118, 15:120-15:122, 15:125-15:126, 15:128-15:129, 15:135-15:150, 15:152, 15:154-15:155, 15:157, 15:159-15:160, 15:162-15:172, 15:174-15:175, 15:177-15:179, 15:181N-15:185, 15:187-15:191, 46:01:01-46:22 |
| 32 | 165 bp, 220 bp, 330 bp | 1070 bp | 5, 8, 12, 21, 22, 37, 41, 42, 44, 45, 48, 51, 56, 57, 60, 62, 70, 71, 72, 82 | *08:01:01-08:05, 08:08N-08:12:03, 08:15-08:19N, 08:21-08:24, 08:26-08:27, 08:29-08:36, 08:38-08:39, 08:41-08:48, 08:50-08:54, 08:56-08:61, 15:03:01-15:03:03, 15:09-15:10:02, 15:18:01-15:18:04, 15:23, 15:29, 15:37, 15:46-15:47, 15:49, 15:51-15:54, 15:61-15:62, 15:64, 15:69, 15:72, 15:74, 15:80, 15:83, 15:90-15:91, 15:93, 15:98-15:99, 15:103, 15:106, 15:108, 15:114-15:115, 15:119, 15:123-15:124, 15:127, 15:131-15:134, 15:143, 15:151, 15:153, 15:156, 15:158, 15:161, 15:173, 15:176, 15:180, 15:186, 35:87, 37:12, 40:12, 41:01-41:03:02, 41:05-41:09, 41:11-41:12, 42:01:01-42:02, 42:04-42:08, 42:10-42:13, 44:14-44:15, 44:18, 44:20, 44:62, 45:01-45:11, 51:08, 51:20, 51:36, 51:44N, 52:19, 55:20, 56:13, 57:09, 82:01-82:03, C*07:46 |
| 338,9,11 | 165 bp, 190 bp, 390 bp | 1070 bp | 5, 17, 21, 35, 51, 5102, 5103, 52, 56, 62, 63, 70, 78 | *15:01:02, 15:09, 15:16-15:17:02, 15:67, 15:95, 15:162, 15:168, 15:177, 35:01:10, 35:04:02, 40:26, 40:28, 44:62, 51:01:01-51:01:03, 51:01:05-51:02:03, 51:03-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03, 51:26-51:33, 51:37-51:41N, 51:43-51:44N, 51:46, 51:48-51:61, 51:63-51:80, 51:82-51:91, 51:94-51:95, 52:01:01-52:01:04, 52:02-52:13, 52:15-52:20, 56:05:01-56:06, 56:21, 58:08, 78:01-78:07 |
| 347,9 | 180 bp | 800 bp | 13, 22, 55, 60 | *07:78, 13:01:01-13:02:07, 13:07N-13:09, 13:11, 13:14-13:20, 13:22:01-13:23, 13:25, 13:27-13:34, 13:36-13:37, 40:48, 45:10, 49:07, 55:09, 55:22, 55:24 |
| 356 | 105 bp, 195 bp | 1070 bp | 8, 12, 27, 38, 39, 3902, 40, 44, 45, 48, 60, 70, 71, 72, 82 | *07:27, 07:50, 08:04, 08:17, 08:54, 15:03:01-15:03:03, 15:47, 15:49, 15:54, 15:61-15:62, 15:64, 15:68-15:69, 15:91, 15:98, 15:103, 15:123, 15:127, 15:131-15:132, 15:151, 15:156, 15:158, 15:173, 18:37, 27:18, 27:29, 38:03, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:23, 39:39, 39:49, 40:12, 40:46, 40:93, 42:11, 44:10, 44:15, 44:18, 44:40, 44:44, 45:01, 45:05-45:07, 45:11, 48:01:01-48:05, 48:07-48:22, 50:02, 52:16, 55:18, 82:01-82:03 |
| 36 | 280 bp | 1070 bp | 5, 7, 8, 13, 15, 16, 17, 18, 22, 35, 40, 42, 44, 46, 48, 49, 51, 53, 55, 56, 57, 58, 60, 61, 62, 63, 70, 71, 72, 75, 76, 77, 78 | *07:09, 07:11, 07:17, 08:28, 08:35, 08:37, 13:04, 13:10, 13:26, 15:01:01:01-15:01:04, 15:01:06-15:08, 15:11:01-15:16, 15:18:01-15:21, 15:23-15:29, 15:31-15:36, 15:38:01-15:40, 15:43-15:44, 15:46-15:47, 15:49-15:57, 15:60-15:62, 15:64-15:72, 15:74-15:76, 15:78:01-15:82, 15:84-15:85, 15:87-15:89, 15:91-15:98, 15:101-15:129, 15:131-15:132, 15:134-15:136, 15:138-15:149N, 15:151-15:161, 15:163-15:167, 15:169-15:176, 15:178-15:187, 15:189-15:191, 18:01:01-18:15, 18:17N-18:25, 18:27-18:40, 18:42-18:49, 27:41, 35:01:01-35:01:19, 35:05:01-35:05:02, 35:07-35:08:04, 35:10-35:11, 35:14:01-35:17, 35:19-35:21, 35:23-35:30, 35:32, 35:35, 35:37, 35:40N-35:43, 35:45-35:54, 35:57-35:58, 35:61-35:64, 35:66-35:69, 35:71-35:72, 35:76-35:80, 35:82, 35:86, 35:89-35:94, 35:97, 35:99, 39:07, 39:43, 40:03, 40:20, 40:38, 40:52, 40:59-40:60, 42:09, 44:17, 44:43, 45:09, 46:01:01-46:10, 46:12-46:17, 46:19-46:22, 48:02:01-48:02:02, 48:14, 49:04-49:05, 51:37, 51:45, 51:63, 53:01:01-53:03, 53:05-53:06, 53:08:01-53:18, 53:20-53:23, 54:06, 54:09, 54:14, 55:14, 55:23, 55:32, 56:03, 56:09, 56:18, 57:01:01-57:01:07, 57:06, 57:08, 57:10-57:11, 57:13-57:16, 57:18-57:27, 57:29-57:31, 58:01:01-58:02, 58:04-58:07, 58:09-58:19, 58:21-58:26 |
| 37 | 195 bp | 1070 bp | 27, 35, 37, 44, 58, 76 | *15:12, 15:14, 15:19, 15:91, 15:131, 15:161, 35:45, 35:71, 37:01:01-37:09, 37:12-37:13, 37:15-37:21, 38:17, 44:17, 44:43, 45:09, 46:17, 53:22, 58:07 |
| 386 | 105 bp, 395 bp, 435 bp | 1070 bp | 5, 7, 15, 41, 42, 62, 63, 70, 71, 75, 77 | *07:04, 07:25, 15:09-15:10:02, 15:13, 15:16-15:18:04, 15:21, 15:23-15:24, 15:37, 15:44, 15:51-15:52, 15:66-15:67, 15:72, 15:80, 15:87, 15:90, 15:93, 15:95, 15:99, 15:108, 15:114-15:115, 15:119, 15:124, 15:133-15:134, 15:153, 15:157, 15:161-15:162, 15:168, 15:176-15:177, 15:186, 15:189, 41:08, 42:05:01-42:05:02 |
| 396 | 115 bp, 150 bp | 1070 bp | 18 | *18:01:01-18:15, 18:17N-18:28, 18:30-18:36, 18:38-18:47, 18:49 |
| 406 | 80 bp | 1070 bp | 7, 27, 2708, 44, 60 | *07:73, 27:01-27:21, 27:23-27:51, 27:53-27:65N, 38:22, 40:46, 40:93, 44:40, 44:44 |
| 41 | 150 bp | 800 bp | 12, 13, 17, 18, 22, 27, 35, 37, 39, 44, 45, 47, 48, 51, 53, 56, 57, 58, 62, 70, 75, 77, 78 | *08:49, 13:01:01-13:01:03, 13:06-13:07N, 13:12-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:36, 14:10, 15:02:01-15:02:04, 15:13, 15:20-15:21, 15:25:01-15:25:02, 15:36, 15:44, 15:62, 15:77, 15:80, 15:85, 15:88-15:89, 15:106, 15:112, 15:121, 15:139, 15:144, 15:154, 15:165, 15:170, 18:22, 27:19, 27:30, 35:01:01-35:04:03, 35:06-35:08:04, 35:10-35:17, 35:19-35:21, 35:23-35:30, 35:33-35:36, 35:38-35:42:02, 35:45-35:50, 35:52, 35:54-35:57, 35:59, 35:61-35:63, 35:65Q, 35:69-35:71, 35:74, 35:76-35:78, 35:80-35:85, 35:90-35:96, 35:98, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:11, 37:13-37:22, 38:20, 39:42, 40:28, 44:02:01:01-44:14, 44:16-44:17, 44:19N, 44:21-44:30, 44:32-44:40, 44:42-44:46, 44:48-44:52N, 44:55-44:64:02, 44:66-44:98, 48:02:01-48:02:02, 51:04, 51:42, 51:46, 51:56, 53:01:01-53:13, 53:15-53:23, 55:14, 56:09, 56:11-56:12, 57:01:01-57:01:04, 57:01:06-57:10, 57:12, 57:14-57:20, 57:22-57:30, 58:01:01-58:01:02, 58:01:04, 58:04-58:05, 58:09-58:15, 58:17N, 58:19, 58:21-58:24, 58:28, 83:01 |
| 42 | 135 bp | 1070 bp | 8, 18, 22, 35, 39, 78 | *07:65, 08:32, 18:01:01-18:08, 18:10-18:15, 18:17N-18:36, 18:38-18:47, 35:01:01-35:03:02, 35:03:04-35:18, 35:20:01-35:24:02, 35:28-35:45, 35:48, 35:50-35:55, 35:57-35:62, 35:64-35:72, 35:74-35:99, 37:11, 39:19:01-39:19:02, 56:06, 78:01-78:05, 78:07 |
| 436,10 | 60 bp, 245 bp, 400 bp | 1070 bp | 5, 15, 44, 48, 60, 62, 63, 70, 71, 72, 75, 76, 77, 82 | *15:01:01:01-15:01:04, 15:01:06-15:21, 15:23-15:40, 15:42-15:58, 15:60-15:67, 15:69-15:99, 15:101-15:129, 15:131-15:191, 40:12, 44:10, 82:01-82:03 |
| 449,10 | 210 bp | 800 bp | 27, 37 | *07:27, 08:49, 37:01:01-37:22, 38:17 |
| 45 | 170 bp | 800 bp | 16, 35, 38, 39, 3901, 3902, 67, 72 | *08:55, 15:69, 15:186, 35:35, 38:01:01-38:09, 38:11-38:21, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:24, 39:26-39:29, 39:31, 39:35, 39:37-39:42, 39:44-39:46, 39:48-39:49, 39:51-39:57, 58:20, 67:01:01-67:02 |
| 466 | 110 bp | 1070 bp | 38 | *38:01:01-38:03, 38:05, 38:08-38:22 |
| 47 | 395 bp | 1070 bp | 14, 16, 38, 39, 3901, 64, 65 | *07:69, 07:85, 14:01-14:18, 38:01:01-38:02:02, 38:04-38:05, 38:08-38:22, 39:01:01:01-39:01:01:02L, 39:01:03-39:01:11, 39:03-39:07, 39:09, 39:11-39:12, 39:14-39:15, 39:18-39:19:02, 39:22, 39:24-39:38Q, 39:40N-39:48, 39:50-39:57 |
| 48 | 425 bp | 1070 bp | 7, 703, 8, 14, 15, 16, 21, 39, 3901, 3902, 40, 41, 42, 45, 48, 50, 60, 61, 64, 65, 67, 70, 73, 81 | *07:02:01-07:26, 07:28-07:35, 07:37, 07:39-07:64, 07:66-07:80, 07:82-07:99, 08:01:01-08:01:10, 08:04-08:35, 08:37-08:51, 08:53-08:61, 14:01-14:02:04, 14:03-14:18, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:57, 40:01:01-40:01:10, 40:07, 40:09?-40:11?, 40:14:01?-40:15?, 40:16, 40:18?, 40:21?-40:22N?, 40:23, 40:24?-40:38?, 40:42?-40:46?, 40:48?-40:50?, 40:51, 40:52?-40:58?, 40:59-40:60, 40:61?-40:65?, 40:66, 40:67?-40:69?, 40:71?, 40:72:01-40:73, 40:74?-40:75?, 40:77, 40:78?, 40:79, 40:80?-40:88?, 40:90?-40:93?, 40:95?, 40:98?-40:99?, 41:01-41:12, 42:01:01-42:02, 42:04-42:13, 45:01-45:11, 48:01:01-48:17, 48:19-48:22, 50:01:01-50:02, 50:04-50:09, 67:01:01-67:02, 73:01, 81:02 |
| 4912 | 175 bp | 1070 bp | 15, 40, 44, 48, 60, 62, 71 | *15:116, 15:124, 40:01:01-40:01:06, 40:01:08-40:01:10, 40:07, 40:10, 40:12, 40:21-40:23, 40:25, 40:33, 40:36, 40:38, 40:42-40:43, 40:46-40:49, 40:51-40:52, 40:54-40:55, 40:60-40:63, 40:65-40:67, 40:69, 40:73-40:74, 40:76, 40:79, 40:81, 40:84, 40:87-40:88, 40:92, 44:31, 46:06, 48:03:01-48:03:02 |
| 50 | 130 bp | 1070 bp | 21, 4005, 41, 50, 60, 61 | *18:48, 40:02:01-40:06:02, 40:08-40:09, 40:11, 40:13, 40:18-40:20, 40:24, 40:26-40:29, 40:35, 40:37, 40:39-40:40, 40:44, 40:50, 40:56-40:57, 40:64, 40:68, 40:70-40:71, 40:74-40:75, 40:78, 40:82, 40:85-40:86, 40:89-40:91, 40:94-40:99 |
| 516 | 105 bp | 1070 bp | 7, 8, 22, 41, 42 | *07:04, 07:25, 08:01:01-08:05, 08:07-08:12:03, 08:14-08:19N, 08:21-08:24, 08:26-08:39, 08:41-08:54, 08:56-08:61, 35:87, 41:01-41:08, 41:10-41:12, 42:01:01-42:02, 42:04-42:07, 42:09-42:13, 55:20 |
| 52 | 325 bp | 1070 bp | 7, 22, 27, 2708, 35, 42, 45, 46, 54, 55, 56, 67, 73, 78, 81, 82 | *07:02:01-07:02:19, 07:04-07:07, 07:09-07:15, 07:17-07:26, 07:28-07:31, 07:33-07:36, 07:39-07:46, 07:47w, 07:48-07:49N, 07:51-07:68, 07:70-07:84, 07:86-07:99, 15:76, 15:101, 27:01-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:65N, 35:76, 42:01:01-42:02, 42:04-42:06, 42:08-42:10, 42:12-42:13, 44:90, 44:97, 45:06, 46:01:01-46:01:04, 46:02w, 46:03-46:22, 54:01-54:22, 55:01:01-55:05, 55:07-55:17, 55:19-55:40, 56:01:01-56:16, 56:18-56:22, 56:24-56:29, 67:01:01-67:02, 73:01, 81:01-81:04N, 82:01-82:03, 83:01 |
| 536 | 115 bp, 195 bp, 225 bp, 260 bp | 1070 bp | 35, 37, 44, 47, 57 | *08:49, 08:60, 35:38, 35:45, 35:71, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:22, 42:13, 44:02:01:01-44:14, 44:16-44:17, 44:19N-44:36, 44:38-44:63, 44:65-44:99, 51:42, 53:22, 57:07, 57:09, 57:24, 83:01, C*15:25 |
| 54 | 215 bp | 800 bp | 5, 8, 12, 17, 21, 22, 41, 42, 44, 45, 51, 5102, 5103, 52, 53, 56, 62 | *08:09, 15:83, 41:01, 41:05-41:07, 41:09, 41:12, 42:04, 44:06, 44:15, 44:18, 44:20, 45:01-45:11, 51:01:01, 51:01:03-51:01:08, 51:01:10-51:02:01, 51:02:03-51:04, 51:06-51:07:01, 51:08-51:09:01, 51:10-51:14, 51:16-51:24:03, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:88, 51:90-51:92, 51:94-51:95, 52:01:02, 52:01:04, 52:02-52:03, 52:06:01-52:06:02, 52:09, 52:19, 53:06, 55:20, 56:13, 58:08 |
| 55 | 130 bp, 270 bp | 1070 bp | 17, 18, 22, 35, 37, 46, 51, 53, 62, 78 | *14:10, 15:57w, 18:22, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 37:04:01-37:04:02, 40:28, 46:01:01-46:22, 51:04, 51:46, 51:56, 53:02, 53:06, 57:14, 58:09 |
| 566,13 | 90 bp, 410 bp | 1070 bp | 15, 22, 27, 35, 47, 54, 55, 56, 78, 81, 82 | *07:65, 27:01w, 27:02-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:65N, 35:76, 44:90, 44:97, 47:01:01:01-47:03, 47:06-47:07, 54:01-54:22, 55:01:01-55:05, 55:07-55:17, 55:19-55:40, 56:01:01-56:16, 56:18-56:22, 56:24-56:29, 81:01, 82:01-82:03, 83:01 |
| 576 | 90 bp, 175 bp | 1070 bp | 15, 27, 35, 48, 60, 62, 75, 77, 81 | *13:36, 15:02:01-15:02:04, 15:08, 15:11:01-15:11:04, 15:13, 15:15, 15:21, 15:31, 15:44, 15:55, 15:76w, 15:88-15:89, 15:112, 15:121, 15:139, 15:144, 15:148, 15:170, 15:189, 15:191, 27:24, 35:46, 40:31, 40:45, 40:80, 48:01:01-48:01:02, 48:04, 48:06-48:07, 48:09, 48:11, 48:15-48:16, 48:18-48:20, 48:22, 55:35w, 81:01-81:04N |
| 58 | 145 bp, 430 bp | 1070 bp | 44, 49, 59, 61 | *40:13, 40:19, 44:18, 44:25, 44:50, 44:95, 49:01:01-49:01:02, 49:04-49:10, 54:12, 56:21, 59:01-59:05 |
| 596 | 120 bp, 210 bp | 800 bp | 5, 13, 15, 17, 22, 35, 45, 49, 50, 51, 5102, 5103, 52, 55, 56, 62, 63, 78 | *07:78, 07:84, 13:16, 13:31w, 15:04w, 15:16, 15:42, 15:67, 15:95, 15:137w, 40:95, 49:01:01, 49:02-49:10, 50:01:01-50:02, 50:04-50:09, 51:01:01-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:14, 51:16-51:24:03, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60, 51:61w, 51:63, 51:65-51:80, 51:82-51:95, 52:01:01-52:20, 54:20, 55:01:01-55:01:06, 55:03, 55:05, 55:09, 55:11, 55:15, 55:17, 55:21w, 55:24-55:25, 55:28-55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 56:05:01-56:06, 56:21, 56:25, 58:08, 78:01-78:07 |
| 60 | 430 bp | 1070 bp | 5, 17, 27, 44, 51, 5102, 5103, 52, 53, 57, 58, 61 | *27:02, 27:30, 27:53, 27:57, 27:62, 27:65N, 40:13, 40:19, 44:06, 44:25, 44:50, 44:95, 51:01:01-51:24:03, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:95, 52:01:01-52:19, 53:01:01-53:02, 53:04-53:08:02, 53:10, 53:14-53:23, 57:01:01-57:11, 57:13-57:31, 58:01:01-58:02, 58:04-58:16, 58:18-58:28 |
| 61 | 145 bp | 1070 bp | 12, 21, 35, 40, 4005, 41, 45, 47, 50, 60, 61 | *15:46, 15:53, 15:106, 15:143, 18:48, 35:19, 35:47, 35:63, 40:01:01-40:11, 40:14:01-40:16, 40:18, 40:20, 40:22N-40:40, 40:42-40:45, 40:48-40:75, 40:77-40:92, 40:94-40:95, 40:97-40:99, 41:01-41:04, 41:05?, 41:06-41:12, 44:09, 44:46, 44:75, 44:90, 45:01-45:11, 47:02, 47:03w, 50:01:01-50:02, 50:04-50:05, 50:07-50:09 |
| 62 | 300 bp | 1070 bp | 7, 12, 13, 15, 17, 18, 21, 27, 2708, 35, 37, 38, 39, 3902, 40, 4005, 41, 44, 45, 47, 48, 49, 50, 52, 60, 61, 62, 70, 72, 76 | *07:54, 08:17, 08:38, 08:54, 13:01:01-13:04, 13:06, 13:08Q-13:23, 13:25-13:35, 13:37, 15:01:01:01-15:01:04, 15:01:06-15:01:16, 15:03:01-15:07, 15:12, 15:14, 15:19-15:20, 15:24-15:27:03, 15:30, 15:32-15:36, 15:38:01-15:40, 15:42-15:43, 15:45-15:50, 15:53-15:54, 15:56-15:58, 15:60-15:63, 15:65-15:66, 15:68-15:71, 15:73-15:75, 15:77-15:79N, 15:81-15:87, 15:91-15:92, 15:94N, 15:96-15:98, 15:101-15:107, 15:109-15:111N, 15:113, 15:116-15:118, 15:122-15:123, 15:125-15:129, 15:131-15:132, 15:135-15:138, 15:140-15:142, 15:145-15:147, 15:150-15:152, 15:154-15:160, 15:163-15:167, 15:169, 15:171-15:175, 15:178-15:179, 15:181N-15:185, 15:187-15:188, 15:190N, 18:12, 27:01-27:21, 27:24-27:65N, 35:10, 35:13, 35:16, 35:28, 35:63, 35:69, 35:80, 37:01:01-37:07, 37:09-37:22, 38:03-38:04, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:22-39:23, 39:49, 40:01:01-40:07, 40:09-40:16, 40:18-40:24, 40:26-40:40, 40:42-40:67, 40:70-40:99, 41:01-41:12, 42:11, 44:02:01:01-44:05:02, 44:07-44:11, 44:13-44:99, 45:01-45:11, 47:01:01:01-47:07, 48:01:01-48:05, 48:07-48:15, 48:17-48:22, 49:01:01-49:10, 50:01:01-50:02, 50:04-50:09, 52:01:01-52:20, 53:17, 55:18, 55:34, 55:40, 78:05-78:06 |
| 63 | 160 bp, 330 bp | 1070 bp | 22, 44, 45, 47, 54 | *44:02:01:01-44:02:13, 44:06, 44:08-44:09, 44:11-44:12, 44:16, 44:19N-44:24, 44:27, 44:33-44:34, 44:41, 44:44, 44:48-44:49, 44:52N-44:53, 44:55, 44:59, 44:63, 44:66-44:68, 44:71-44:74, 44:80, 44:83-44:84, 44:86-44:87, 44:89-44:91, 44:93, 44:95, 44:97, 44:99, 54:01-54:22, 83:01 |
| 64 | 180 bp, 210 bp | 1070 bp | 13, 15, 18, 35, 39, 40, 54, 55, 56, 59, 78 | *07:65, 13:06, 15:42, 15:86, 18:04, 35:42:02, 35:60, 39:17, 40:58, 45:08, 46:18, 54:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:16-54:22, 55:01:01-55:03, 55:05, 55:07, 55:10-55:12, 55:15-55:16, 55:18-55:19, 55:21, 55:25-55:26, 55:29-55:31, 55:33-55:40, 56:05:01, 56:06, 56:10, 56:23, 59:01, 59:05, 78:01, 78:02:02-78:03, 78:07, C*15:02:04 |
| 65 | 180 bp | 1070 bp | 12, 15, 21, 22, 44, 45, 49, 50, 51, 56, 61, 62, 82 | *13:03, 15:73, 40:71, 44:10, 44:15, 44:18, 45:01, 45:04-45:07, 45:11, 46:11, 49:01:01-49:03, 49:06, 49:08-49:10, 50:01:01-50:02, 50:04-50:08, 51:15, 51:62, 54:03, 56:01:01-56:02, 56:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:20, 56:24-56:29, 59:04, 82:01-82:03, C*03:12, C*03:19 |
| 666,10 | 90 bp, 240 bp | 800 bp | 57 | *55:14, 57:01:01-57:31 |
| 676,13 | 90 bp | 1070 bp | 5, 13, 15, 17, 39, 55, 56, 58, 60, 61, 62, 71 | *13:11, 15:18:03, 15:73, 39:17, 40:48, 40:71, 51:62, 56:02, 56:04, 56:10, 58:01:01-58:02, 58:04-58:28, C*02:06w |
| 686,8 | 95 bp | 1070 bp | 8, 18, 22, 35, 37, 38, 39, 44, 51, 5102, 5103, 53, 78 | *07:65w, 08:32, 18:01:01-18:11, 18:13-18:15, 18:17N-18:36, 18:38-18:47, 18:49, 35:01:01-35:09:02, 35:11-35:12:02, 35:14:01-35:15, 35:17-35:18, 35:20:01-35:24:02, 35:27, 35:29-35:45, 35:48, 35:50-35:62, 35:64-35:68:02, 35:70-35:72, 35:74-35:75, 35:76w, 35:77-35:79, 35:81-35:99, 37:08, 38:06-38:07, 39:19:01-39:19:02, 44:06, 51:01:01-51:24:03, 51:26-51:46, 51:48-51:95, 53:01:01-53:16, 53:18-53:23, 56:06w, 78:01-78:04, 78:07 |
| 696 | 115 bp, 150 bp | 1070 bp | 15, 18,22, 27, 35, 39, 46, 55, 62, 72, 73, 75, 76 | *13:31, 15:01:01:01-15:01:04, 15:01:06-15:01:16, 15:04, 15:07-15:08, 15:11:01-15:12, 15:14-15:15, 15:19, 15:24, 15:26N-15:28, 15:30, 15:32, 15:34-15:35, 15:38:01-15:38:02, 15:43, 15:45-15:46, 15:50, 15:53-15:54, 15:56-15:58, 15:60, 15:63, 15:66, 15:68, 15:70-15:71, 15:73, 15:75-15:77, 15:79N, 15:81-15:82, 15:85, 15:87, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:105, 15:109-15:111N, 15:113, 15:117-15:118, 15:120, 15:122, 15:125-15:126, 15:128-15:129, 15:135, 15:137, 15:140, 15:142-15:149N, 15:152, 15:154, 15:157, 15:159-15:160, 15:163-15:167, 15:169, 15:171-15:172, 15:174-15:175, 15:178, 15:180-15:184, 15:187, 15:189-15:191, 18:19, 27:25, 35:14:01-35:14:02, 35:43-35:44, 35:62, 35:67, 35:79, 35:86, 39:18, 39:36, 46:01:01-46:02, 46:04-46:05, 46:07N, 46:09-46:10, 46:12, 46:14-46:17, 46:20, 46:22, 51:61, 54:06, 55:21, 56:03, 73:01, C*06:20 |
| 7014 | 360 bp | 1070 bp | Bw4 |  |
| 71 | 350 bp | 1070 bp | Bw6 |  |
| 7210 | 285 bp | 800 bp | 7, 703, 8, 35, 40, 41, 42, 48, 53, 60, 61, 81 | *07:02:01-07:06, 07:08-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:52, 07:54-07:59, 07:61-07:76, 07:79-07:83, 07:85-07:99, 08:01:01-08:08N, 08:10-08:11, 08:13-08:15, 08:17-08:61, 35:66, 35:87, 37:07, 40:15-40:16, 40:30-40:32, 40:34, 40:45, 40:59, 40:80, 40:98, 41:02:01-41:02:02, 41:04, 41:10-41:11, 42:01:01-42:02, 42:05:01-42:07, 42:09-42:13, 48:01:01-48:01:02, 48:05-48:12, 48:14-48:20, 48:22, 53:15, 81:01-81:04N |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B low resolution typing.

In addition, wells number 28 to 30, 34, 41, 44, 45, 54, 59, 66 and 72 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-B alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

4Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

The B*08:26 and 08:50 and the B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:47 and 39:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*18:29 and the B*35:32, 35:37, 35:53N, 35:64, 35:68:01-35:68:02 and 35:99 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*41:09 and the B*45:02 and 45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:04, 55:08, 55:13 and 55:27 and the B*56:15, 56:19N and 56:22 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5The C*01:30 allele will be amplified by primer mix 28, the C*02:06 allele will be weakly amplified by primer mix 67, the C*02:23 allele will be amplified by primer mix 25, the HLA-C*03:05, 03:25 and 03:27 alleles will be amplified by primer mix 27, the HLA-C*03:12 and 03:19 alleles will be amplified by primer mix 65, the C*06:20 allele will be amplified by primer mix 69, the C*07:46 allele will be amplified by primer mix 32, the C*15:02:04 allele will amplified by primer mix 64, the C*15:25 allele will amplified by primer mix 53 and the C*16:10 allele will be amplified by primer mix 30

6Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

7Primer mix 28 may yield less specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

8Primer mixes 29, 33 and 68 give a lower yield of specific PCR product than the other HLA-B low resolution primer mixes.

9Primer mixes 33, 34 and 44 may give rise to nonspecific amplifications, most pronounced in primer mix 33.

10Primer mixes 25, 43, 44, 66 and 72 may give rise to a primer oligomer artifact.

11The B*57 and B*58 alleles might be faintly amplified by primer mix 33.

12The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 49.

13Primer mixes 56 and 67 may generate a false positive band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

14The Bw4-associated HLA-A specificities A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 70.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the B*07:01 allele has been shown to be in error. The sequence of the B*13:05 allele has been shown to be identical to B*13:04. The B*13:24 allele has been renamed B*13:22:02The B*15:01:05 allele has been corrected and renamed B*15:120The B*15:22 allele has been renamed B*35:43. The sequence of the B*15:41 allele has been shown to be identical to B*15:39. The B*15:59 allele has been renamed B*35:44. The sequence of the B*18:16 allele has been shown to be identical to B*18:14. The sequence of the B*27:051 allele has been shown to be identical to B*27:05:02. The sequence of the B*27:22 allele has been shown to be identical to the corrected B*27:06 sequence. The B*35:73 allele has been renamed B*35:08:03. The sequence of the B*39:012 allele has been shown to be identical to B*39:01:01. The sequence of the B*39:21 allele has been shown to be identical to B*39:24.The sequence of the B*40:17 allele has been shown to be identical to B*40:16. The sequence of the B*40:41 allele has been shown to be identical to B*40:40. The sequence of the B*42:03 allele has never been assigned.The sequence of the B*44:01 allele has been shown to be identical to B*4402. The sequence of the B*50:03 allele has been shown to be identical to B*50:02. The sequence of the B*51:25 allele has been shown to be identical to B*51:22. The B*51:47 allele has been renamed B*51:09:02. The sequence of the B*55:06 allele has been shown to be identical to B*55:04. The sequence of the B*58:03 allele has never been assigned. The B*79:01 allele has been renamed B*15:18. The B*9530 allele has been renamed B*15:27:02.

The B*08:26 and 08:50 and the B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:47 and 39:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*18:29 and the B*35:32, 35:37, 35:53N, 35:64, 35:68:01-35:68:02 and 35:99 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*41:09 and the B*45:02 and 45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:04, 55:08, 55:13 and 55:27 and the B*56:15, 56:19N and 56:22 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

2The serological reactivity of all HLA-B alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

‘ser’, serological HLA specificity.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Specificities and sizes of the PCR products of the 23 primer mixes used for HLA-Cw low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA class I 3,4 alleles |
| --- | --- | --- | --- |
| 73 | 155 bp | 800 bp | *01:02:01-01:37N, B*54:18 |
| 74 | 270 bp | 800 bp | *01:10, 02:02:01-02:02:03, 02:02:05-02:36, 04:32, 06:08, 08:31, 17:01:01:01-17:06, 18:03 |
| 7511 | 280 bp | 800 bp | *02:02:01-02:02:03, 02:02:05-02:03, 02:04w, 02:05-02:13, 02:14w, 02:15-02:25Q, 02:26:02-02:36, 03:02:01-03:02:06, 03:04:01:01-03:10, 03:14-03:17, 03:19, 03:23-03:29, 03:32-03:38:02, 03:40-03:42, 03:44-03:48, 03:51, 03:54, 03:57, 03:60, 03:63-03:65, 03:70-03:74, 03:77-03:78, 03:80, 03:82, 03:84, 04:03, 04:06, 04:16, 06:03, 07:96, 15:02:01-15:09, 15:10:02-15:11, 15:13, 15:15-15:22, 15:24-15:28 |
| 7611 | 155 bp | 1070 bp | *03:02:01-03:04:06, 03:04:08-03:15, 03:17-03:40, 03:42-03:57, 03:59-03:74, 03:76-03:79, 03:81-03:84 |
| 77 | 280 bp | 800 bp | *03:03:01-03:03:11, 03:11:01-03:13, 03:20N-03:22Q, 03:30-03:31, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66, 03:67w, 03:68-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 15:12 |
| 785,10 | 125 bp, 325 bp | 800 bp | *02:02:01-02:02:03, 02:02:05-02:20, 02:22-02:25Q, 02:27-02:36, 04:01:01:01-04:01:16, 04:03-04:15:02, 04:17-04:20, 04:23-04:41, 04:43-04:62, 05:26, 07:02:09, 15:11 |
| 79 | 390 bp | 1070 bp | *05:01:01:01-05:01:13, 05:03-05:42, 08:10 |
| 809 | 130 bp, 355 bp | 800 bp | *02:06, 06:02:01:01-06:02:01:02, 06:02:03-06:16N, 06:18-06:31, 06:33-06:37, 12:03:09, 12:15, 15:02:01-15:03, 15:07-15:08, 15:10:01-15:13, 15:15-15:18, 15:21, 15:26, 15:28 |
| 816 | 245 bp, 425 bp | 800 bp | *07:01:01-07:33N, 07:35-07:99 |
| 825,7,11 | 115 bp, 165 bp, 265 bp, 390 bp | 800 bp | *08:01:01-08:31 |
| 83 | 340 bp | 1070 bp | *01:14, 02:02:01-02:02:03, 02:02:05-02:11, 02:13-02:26:02, 02:28-02:36, 03:07, 03:15, 03:45, 04:01:01:01-04:01:16, 04:03-04:10, 04:12-04:20, 04:23-04:28, 04:30-04:35, 04:37-04:54, 04:56-04:62, 05:01:01:01-05:01:13, 05:03-05:42, 06:02:01:01-06:02:01:02, 06:02:03-06:10, 06:12-06:37, 07:07, 07:09, 07:49, 07:76, 08:10, 12:04:01-12:05, 12:09, 12:21, 12:33, 14:04, 14:12, 15:02:01-15:06:02, 15:08-15:13, 15:15-15:20, 15:22-15:24, 15:26-15:28, 16:02:01-16:02:04, 16:09, 16:12, 16:19, 17:01:01:01-17:06, 18:01-18:03 |
| 845 | 100 bp, 160 bp, 220 bp | 1070 bp | *01:04, 01:21, 12:02:01-12:03:01:02, 12:03:03-12:03:07, 12:03:09-12:03:11, 12:04:02-12:08, 12:10:01-12:13, 12:14:02-12:25, 12:27-12:32, 12:34-12:35 |
| 8510,11 | 130 bp, 250 bp | 800 bp | *01:21, 02:12, 04:01:01:01-04:01:16, 04:03-04:09N, 04:12-04:20, 04:23-04:35, 04:37-04:54, 04:56-04:62, 05:42, 07:02:09, 08:05, 08:21, 08:25, 12:02:01-12:03:03, 12:03:05-12:03:08, 12:03:10-12:03:11, 12:04:02, 12:06-12:08, 12:10:01-12:20, 12:22-12:32, 12:34-12:35, 15:03, 15:16, 15:25, 16:01:01-16:02:04, 16:06-16:21, B*67:02 |
| 8612 | 160 bp, 220 bp | 800 bp | *01:04, 01:09, 02:05, 02:17, 04:42, 06:02:01:01-06:02:01:02, 06:02:03-06:03, 06:07-06:13, 06:15-06:34, 06:36-06:37, 12:03:01:01-12:07, 12:11-12:13, 12:15, 12:19, 12:23, 12:25-12:26, 12:28-12:29, 12:31-12:35, 14:16, 16:04:01 |
| 8711 | 130 bp, 255 bp | 1070 bp | *02:02:01w, 02:02:02-02:02:03, 02:02:05-02:13, 02:15-02:26:02, 02:28-02:36, 03:07, 03:10, 03:15, 03:29, 03:45, 04:03, 04:06, 04:16, 04:37, 05:01:01:01-05:01:13, 05:03-05:19, 05:21-05:42, 06:03, 08:10, 12:04:01-12:05, 12:09, 12:21, 12:33, 15:02:01-15:06:02, 15:08-15:13, 15:15-15:22, 15:24, 15:26-15:28, 16:02:01-16:02:04, 16:09, 16:12, 16:19, 17:01:01:01-17:06 |
| 885,10,11 | 95 bp, 255 bp | 1070 bp | *04:11, 04:29, 04:36, 04:55, 05:11, 05:17, 05:27, 07:64, 08:04, 08:13, 14:02:01-14:11, 14:13-14:19 |
| 895,8,11 | 110 bp, 325 bp | 1070 bp | *02:06, 03:81, 05:36, 12:08, 12:15, 15:02:01-15:13, 15:15-15:24, 15:26-15:28, 16:20 |
| 9010 | 190 bp, 240 bp | 1070 bp | *02:13, 02:18, 02:33, 04:01:01:01-04:01:16, 04:03-04:10, 04:12-04:20, 04:23-04:32, 04:34-04:62, 05:17, 05:25, 05:42, 06:05, 06:31, 07:02:09, 07:31, 08:01:01-08:01:02, 08:03:01-08:03:02, 08:06, 08:08-08:11, 08:14, 08:16, 08:20-08:22, 08:24, 08:26N, 08:28, 12:14:01-12:14:02, 12:28, 14:10, 14:15, 15:12, 15:25, 16:01:01-16:02:04, 16:04:01, 16:06-16:21 |
| 919 | 225 bp, 270 bp | 800 bp | *04:11, 04:33, 04:58, 05:23, 08:07, 12:14:01-12:14:02, 14:17, 15:25, 17:01:01:01-17:06 |
| 92 | 220 bp, 425 bp | 800 bp | *01:02:01-01:03, 01:06-01:08, 01:10-01:20, 01:23-01:34, 01:37N, 03:58, 04:37, 05:16, 06:05-06:06, 06:17, 07:07, 07:09, 07:49, 07:76, 08:12, 12:09, 12:24, 14:02:01-14:05, 14:07N, 14:10-14:14, 14:17-14:19, 18:01-18:03, B*14:03 |
| 93 | 325 bp, 380 bp | 1070 bp | *01:03, 01:24, 01:34, 02:22, 03:03:01-03:04:19, 03:06-03:12, 03:14, 03:18-03:24, 03:26, 03:28-03:32, 03:34, 03:37-03:59, 03:61-03:70, 03:72-03:83, 04:01:01:01-04:01:15, 04:03-04:20, 04:24-04:53, 04:55-04:62, 05:01:01:01-05:01:13, 05:03, 05:05-05:21, 05:23-05:42, 06:09, 06:14, 06:35, 07:10, 07:28, 07:41, 07:43, 08:01:01-08:08, 08:10, 08:12-08:31, 12:31, 15:02:01-15:13, 15:15-15:21, 15:23-15:28, 17:01:01:01-17:06, 18:01-18:03 |
| 9410 | 135 bp | 1070 bp | *03:02:01-03:17, 03:19-03:38:02, 03:40-03:66, 03:67w, 03:68-03:84, 04:32, 06:03, 07:96, 18:03 |
| 95 | 160 bp, 235 bp | 1070 bp | *04:42, 06:02:01:01-06:02:01:02, 06:02:03-06:02:05, 06:04, 06:06-06:37, 07:01:01-07:02:07, 07:02:09-07:25, 07:27:01-07:32N, 07:35-07:38, 07:41-07:63, 07:65-07:91, 07:93-07:95, 07:97-07:99, 12:16, 16:01:01-16:02:04, 16:06-16:21, 18:01-18:03 |
| 9613 | Negative control |  | - |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-C low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 73 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-C low resolution typing.

In addition, wells number 74, 75, 77, 78, 80, 81, 82, 85, 86, 91 and 92 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-C alleles only partial 1st and 4th exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the HLA-C low resolution SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

4Due to sharing of sequence motifs, primer mix 73 will amplify the B*54:18 allele, primer mix 85 will amplify the B*67:02 allele and primer mix 92 will amplify the B*14:03 allele.

The C*01:05, 01:22, 01:35 and 01:36 and the B*54:18 alleles give rise to identical amplification patterns with the HLA-C low resolution primer set. These alleles are separated by the HLA-B low primer set.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Primer mix 81 will for most C*07 alleles give rise to two specific PCR fragments.

7Primer mix 82 will for most C*08 alleles give rise to multiple specific PCR fragments.

8Primer mix 89 will for most C*15 alleles give rise to two specific PCR fragments.

9Primer mixes 80 and 91 may give rise to unspecific amplification.

10Primer mixes 78, 85, 88, 90 and 94 have a tendency of giving rise to primer oligomer artifacts.

11Primer mixes 75, 76, 82, 85, 87, 88 and 89 yield somewhat less intense specific PCR fragments than the other HLA-C low resolution primer mixes.

12Primer mix 86 might faintly amplify the C*01:02 to 01:03, the C*01:05 to C*01:19 and the C*14:02:01 to 14:09 alleles.

13Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘w’, might be weakly amplified.

1The sequence of the C*01:01 allele has been shown to be identical to C*01:02.

The sequence of the C*02:01 allele has been shown to be identical to C*02:022.

The sequence of the C*02:02:04 allele has been shown to be identical to C*02:10.

The sequence of the C*02:16:03 allele has been shown to be in error.

The sequence of the C*03:01 allele has been shown to be identical to C*03:04.

The sequence of the C*04:02 allele has been shown to be identical to C*04011.

The sequence of the C*04:21 allele has been shown to be identical to C*04:15:02.

The sequence of the C*04:22 allele has been shown to be identical to C*04:21.

The sequence of the C*05:02 allele has been shown to be identical to C*05:09.

The sequence of the C*06:01 allele has been shown to be identical to C*06:02.

The sequence of the C*06:02:02 allele has been shown to be identical to C*06:17.

The sequence of the C*07:34 allele has been shown to be identical to C*07:27:02.

The sequence of the C*11:01 allele has been shown to be in error.

The sequence of the C*12:01 allele has been shown to be identical to C*12:022.

The sequence of C*13:01 has been shown to be in error.

The sequence of the C*14:01 allele has been shown to be identical to C*14:02.

The sequence of the C*15:01 allele has been shown to be identical to C*15:02.

The sequence of the C*15:14 allele has been shown to be identical to C*15:10:02.

The sequence of the C*16:03 allele has been shown to be identical to C*14:03.

The sequence of the C*16:042 allele has been shown to be identical to C*16:041.

The sequence of the C*16:05 allele has been shown to be identical to C*16:041.

The C*01:05, 01:22, 01:35 and 01:36 and the B*54:18 alleles give rise to identical amplification patterns with the HLA-C low resolution primer set. These alleles are separated by the HLA-B low primer set.

Change in revision R01 compared to R00:

- The HLA-C*03:12 and 03:19 alleles are amplified in primer mix 65.

Change in revision R02 compared to R01:

- The HLA-C*04:03, 04:06 and 04:16 alleles are amplified in primer mix 87.

Change in revision R03 compared to R02:

- Primer mix 6 may give a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.