Gel Documentation Form and Worksheet

HLA-A-B-C Combi Tray	Lot No: 3D7	Expiry Date: 2018-June-01

(101.702-24/06 – 24u/06u)

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):__________

Test  Date:_______________

Tested By:

Review Date:_________________

Reviewed By:_________________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

HLA-A low  resolution

HLA-B low  resolution

HLA-C low  resolution

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-A low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

The primer pairs in wells 1, 2, 5, 7, 8, 10, 11, 14, 15 and 18 will in many samples give rise to two or three HLA-specific PCR fragments.

Primer mix 3 may faintly amplify the A*30:04:01-30:04:02, 30:06, 30:17, 30:29, 30:46 and 30:77 alleles.

Primer mix 6 may give rise to a lower yield of A*23 alleles than the other A low primer mixes.

Primer mix 9 may weakly amplify the A*34 alleles.

Primer mixes 6, 10, 14, 20 and 23 have a tendency to giving rise to primer oligomer formation.

Primer mix 22 might faintly amplify most A*11 alleles.

Primer mixes 15 and 24 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.

Primer mix 15 may have tendency of unspecific amplification.

In primer mix 11 the positive control band may be weaker than for other HLA-A low primer mixes.

Primer mix 21 may generate a false positive band of about 560 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings. A footnote has been added in the Specificity Table.

HLA-B low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 28 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

Primer mixes 29, 33 and 63 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes.

Primer mixes 33, 52 and 62 may have tendencies of unspecific amplifications, most pronounced in primer mix 33.

Primer mixes 25, 27, 38, 41, 43, 44, 50, 58, 67 and 69 have a tendency to giving rise to primer oligomer formation.

The B*57 and B*58 alleles might be faintly amplified by primer mix 33.

The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 49.

Primer mixes 56 and 67 may generate a false positive band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

Primer mix 63 may give rise to a lower yield of B*54 alleles than the other B low primer mixes.

The Bw4-associated HLA-A specificities A9, A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 70.

HLA-C low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 74 will for most C*02 alleles give rise to two specific PCR fragments.

Primer mix 76 will for most C*03 alleles give rise to two specific PCR fragments.

Primer mix 81 will for most C*07 alleles give rise to two specific PCR fragments.

Primer mix 82 will for most C*08 alleles give rise to multiple specific PCR fragments.

Primer mix 89 will for most C*15 alleles give rise to two specific PCR fragments.

Primer mixes 84, 85 and 91 may have tendencies of unspecific amplifications.

Primer mixes 88 and 90 have a tendency to giving rise to primer oligomer formation.

Primer mixes 75, 78, 82, 87, 94 and 95 may yield less HLA-specific PCR fragments than the other HLA-C low resolution primer mixes.

Primer mix 86 might faintly amplify most C*01 and the C*14 alleles.

Primer mix 91 might generate a false band of about 500 base pairs. This band should be disregarded when interpreting HLA-C low resolution typings.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

HLA-A low resolution Interpretation Table

1HLA-A, HLA-B and HLA-C alleles listed on the IMGT/HLA web page 2015-August-11, release 3.21.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-A alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4The A*01:136 and the A*11:94, 11:112 and 11:211 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mixes 1 and 11.

The A*03:88 and the A*11:130 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mixes 11.

The A*11:199:01 and 11:222 and the A*66:23 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 11.

The A*24:14-24:15, 24:51-24:53, 24:57, 24:64, 24:114, 24:222N, 24:291 and 24:296 and the C*04:01:03 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 5.

5The HLA-A alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns.

These alleles can be separated by the full A-B-C primer set or by respective high resolution SSP primer sets.

| Alleles | Alleles |
| --- | --- |
| A*03:152, A*24:92 | A*31:08, A*33:53 |
| A*23:14:01-23:14:02, A*24:24, 24:71 | A*32:01:18, B*07:81, B*08:52, B*18:67, B*38:41, B*51:185, B*53:05, 53:16, 53:33 |
| A*23:57, 23:66, A*24:94, 24:138, 24:188, 24:228 |

‘w’, might be weakly amplified.

HLA-B low resolution Interpretation Table

1HLA-A, HLA-B and HLA-C alleles listed on the IMGT/HLA web page 2015-August-11, release 3.21.1, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-B alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

page http://hla.alleles.org/alleles/deleted.html.

4The HLA-B alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns.

These alleles can be separated by the full A-B-C primer set or by respective high resolution SSP primer sets.

| Alleles | Alleles |
| --- | --- |
| B*07:202, B*81:02 | B*41:06, 41:09, 41:29, 41:32, B*45:02-45:03 |
| B*08:26:01-08:26:02, 08:50, 08:62, 08:85, 08:94, B*42:07 | B*53:30, 53:39, B*57:45, 57:51, 57:69 |
| B*13:21, 13:35, 13:59, 13:71, B*44:135, 44:158, 44:184 | B*55:04, 55:08, 55:13, 55:27, 55:46, 55:49, B*56:01:05, 56:15, 56:19N, 56:22 |
| B*14:08:01-14:08:02, B*39:01:19, 39:25N, 39:30, 39:32-39:34, 39:47, 39:50, 39:74, 39:82, 39:102 | B*55:23, 55:32, 56:18, 56:31-56:32 |
| B*18:29, 18:72, 18:92, 18:102, B*35:32:01-35:32:02, 35:37, 35:53N, 35:64:01-35:64:02, 35:68:01-35:68:02, 35:99, 35:118-35:119, 35:174 | B*57:67:01, B*58:36 |

‘?’, nucleotide sequence information not available for the primer matching sequence.

‘w’, might be weakly amplified.

HLA-C low resolution Interpretation Table

1HLA-A, HLA-B and HLA-C alleles listed on the IMGT/HLA web page 2015-August-11, release 3.21.1. www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

page http://hla.alleles.org/alleles/deleted.html.

3The HLA-C alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns.

These alleles can be separated by the full A-B-C primer set or by respective high resolution SSP primer sets.

Alleles

C*01:05, 01:07:02, 01:22, 01:35-01:36, 01:79:01-01:79:02, B*39:76, B*54:18

C*01:67, C*14:58

C*12:18:02, B*67:02

C*14:06, 14:08-14:09, 14:28:01-14:28:02, 14:53, 14:63, A*30:62, B*35:252

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Changes in revision R01 compared to R00:

- Primer mix 34 amplifies the C*02:06 and C*15:42 alleles and primer mix 65 amplifies the B*35:291 and the C*15:15 alleles. This has been corrected in the Specificity and Interpretation Tables.

- Primer mix 21 may generate a false positive band of about 560 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings. A footnote has been added in the Specificity Table.

Change in revision R02 compared to R01:

- Primer mix 51 does not amplify the B*08:31, 41:06, 41:15 and 42:14 alleles, and primer mix 72 does not amplify the B*08:31 allele. This has been corrected in the Specificity and Interpretation Tables.