PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A-B-DR Combi Tray

Prod. No:	101.701-24/24u

101.701-06/06u

Lot No:    	23G

Expiry Date:    	2011-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-A low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-A serology3 | Amplified HLA-A alleles4,10,11 |
| --- | --- | --- | --- | --- |
| 15 | 120, 145 bp | 800 bp | A1, A36 | *01010101-0104N, 0106-0135, 0137, 3601-3604 |
| 26 | 135, 200, 255 bp | 1070 bp | A2 | *02010101-020115, 020117-020119, 020121-022202, 0224-023501, 023503-0247, 0249-0297, 0299, 9201-9221, 9223-9228, 9230-9257, 9269-9271, 9277-9279 |
| 312 | 205, 235 bp | 1070 bp | A1, A3, A11, A24, A32, A34, A36 | *0112, 0119, 0121, 03010101-0317, 0319-0347, 0350, 0351, 0355, 1125, 2492, 3204, 3402-3404, 3407, 3408, 3602 |
| 4 | 190 bp | 800 bp | A1, A2, A3, A11, A26, A30, A36 | *01010101-0104N, 0106, 0107, 0109-0111N, 0113, 0116N-0118N, 0120-0129, 0131N-0135, 0137, 0278, 9269, 0312, 0318, 110101-1127, 1129-1144, 2619, 3008, 3604 |
| 5 | 160, 535 bp | 800 bp | A3, A23, A24, A31, A32 | *0330, 230101-2321, 24020101-240217, 240220-2411N, 241301-2415, 2417-2464, 2466-2488, 2490N-2499, 3108, 3205, 3213 |
| 6 | 165, 200 bp | 800 bp | A2, A23, A24, A29, A80 | *9222, 230101-2321, 2405, 241302, 2424, 2907, 8001, B*1827 |
| 77 | 175, 200 bp | 1070 bp | A2, A23, A24, A26, A33, A68 | *021701w, 021702w, 2314, 24020101-240217, 240220-2411N, 241301, 241302, 2417-2450, 2454-2456, 2458-2463, 2466-2491, 2493, 2495-2499, 2616, 3319, 6845 |
| 8 | 165, 200 bp | 800 bp | A2, A3, A11, A25, A26, A32, A34, A66, A68, A69 | *0255, 0324, 0350, 1110, 250101-2508, 260101-2606, 2608-2615, 2617, 2618, 2620-2638, 2640, 3215, 340101-3408, 6601-6610, 680101-6848, 6901 |
| 95 | 75 bp | 800 bp | A25, A32 | *250101-2508, 320101-3202, 3204, 3206-3219N |
| 105,7 | 85 bp | 1070 bp | A2, A26, A33 | *9246, 260101-2602, 2604, 260701-2618, 2620, 2622-2629, 2632-2638, 2640, 3313 |
| 115,7,8 | 80, 175, 500 bp | 1070 bp | A1, A11, A24, A26, A31, A34, A66 | *0113, 0128, 110101-1127, 1129-1144, 2419, 2444, 2603, 2606, 2621, 3103, 340101-3408, 6601, 6604-6610 |
| 12 | 185 bp | 800 bp | A11, A25, A26, A31, A34, A43, A66 | *1111, 2506, 2609, 3103, 3104, 340101-3408, 4301, 6602, 6603 |
| 13 | 175, 225 bp | 1070 bp | A1, A2, A3, A25, A26, A34, A43, A66, A74 | *0113, 0234-023503, 025601, 025602, 0262, 9203, 9235, 03010101-0307, 0309-0311N, 0313-0331, 0333-0335, 0337-0340, 0342-0347, 0350, 0351, 0355, 250101-2505, 2507, 2508, 260101-260108, 2602w, 2603, 2605-2608, 2610-2633, 2635-2638, 2640, 3408, 4301, 6601, 6604-6610, 7413 |
| 14 | 200, 240 bp | 1070 bp | A29, A33 | *29010101-2919, 3313 |
| 155 | 90, 135, 205 bp | 1070 bp | A1, A3, A30 | *0113, 0128, 0343, 300101-3004, 3006-3020, 3022-3028 |
| 16 | 370, 395 bp | 1070 bp | A31, A32 | *2914, 310102-3125, 3205 |
| 17 | 150, 180 bp | 1070 bp | A29, A32 | *0343, 2913, 320101-3203, 3205-3219N |
| 18 | 200 bp | 1070 bp | A33, A68 | *330101, 330301-3325, 6829 |
| 19 | 160, 200 bp | 800 bp | A29, A74 | *2919, 7401-7414N |
| 20 | 155, 200, 240 bp | 800 bp | A2, A210, A24, A25, A30, A68 | *0210, 0234-023503, 0239, 0246, 0248, 0250, 0252, 025601, 025602, 0262, 0270, 0273, 0278, 0293, 0295, 9203, 9210, 9214, 9217, 9222, 9229, 9256, 2494, 2505, 680101-6848 |
| 21 | 240, 375 bp | 800 bp | A2, A24, A26, A68, A69 | *0255, 2482, 2622, 3322, 6609, 6829, 6901 |
| 225,9 | 85, 240 bp | 800 bp | A2, A30, A36 | *0234-023503, 0246, 0248, 025601, 025602, 0262, 0270, 0278, 9203, 9229, 3601-3604 |
| 235,7 | 75, 160 bp | 800 bp | A2, A26, A68, A80 | *0255, 2603, 2605, 2606, 2621, 2630, 3324, 6805, 6815, 6820, 8001 |
| 24 | 360 bp | 1070 bp | A1, A3, A11, A33, A36, A74, A80 | *01010101-0104N, 0106, 0108-0129, 0131N-0135, 0137, 0341, 110101-1111, 1113-1117, 1119-1127, 1129-1140, 1142-1144, 3313, 3601-3604, 7410, 8001 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A low resolution typing.

In addition, wells number 4, 5, 6, 8, 9, 12 and 19 to 23 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-A alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

4For several HLA-A alleles only partial 1st exon nucleotide sequences are available. We assume that unknown sequences are conserved within allelic groups.

Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6The primer pairs in vial 2 will in most samples give rise to two specific PCR fragments.

7Primer mixes 7, 10, 11 and 23 may yield less specific PCR product than the other primer mixes.

8The primer pairs in vial 11 will in most samples give rise to two to three specific PCR fragments.

9Primer mix 22 might faintly amplify most A*11 alleles.

10The A*0318 and A*2619 alleles will give rise to identical amplification patterns. These two alleles can be separated by the respective high resolution SSP primer sets.

The A*2314, A*2405, A*241302 and A*2424 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

11Primer mix 6 will amplify the B*1827 allele.

12Primer mix 3 may faintly amplify the A*30:04, 30:06, 30:17 and 30:29 alleles.

‘w’, might be weakly amplified.

1The sequence of the A*0105N has been shown to be identical to A*0104N.

The A*020116 allele has been renamed to A*9234.

The A*020120 has been renamed to A*020118.

The sequence of the A*0223 allele has been shown to be identical to A*0222.

The sequence of the A*0298 allele has been shown to be identical to A*0296.

The A*1128 allele has been renamed to A*111502.

The  sequence of the A*2401 allele has been shown to be in error.

The sequence of the A*2412 allele has been shown to be identical to A*2408.

The A*2416 allele has been renamed to A*3108.

The A*2465 allele has been renamed to A*241302.

The sequence of the A*3005 allele has been shown to be identical to A*3004.

The A*3021 allele has been renamed to A*301102.

The sequence of the A*31011 allele has been shown to be identical to A*310102.

The sequence of the A*3302 allele has been shown to be identical to A*3303.

2The serological reactivity of all HLA-A alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

3The A*0318 and A*2619 alleles will give rise to identical amplification patterns. These two alleles can be separated by the respective high resolution SSP primer sets.

4The A*2314, A*2405, A*241302 and A*2424 alleles will give rise to identical amplification patterns. These two alleles can be separated by the respective high resolution SSP primer sets.

5The primer pairs in vials 2 and 11 will in most samples give rise to two specific PCR fragments.

‘w’, may be weakly amplified.

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-B low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-B serol.3 | Amplified HLA-B alleles4,5 |
| --- | --- | --- | --- | --- |
| 256 | 110 bp | 800 bp | 7, 15, 37, 40, 41, 42, 48 | *070201-070207, 070209-071802, 0720-0732, 0734-0739, 0741-0747, 0749N, 0750, 0752, 0754-0759, 0761-0772, 0775-0777, 9538, 3707, 4015, 4016, 4023, 4032, 4098, 4108, 420501, 420502, 4805, 4808, 4815, Cw*0223 |
| 26 | 215 bp | 1070 bp | 8, 15, 44 | *080101-0825, 0827-0838, 0840, 9542, 4449 |
| 27 | 140, 235 bp | 1070 bp | 7, 8, 13, 15, 35, 40, 4005, 44, 46, 53, 57, 58 | *0720, 0724, 0760, 0821, 0825, 130101-1304, 1306-1308Q, 1310-1323, 1325, 1326, 1507, 1536w, 1545, 1555, 1568, 1589w, 9526, 3505, 3516, 3517, 3522, 3530, 3551, 3558, 3572, 3589, 3597, 4005, 4071, 4408w, 4454, 4457w, 4460, 4612, 5164, 5314, 5818, Cw*0305, Cw*0325, Cw*0327 |
| 287 | 130, 265 bp | 1070 bp | 13, 14, 15, 35, 40, 41, 44, 45, 47, 49, 50 | *130101-1304, 1306-1313, 1315-1323, 1325, 1326, 1401-1404, 1407N, 1409, 1546, 1553, 9506, 3546, 3547, 3563, 400101-400106, 4007, 4010, 401401-4016, 4022N, 4023, 4025, 4030-4034, 4036, 4038, 4042, 4043, 4045, 4047-4049, 4051-4055, 4058-4063, 4065-4067, 4069, 407201-4073, 4076, 4077, 4079-4081, 4084, 4087, 4088, 4092, 4101-4108, 44020101-440205, 440208-4405, 4409-4439, 4441-4443, 4445-4456N, 4458N-4467, 4469, 4501-4509, 47010101-4705, 490101-4902, 4904, 4905, 5001, 5002, 5004 |
| 298 | 185, 235 bp | 800 bp | 14, 16, 27, 37, 38, 39, 47, 64, 67 | *1401, 1407N, 1408, 1410, 2701-2706, 2708-2710, 2712, 2713, 2716-2718, 2720, 2723, 2726, 2727, 2729, 2731, 2735-2737, 2739-2742, 2744-2746, 2750-2752, 3702, 380101-3803, 3807-3819, 39010101, 39010102L, 390103-390201, 3903, 3905-390602, 3907?, 3908, 3909, 3911, 3914, 3915, 3918, 391902, 3922, 3924-3944, 3946-3948, 3950, 4704, 4705, 670101 |
| 308 | 185 bp | 800 bp | 14, 35, 38, 39, 65 | *140201-140602, 1409, 3526, 3805, 3904, 3907? |
| 31 | 290 bp | 1070 bp | 15, 62, 63, 75, 76, 77, 46 | *15010101-150103, 150106-1502, 1504-1508, 151101-1515, 15170101-151702, 1519-1521, 1524-1528, 1530-1536, 1538-1540, 1542-1546, 1548, 1550, 1555-1558, 1560, 1563, 1565, 1566, 1570, 1571, 1573, 1575-1579N, 1581-1589, 1592, 1594N, 1596, 1597, 9501, 9502, 9504-9507, 9509-9513, 9516-9518, 9520-9522, 9525, 9526, 9528, 9529, 9535-9550, 9552, 9554, 9555, 9557, 9559, 9560, 9563, 9564, 460101-4619, 4621 |
| 32 | 165, 220, 330 bp | 1070 bp | 8, 15, 35, 37, 40, 41, 42, 44, 45, 51, 56, 57, 70, 71, 72 | *080101-0805, 0808N-081201, 0815-0819N, 0821-0824, 0826, 0827, 0829-0836, 0838, 150301, 1509, 1510, 1518, 1523, 1529, 1537, 1546, 1547, 1549, 1551-1554, 1561, 1562, 1564, 1569, 1572, 1574, 1580, 1583, 1590, 1591, 1593, 1598, 1599, 9503, 9506, 9508, 9514, 9515, 9519, 9523, 9524, 9527, 9531-9534, 9543, 9551, 9553, 9556, 9558, 9561, 3587, 3712, 4012, 4101-410302, 4105-4108, 4201, 4202, 4204-4208, 4210-4212, 4414, 4415, 4418, 4420, 4462, 4501-4509, 5108, 5120, 5136, 5144N, 5520, 5613, 5709, 8201, 8202, Cw*0746 |
| 339,12 | 165, 390 bp | 1070 bp | 15, 35, 40, 44, 51, 52, 56, 58, 62, 63, 70, 78 | *150102, 1509, 1516-151702, 1567, 1595, 350402, 4026, 4028, 4462, 510101-510103, 510105-510203, 5103-510902, 5111N, 5112, 511302, 5114, 5117-5120, 5122, 5123, 5126-5130, 5132, 5133, 5137-5141N, 5143, 5144N, 5146, 5148-5161, 5163-5165, 520101-5207, 5209-5213, 560501-5606, 5621, 5808, 7801-7806 |
| 349,10 | 180, 395 bp | 800 bp | 13, 15, 40, 55, 62, 70, 71, 72, 75 | *130101-130205, 1307N-1309, 1311, 1314-1320, 132201-1323, 1325, 1509, 1510, 1518, 1521, 1523, 1537, 1544, 1551, 1552, 1566, 1572, 1580, 1590, 1593, 1599, 9508, 9514, 9515, 9519, 9524, 9533, 9534, 9553, 9561, 4048, 5509, 5522, 5524 |
| 356 | 105, 195 bp | 1070 bp | 7, 8, 15, 27, 38, 39, 40, 44, 45, 48, 55, 72 | *0727, 0750, 0804, 0817, 150301, 1547, 1549, 1554, 1561, 1562, 1564, 1568, 1569, 1591, 1598, 9503, 9523, 9527, 9531, 9532, 9551, 9556, 9558, 9561, 2718, 2729, 2750, 3803, 390201, 390202, 3908, 391301, 391302, 3923, 3939, 3949, 4012, 4046, 4093, 4211, 4410, 4415, 4418, 4440, 4444, 4501, 4505-4507, 480101-4805, 4807-4819, 5002, 5518, 8201, 8202 |
| 36 | 280 bp | 1070 bp | 7, 8, 13, 15, 18, 27, 35, 39, 40, 42, 44, 46, 48, 49, 51, 53, 54, 55, 56, 57, 58, 62, 63, 70, 71, 72, 75, 76, 77 | *0709, 0711, 0717, 0828, 0835, 0837, 1304, 1310, 1326, 15010101-150104, 150106-1508, 151101-1516, 1518-1521, 1523-1529, 1531-1536, 1538-1540, 1543, 1544, 1546, 1547, 1549-1557, 1560-1562, 1564-1572, 1574-1576, 157801-1582, 1584, 1585, 1587-1589, 1591-1598, 9501-9529, 9531, 9532, 9534-9536, 9538-9549N, 9551-9561, 9563, 9564, 180101-1815, 1817N-1825, 1827-1833, 1835, 2741, 350101-350107, 3505, 3507-350803, 3510, 3511, 351401-3517, 3519-3521, 3523-3530, 3532, 3535, 3537, 3540N-3543, 3545-3554, 3557, 3558, 3561-3564, 3566-3569, 3571, 3572, 3576-3580, 3582, 3586, 3589-3594, 3597, 3599, 3907, 3943, 4003, 4020, 4038, 4052, 4059, 4060, 4209, 4417, 4443, 4509, 460101-4610, 4612-4617, 4619, 4621, 4802, 4814, 4904, 4905, 5137, 5145, 5163, 530101-5303, 5305, 5306, 5308-5318, 5406, 5409, 5414, 5514, 5523, 5532, 5603, 5609, 5618, 570101-570103, 5706, 5708, 5710, 5711, 5713-5716, 5718, 5719, 580101-5802, 5804-5807, 5809-5819, 5821 |
| 37 | 195 bp | 1070 bp | 15, 35, 37, 38, 44, 45, 58, 76 | *1512, 1514, 1519, 1591, 9531, 9561, 3545, 3571, 370101-3709, 3712, 3713, 3817, 4417, 4443, 4509, 4617, 5807 |
| 386 | 105, 435 bp | 1070 bp | 7, 15, 41, 42, 59, 62, 63, 77 | *0704, 0725, 1513, 1516-151702, 1523, 1524, 1567, 1587, 1595, 9557, 4108, 420501, 420502 |
| 396 | 115 bp | 1070 bp | 18 | *180101-1815, 1817N-1828, 1830-1833, 1835 |
| 406 | 80 bp | 1070 bp | 27, 40, 44 | *2701-2721, 2723-2746, 2750, 2751, 4046, 4093, 4440, 4444 |
| 41 | 150 bp | 800 bp | 13, 15, 18, 27, 35, 37, 39, 40, 44, 48, 51, 53, 55, 56, 57, 58, 62, 75, 77 | *130101, 130102, 1306, 1307N, 1312, 1313, 1317, 1320, 132201-1323, 1325, 1326, 1410, 1502, 1513, 1520, 1521, 152501, 1536, 1544, 1562, 1577, 1580, 1585, 1588, 1589, 9506, 9512, 9521, 9539, 9544, 9554, 1822, 2719, 2730, 350101-350403, 3506-350803, 3510-3517, 3519-3521, 3523-3530, 3533-3536, 3538-3542, 3545-3550, 3552, 3554-3557, 3559, 3561-3563, 3565Q, 3569-3571, 3574, 3576-3578, 3580-3585, 3590-3595, 370101-370105, 3703N-3706, 3708, 3710, 3711, 3713, 3714, 3942, 4028, 44020101-440205, 440208-4414, 4416, 4417, 4419N, 4421-4430, 4432-4440, 4442-4446, 4448-4452N, 4455-446402, 4466, 4467, 4469, 4802, 5104, 5142, 5146, 5156, 530101-5313, 5315-5318, 5514, 5609, 5611, 5612, 570101-5710, 5712, 5714-5719, 580101, 580102, 5804, 5805, 5809-5815, 5817N, 5819, 5821, 8301 |
| 42 | 135 bp | 1070 bp | 7, 8, 18, 35, 37, 39, 56, 78 | *0765, 0832, 180101-1808, 1810-1815, 1817N-1833, 1835, 350101-3518, 352001-3524, 3528-3545, 3548, 3550-3555, 3557-3562, 3564-3572, 3574-3595, 3597, 3599, 3711, 391901, 391902, 5606, 7801-7805 |
| 43 | 245, 400 bp | 1070 bp | 15, 62, 63, 70, 71, 72, 75, 76, 77, 40, 44 | *15010101-150104, 150106-1521, 1523-1540, 1542-1558, 1560-1567, 1569-1575, 1576w, 1577-1599, 9501-9529, 9531-9561, 9563, 9564, 4012, 4410, 8201, 8202 |
| 44 | 210 bp | 800 bp | 7, 8, 15, 18, 37, 38 | *0727, 370101-3714, 3817 |
| 45 | 170 bp | 800 bp | 15, 35, 38, 39, 67 | *1569, 3535, 380101-3809, 3811-3819, 39010101, 39010102L, 390103-3920, 3922-3924, 3926-3929, 3931, 3935, 3937-3942, 3944-3946, 3948, 3949, 5820, 670101-6702 |
| 466 | 110 bp | 1070 bp | 38 | *380101-3803, 3805, 3808-3819 |
| 47 | 395 bp | 1070 bp | 14, 38, 39, 64, 65 | *0769, 1401-1410, 380101-380202, 3804, 3805, 3808-3819, 39010101, 39010102L, 390103-390106, 3903-3907, 3909, 3911, 3912, 3914, 3915, 3918-391902, 3922, 3924-3938Q, 3940N-3948, 3950 |
| 48 | 425 bp | 1070 bp | 7, 8, 14, 39, 40, 41, 42, 45, 48, 50, 60, 64, 65, 67, 73 | *070201-070207, 070209-0726, 0728-0735, 0737, 0739-0764, 0766-0772, 0775-0777, 080101-080106, 0804-0835, 0837, 0838, 0840, 1401-1410, 39010101, 39010102L, 390103-3920, 3922-3950, 400101-400106, 4007, 4009?-4011?, 401401?-4015?, 4016, 4018?, 4021?, 4022N?, 4023, 4024?-4038?, 4040?, 4042?-4046?, 4048?-4050?, 4051, 4052?-4058?, 4059, 4060, 4061?-4065?, 4066, 4067?-4069?, 4071?, 407201-4073, 4074?, 4075?, 4077, 4078?, 4079, 4080?-4088?, 4090?-4093?, 4095?, 4098?, 4101-4108, 4201, 4202, 4204-4212, 4501-4509, 480101-4817, 4819, 5001, 5002, 5004, 670101-6702, 7301, 8102 |
| 4913 | 175 bp | 1070 bp | 15, 40, 44, 46, 48, 60 | *9516, 9524, 400101-400106, 4007, 4010, 4012, 4021-4023, 4025, 4033, 4036, 4038, 4042, 4043, 4046-4049, 4051, 4052, 4054, 4055, 4060-4063, 4065-4067, 4069, 4073, 4074, 4076, 4079, 4081, 4084, 4087, 4088, 4092, 4431, 4606, 480301, 480302 |
| 50 | 130 bp | 1070 bp | 40, 61 | *400201-400602, 4008, 4009, 4011, 4013, 4018-4020, 4024, 4026-4029, 4035, 4037, 4039, 4040, 4044, 4050, 4056, 4057, 4064, 4068, 4070, 4071, 4074, 4075, 4078, 4082, 4085, 4086, 4089-4091, 4094-4098 |
| 516 | 105 bp | 1070 bp | 7, 8, 35, 41, 42, 55 | *0704, 0725, 080101-0805, 0807-081201, 0814-0819N, 0821-0824, 0826-0838, 3587, 4101-4108, 4201, 4202, 4204-4207, 4209-4212, 5520 |
| 52 | 325 bp | 1070 bp | 7, 15, 27, 35, 42, 45, 46, 54, 55, 56, 67, 73, 81 | *070201-070207, 070209, 0704-0707, 0709-0712, 0713w, 0714, 0715w, 0717-0726, 0728-0731, 0733-0736, 0739-0749N, 0751-0768, 0770-0772, 0775-0777, 1576, 9501, 2701-2711, 2713-2715, 2717, 2719-2721, 2724-2728, 2730-2738, 2740-2746, 2750-2752, 3576, 4201, 4202, 4204-4206, 4208-4210, 4212, 4506, 460101, 460102, 4602w, 4603-4619, 4621, 5401-5419, 550101-5505, 5507-5517, 5519-5535, 5601-5616, 5618-5622, 5624, 670101, 670102, 6702w, 7301, 8101-8104N, 8201, 8202, 8301 |
| 536,14 | 115, 195, 225, 260 bp | 1070 bp | 35, 37, 44, 47, 51, 57 | *3538, 3545, 3571, 370101-370105, 3703N-3706, 3708, 3710-3714, 44020101-440205, 440208-4414, 4416, 4417, 4419N-4436, 4438-4463, 4465-4467, 4469, 5142, 5707, 5709, 8301 |
| 54 | 215 bp | 800 bp | 8, 12, 15, 41, 42, 44, 45, 51, 52, 53, 55, 56, 58 | *0809, 1583, 4101, 4105-4107, 4204, 4406, 4415, 4418, 4420, 4501-4509, 510101, 510103-510108, 510201, 510203-5104, 5106-510901, 5110-5114, 5116-512402, 5126-5146, 5148, 5149, 5151-5153, 5155-5165, 520102, 520104-5203, 520601, 520602, 5209, 5306, 5520, 5613, 5808 |
| 55 | 130, 270 bp | 1070 bp | 18, 35, 37, 40, 46, 51, 53, 57, 58 | *1410, 1557w, 1822, 3521, 3524, 3581, 3704, 4028, 460101-4619, 4621, 5104, 5146, 5156, 5302, 5306, 5714, 5809 |
| 566 | 90, 410 bp | 1070 bp | 7, 27, 35, 47, 54, 55, 56, 81 | *0765, 2702-2711, 2713-2715, 2717, 2719-2721, 2724-2728, 2730-2738, 2740-2746, 2750-2752, 3576, 47010101-4703, 5401-5419, 550101-5505, 5507-5517, 5519-5535, 5601-5616, 5618-5622, 5624, 8101, 8201, 8202, 8301 |
| 576 | 90, 175 bp | 1070 bp | 15, 27, 35, 40, 48, 62, 75, 77, 81 | *1502, 1508, 151101-151103, 1513, 1515, 1521, 1531, 1544, 1555, 1576w, 1588, 1589, 9512, 9521, 9539, 9544, 9548, 2724, 3546, 4031, 4045, 4080, 480101, 480102, 4804, 4806, 4807, 4809, 4811, 4815, 4816, 4818, 4819, 5535w, 8101-8104N |
| 58 | 145, 430 bp | 1070 bp | 40, 44, 49 | *4013, 4019, 4418, 4425, 4450, 490101, 490102, 4904, 4905, 5412, 5621, 5901-5905 |
| 596 | 120, 210 bp | 800 bp | 13, 15, 45, 49, 50, 51, 52, 55, 56, 58, 62, 63, 78 | *1316, 1504w, 1516, 1542, 1567, 1595, 9537w, 4095, 490101, 4902-4905, 5001, 5002, 5004, 510101-5103, 5105, 5107-510902, 5111N-5114, 5116-512402, 5126-5141N, 5143, 5144N, 5148-5155, 5157, 5158, 5160, 5161w, 5163, 5165, 520101-5214, 550101-550104, 5503, 5505, 5509, 5511, 5515, 5517, 5521w, 5524, 5525, 5528, 5529, 5531, 5533, 560501-5606, 5621, 5808, 7801-7806 |
| 60 | 430 bp | 1070 bp | 27, 40, 44, 45, 51, 52, 53, 57, 58, 59 | *2702, 2730, 4013, 4019, 4406, 4425, 4450, 510101-512402, 5126-5146, 5148, 5149, 5151-5153, 5155-5165, 520101-5214, 530101-5302, 5304-5308, 5310, 5314-5318, 570101-5711, 5713-5719, 580101-5802, 5804-5816, 5818-5821 |
| 61 | 145 bp | 1070 bp | 15, 35, 40, 41, 44, 45, 47, 50, 72 | *1546, 1553, 9506, 9543, 3519, 3547, 3563, 400101-4011, 401401-4016, 4018, 4020, 4022N-4040, 4042-4045, 4048-4075, 4077-4092, 4094, 4095, 4097, 4098, 4101-4104, 4105?, 4106-4108, 4409, 4446, 4501-4509, 4702, 4703w, 5001, 5002, 5004 |
| 62 | 300 bp | 1070 bp | 7, 8, 13, 15, 18, 27, 35, 37, 38, 39, 40, 41, 44, 45, 46, 47, 48, 49, 50, 52, 55, 57, 58, 62, 72, 76, 78 | *0754, 0817, 0838, 130101-1304, 1306, 1308Q-1323, 1325, 1326, 15010101-150104, 150106-150108, 150301-1507, 1512, 1514, 1519, 1520, 1524-152703, 1530, 1532-1536, 1538-1540, 1542, 1543, 1545-1550, 1553, 1554, 1556-1558, 1560-1563, 1565, 1566, 1568-1571, 1573-1575, 1577-1579N, 1581-1587, 1591, 1592, 1594N, 1596-1598, 9501-9507, 9509-9511N, 9513, 9516-9518, 9522, 9523, 9525-9529, 9531, 9532, 9535-9538, 9540-9542, 9545-9547, 9550-9552, 9554-9560, 9563, 9564,1812, 2701-2721, 2724-2746, 2750-2752, 3510, 3513, 3516, 3528, 3563, 3569, 3580, 370101-3707, 3709-3714, 3803, 3804, 390201, 390202, 3908, 391301, 391302, 3922, 3923, 3949, 400101-4007, 4009-4016, 4018-4024, 4026-4040, 4042-4067, 4070-4098, 4101-4108, 4211, 44020101-440205, 440208-4405, 4407-4411, 4413-4467, 4469, 4501-4509, 47010101-4705, 480101-4805, 4807-4815, 4817-4819, 490101-4905, 5001, 5002, 5004, 520101-5214, 5317, 5518, 5534, 7805, 7806 |
| 63 | 160, 215 bp | 1070 bp | 15, 37, 44, 54 | *150104, 370104, 4094, 44020101-440205, 440208, 4406, 4408, 4409, 4411, 4412, 4416, 4419N-4424, 4427, 4433, 4434, 4441, 4444, 4448, 4449, 4452N, 4453, 4455, 4459, 4463, 4466, 4467, 5401-5419, 8301 |
| 64 | 180, 210 bp | 1070 bp | 7, 13, 15, 18, 35, 39, 40, 45, 54, 55, 56, 59, 78 | *0765, 1306, 1542, 1586, 1804, 3560, 3917, 4058, 4508, 4618, 5401, 5402, 5405N, 5407, 5408N, 5410, 5412, 5413, 5416-5419, 550101-5503, 5505, 5507, 5510-5512, 5515, 5516, 5518, 5519, 5521, 5525, 5526, 5529-5531, 5533-5535, 560501, 5606, 5610, 5623, 5901, 5905, 7801, 780202, 7803, Cw*150204 |
| 6515 | 180 bp | 1070 bp | 13, 15, 40, 44, 45, 46, 49, 50, 51, 54, 56 | *1303, 1573, 4071, 4410, 4415, 4418, 4501, 4504-4507, 4611, 490101-4903, 5001, 5002, 5004, 5115, 5162, 5403, 5601, 5602, 5604, 5607, 5608, 5613, 5614, 5616, 5617, 5620, 5624, 5904, 8201, 8202 |
| 666,11 | 90, 240 bp | 800 bp | 55, 57 | *5514, 570101-5715, 5717-5719 |
| 676 | 90 bp | 1070 bp | 13, 15, 39, 40, 56, 58 | *1311, 1573, 3917, 4048, 4071, 5162, 5602, 5604, 5610, 580101-5802, 5804-5821, Cw*0206w |
| 686,8 | 95 bp | 1070 bp | 7, 8, 18, 35, 37, 38, 39, 44, 51, 52, 53, 78 | *0765w, 0832, 180101-1811, 1813-1815, 1817N-1833, 1835, 350101-350902, 3511, 3512, 351401-3515, 3517, 3518, 352001-3524, 3527, 3529-3545, 3548, 3550-3562, 3564-356802, 3570-3572, 3574, 3575, 3576w, 3577-3579, 3581-3595, 3597, 3599, 3708, 3806, 3807, 391901, 391902, 4406, 510101-512402, 5126-5146, 5148-5165, 530101-5316, 5318, 5606w, 7801-7804 |
| 696 | 115, 150 bp | 1070 bp | 15, 18, 27, 35, 39, 46, 54, 55, 56, 62, 73, 75, 76 | *15010101-150104, 150106-150108, 1504, 1507, 1508, 151101-1512, 1514, 1515, 1519, 1524, 1526N-1528, 1530, 1532, 1534, 1535, 1538, 1543, 1545, 1546, 1550, 1553, 1554, 1556-1558, 1560, 1563, 1566, 1568, 1570, 1571, 1573, 1575-1577, 1579N, 1581, 1582, 1585, 1587, 1592, 1594N, 1596, 1597, 9501, 9502, 9504, 9505, 9509-9511N, 9513, 9517, 9518, 9520, 9522, 9525, 9526, 9528, 9529, 9535, 9537, 9540, 9542-9549N, 9552, 9554, 9557, 9559, 9560, 9563, 9564, 1819, 2725, 351401, 351402, 3543, 3544, 3562, 3567, 3579, 3586, 3918, 3936, 460101-4602, 4604, 4605, 4607N, 4609, 4610, 4612, 4614-4617, 5161, 5406, 5521, 5603, 7301, Cw*0620 |
| 7016 | 360 bp | 1070 bp | Bw4 |  |
| 71 | 350 bp | 1070 bp | Bw6 |  |
| 7211 | 290 bp | 800 bp | 7, 8, 35, 37, 40, 41, 42, 48, 53, 81 | *070201-070207, 070209-0706, 0708-071802, 0720-0732, 0734-0739, 0741-0752, 0754-0759, 0761-0772, 0775, 0776, 080101-0808N, 0810, 0811, 0813-0815, 0817-0838, 0840, 3566, 3587, 3707, 4015, 4016, 4030-4032, 4034, 4045, 4059, 4080, 4098, 410201, 4104, 4201, 4202, 420501-4207, 4209-4212, 480101, 480102, 4805-4812, 4814-4819, 5315, 8101-8104N |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B low resolution typing.

In addition, wells number 29, 30, 34, 41, 44, 45, 54, 59, 66 and 72 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-B alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

4Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

The B*0826 and the B*4207 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*1408 and the B*3925N, B*3930, B*3932 to 3934, 3947 and 3950 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*1829 and the B*3532, B*3537, B*3553N, B*3564, B*356801-356802 and 3599 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*5504, B*5508, B*5513, B*5527 and the B*5615, B*5619N and B*5622 give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*5523, B*5532 and the B*5618 give rise to identical amplification patterns wih the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5Primer mix 25 will amplify the Cw*0223 allele, primer mix 27 will amplify the Cw*0305, Cw*0325 and Cw*0327 alleles, primer mix 32 will amplify the Cw*0746 allele, primer mix 64 will amplify the Cw*150204 allele, primer mix 67 will weakly amplify the Cw*0206 allele and primer mix 69 will amplify the Cw*0620 allele.

6Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

7Primer mix 28 may yield less specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

8Primer mixes 29, 30, and 68 may yield less specific PCR product than the other HLA-B low resolution primer mixes.

9Primer mixes 33 and 34 have a higher tendency than other primer mixes to give rise to nonspecific amplifications.

10In primer mix 34, the positive control band may be weaker than for other HLA-B low primer mixes.

11Primer mixes 66 and 72 may give rise to a primer oligomer artifact.

12The B*57 and B*58 alleles might be faintly amplified by primer mix 33.

13The Cw*1701 to Cw*1704 alleles might be faintly amplified by primer mix 49.

14Primer mix 53 might generate a false band of about 500 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

15The Cw*0312 allele might be faintly amplified by primer mix 65.

16The Bw4-associated HLA-A specificities A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 70.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the B*0701 allele has been shown to be in error. The sequence of the B*1305 allele has been shown to be identical to B*1304. The B*1324 allele has been renamed B*132202The B*150105 allele has been corrected and renamed B*9520The B*1522 allele has been renamed B*3543. The sequence of the B*1541 allele has been shown to be identical to B*1539. The B*1559 allele has been renamed B*3544. The sequence of the B*1816 allele has been shown to be identical to B*1814. The sequence of the B*27051 allele has been shown to be identical to B*270502. The sequence of the B*2722 allele has been shown to be identical to the corrected B*2706 sequence. The B*3573 allele has been renamed B*350803. The sequence of the B*39012 allele has been shown to be identical to B*390101. The sequence of the B*3921 allele has been shown to be identical to B*3924.The sequence of the B*4017 allele has been shown to be identical to B*4016. The sequence of the B*4041 allele has been shown to be identical to B*4040. The sequence of the B*4203 allele has been shown to be in error.The sequence of the B*4401 allele has been shown to be identical to B*4402. The sequence of the B*5003 allele has been shown to be identical to B*5002. The sequence of the B*5125 allele has been shown to be identical to B*5122. The B*5147 allele has been renamed B*510902. The sequence of the B*5506 allele has been shown to be identical to B*5504. The sequence of the B*5803 allele has never been assigned. The B*7901 allele has been renamed B*1518. The B*9530 allele has been renamed B*152702.

The B*0826 and the B*4207 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets. The B*1408 and the B*3925N, B*3930, B*3932 to 3934, 3947 and 3950 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets. The B*1829 and the B*3532, B*3537, B*3553N, B*3564, B*356801-356802 and 3599 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets. The B*5504, B*5508, B*5513, B*5527 and the B*5615, B*5619N and B*5622 give rise to identical amplification patterns wih the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.The B*5523, B*5532 and the B*5618 give rise to identical amplification patterns wih the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

2The serological reactivity of all HLA-B alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

‘ser’, serological HLA specificity.

‘w’, might be weakly amplified. ‘?’, nucleotide sequence information not available for the primer matching sequence.

Specificities and sizes of the PCR products of the 23 primer mixes used for DR low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DR serology3 | Amplified HLA-DRB4 alleles |
| --- | --- | --- | --- | --- |
| 737 | 205 bp | 515 bp | 1 | *010101-010205, 0104-0122 |
| 74 | 200 bp | 430 bp | 103 | *0103 |
| 756 | 215, 260 bp | 430 bp | 15 | *15010101-1533 |
| 76 | 210 bp | 430 bp | 16 | *160101-160502, 1607-1613N |
| 775,6 | 120, 220 bp | 430 bp | 3, 17, 18, 11 | *03010101-0343, 1107, 1153, 1525 |
| 785,6 | 80, 210 bp | 430 bp | 3, 17, 11, 13, 14 | *03010101-030106, 0304-0306, 0308-0316, 0318-0320, 0322, 0323, 0325, 0326, 0328, 0330, 0331, 0333, 0334, 0336, 0337, 0343, 110201-1103, 111101, 111102, 111401, 111402, 1116, 1120, 1121, 1136, 1140, 1141, 1148, 1159, 1163, 116501, 116502, 1168, 1170, 1173, 130101-1304, 1308, 1310, 1315-1317, 1319, 1320, 1322-1324, 1327-1329, 1331-1341, 1343, 1345, 1348, 1351-1354, 1357, 1359, 1361, 1363-1366, 1368-1376, 1378-1381, 1383-1385, 1387, 1388, 1416, 1419, 1421, 1482 |
| 795,6 | 85, 210 bp | 430 bp | 3, 18, 11, 13, 14 | *030201-0303, 0327, 0329, 0338, 111301w, 111302w, 1126, 1134, 1315, 1319, 1326, 1344, 1353, 1357, 1385, 1386, 1402-140302, 1406, 1409, 1412, 1413, 1417-1421, 1424, 1427, 1429, 1430, 143201w, 143202w, 1433, 1440, 1441, 1447-1449, 1451, 1463, 1465w, 1467, 1477, 1478, 1480, 1481, 1483, 1485 |
| 805,6 | 100, 175 bp | 430 bp | 4 | *040101-0478 |
| 816 | 205, 230 bp | 430 bp | 7, 13, 14 | *07010101-070103, 0703-0717, 1317, 1450 |
| 826 | 170, 215, 250 bp | 515 bp | 8, 11, 12, 14 | *080101-0819, 0821-0836, 1167, 1204, 1216, 1411, 1415, 1468 |
| 835,6,8 | 85, 135, 180 bp | 430 bp | 3, 9, 11 | *0308, 090102-0908, 1107, 1153 |
| 84 | 205 bp | 430 bp | 10 | *100101-1003 |
| 855,6 | 100, 170 bp | 430 bp | 11, 3, 8 | *0308, 0831, 110101-1170, 1172, 1173 |
| 865,6 | 85, 105 bp | 430 bp | 12, 8 | *0832, 120101-120204, 120302-1219 |
| 87 | 210 bp | 430 bp | 13, 8, 11, 14 | *0820, 110101-110404, 110601, 110602, 110801-111202, 111401-1116, 1118-1121, 1123-1125, 112701-1133, 1135-1151, 115401, 115402, 1156-1166, 1168, 1170, 1172, 1173, 130101-1308, 1310-1316, 1318-1343, 1345-1385, 1387, 1388, 140301, 140302, 1412, 1416, 1419, 1421, 1422, 1425, 1427, 1440, 1453, 1463, 1467, 1469, 1474, 1477, 1478, 1484, 1485 |
| 886 | 195, 210 bp | 430 bp | 13, 8, 11, 12, 14 | *080101-080203, 080401-0809, 0811, 0816, 0817, 0820-0822, 0824, 0826, 0828, 0831, 110101-110602, 1109-111202, 111401-1116, 1120, 1121, 1123-1125, 112701-1130, 1132, 1133, 1135-1141, 1143, 1144, 1146-1151, 115401-1156, 1158-1163, 116501-1170, 1172, 120201-120204, 1213, 1215, 1216, 1218, 1219, 130101-130201, 130203, 1304-130502, 130701-1309, 131101, 131102, 131401-1324, 1326-1329, 1331, 1332, 1334-1336, 1338-1343, 1345-1355, 1357, 1359, 1361-1365, 1367-1376, 1378-1380, 1383, 1384, 1387, 1415, 1416, 1422, 1424, 1425, 1427, 1437, 1453, 1473 |
| 89 | 175 bp | 430 bp | 3, 8, 13, 14 | *03010101-0307, 0309, 0311-0341, 0343, 0820, 130101-1316, 1318-1342, 1344, 1346-1366, 1368-1388, 1402-140302, 140501-1406, 1409, 1412-1414, 1417-1421, 142301, 1424, 1427, 1429, 1430, 1433, 1436, 1437, 1440-1445, 1447, 1448, 1451, 1456, 1459, 1463-1465, 1467, 1477, 1478, 1480, 1481, 1483-1485 |
| 905-7 | 100, 140 bp | 430 bp | 14, 4, 8,13 | *0462, 0469, 0473, 0808, 1169, 1345, 140101, 140102, 1404, 140701, 140702, 1410, 1416, 1422, 1425, 1426, 1428, 1431-143202, 1435, 1437-1439, 1449, 1450, 1453-1455, 1457, 1458, 1460-1462, 1468-1471, 1473-1476, 1479, 1482, 1486 |
| 915-7,9 | 110, 135, 170 bp | 430 bp | 14, 3, 9, 11, 12, 13, 15 | *0310, 090102-090202, 0904-0908, 111301, 111302, 1117, 1152, 1343, 140101-1402, 1404-1411, 1413, 1414, 1416-1418, 1420, 1422-142302, 1426, 1428-1436, 1438, 1439, 1441, 1443-1452, 1454-1457, 1459-1462, 1464, 1465, 1468, 1470-1476, 1479-1483, 1486, 1527 |
| 925-7 | 110, 175, 225 bp | 430 bp | 14, 3, 8, 11, 13, 15, 16 | *0310, 0809, 0820, 0821, 0832, 0835, 111301, 111302, 1117, 1123, 1125, 1131, 1145, 1152, 1155, 1164, 1313, 1318, 1343, 1345, 1347, 1355, 140101-140103, 140301-140503, 140701-1408, 1410-1412, 1414-1416, 1418, 1422-142302, 1425-1428, 1431-143202, 1434-1436, 1438-1440, 1442-1445, 1449, 1450, 1453-1465, 1467-1479, 1481, 1482, 1484-1486, 1521w, 1604w |
| 936 | 165, 240 bp | 430 bp | 52 | DRB3*01010201-0113, DRB3*0201-0224, DRB3*030101-0303 |
| 9410 | 215 bp | 430 bp | 53 | DRB4*01010101-0107 |
| 95 | 175 bp | 430 bp | 51 | DRB5*010101-0113, DRB5*0202-0205 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 75, 80, 90, 91 and 92.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 73 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DR low resolution typing.

In addition, well number 82 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DRB alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

4For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP set are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

The DRB1*0809, DRB1*0821 and DRB4*1415 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1415 allele.

The DRB1*0820, DRB1*1318, DRB1*1347 and DRB1*1355 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1318, DRB1*1347 and DRB1*1355 alleles.

The DRB1*0831 and DRB1*1167 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1167 allele.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to multiple differently sized specific PCR fragments in these primer mixes.

7Primer mixes 73, 90, 91 and 92 may give rise to primer oligomer formation.

8Primer mix 83 may have a higher tendency than other primer mixes to give rise to nonspecific amplifications.

9Primer mix 91 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

10The DRB4*01030102N allele is amplified by the primer pair in well No. 94, whereas the DRB4*0201N and DRB4*0301N null alleles are not amplified by this primer pair.

‘w’, might be weakly amplified.

1The DRB1*0809, DRB1*0821 and DRB4*1415 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1415 allele.

The DRB1*0820, DRB1*1318, DRB1*1347 and DRB1*1355 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1318, DRB1*1347 and DRB1*1355 alleles.

The DRB1*0831 and DRB1*1167 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1167 allele.

The sequence of the DRB1*0702 allele has been shown to be identical to DRB1*070101. The sequence of the DRB1*08031 allele has been shown to be identical to DRB1*080302. The sequence of the DRB1*09011 allele has been shown to be identical to DRB1*090102. The sequence of the DRB1*12031 allele has been shown to be identical to DRB1*1201. The sequence of the DRB1*1466 allele has been shown to be identical to DRB1*142302. The sequence of the DRB1*1606 allele has been shown to be identical to DRB1*1605. The sequence of the DRB3*010101 allele has been shown to be identical to DRB3*01010201. The DRB4*0101102N allele has been renamed DRB4*01030102N. The sequence of the DRB5*0201 allele has been shown to be identical to DRB5*0202.

2The serological reactivity of all DRB alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

‘ser’, serological HLA specificity.

‘w’, may be weakly amplified.

‘f’, may  be faintly amplified.

‘?’, nucleotide sequence information for the primer matching region(s) is not known.

Change in revision R01 compared to R00:

- In primer mix 34, the positive control band may be weaker than for other HLA-B low primer mixes.

- The numbers of the wells referred to in footnote 7 of the specificity table have been corrected.

Change in revision R02 compared to R01:

- Primer mix 3 may faintly amplify the A*30:04, 30:06, 30:17 and 30:29 alleles.