PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A-B-DR Combi Tray

Prod. No:	101.701-24/24u

101.701-06/06u

Lot No:    	74K

Expiry Date:    	2013-June-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-A low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-A serology3 | Amplified HLA-A alleles4,5, 6 |
| --- | --- | --- | --- | --- |
| 17,8 | 120 bp, 140 bp, 225 bp | 800 bp | A1, A36 | *01:01:01:01-01:04N, 01:06-01:81, 03:18, 36:01-36:05 |
| 28 | 210 bp, 255 bp, 365 bp, 415 bp | 800 bp | A2, A19, A28, A203, A210 | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:22:02, 02:24-02:35:01, 02:35:03-02:47, 02:49-02:77, 02:78w, 02:79:01-02:97:02, 02:99, 02:101-02:128, 02:130-02:282 |
| 39 | 205 bp, 235 bp | 1070 bp | A1, A3, A11, A32, A34, A36 | *01:12, 01:19, 01:21, 03:01:01:01-03:17, 03:19-03:74, 03:76-03:94, 03:96-03:112, 11:25, 11:60, 24:92, 32:04, 34:02-34:04, 34:07-34:09, 36:02 |
| 4 | 190 bp | 800 bp | A1, A2, A3, A11, A26, A30, A36 | *01:01:01:01-01:01:22, 01:01:24-01:04N, 01:06-01:07, 01:09-01:11N, 01:13, 01:16N-01:18N, 01:20-01:29, 01:31N-01:78, 01:80-01:81, 02:78, 02:169, 03:12, 03:18, 03:88, 11:01:01-11:27, 11:29-11:80, 26:19, 30:08, 36:04, 68:66 |
| 5 | 160 bp, 535 bp | 800 bp | A3,A9,A23, A24, A31, A32 | *03:30, 23:01:01-23:31, 24:02:01:01-24:11N, 24:13:01-24:15, 24:17-24:64, 24:66-24:88, 24:90N-24:128, 24:130-24:155N, 31:08, 32:05, 32:13 |
| 6 | 135 bp, 175 bp, 210 bp | 800 bp | A9,A23, A24, A29, A80 | *23:01:01-23:31, 24:05, 24:13:02, 24:24, 29:07, 31:29, 80:01-80:02, B*18:27 |
| 7 | 175 bp, 200 bp | 1070 bp | A2, A23, A24, A26, A33 | *02:17:01w-02:17:02w, 23:14, 24:02:01:01-24:11N, 24:13:01-24:13:02, 24:17-24:50, 24:54-24:56, 24:58-24:63, 24:66-24:91, 24:93, 24:95-24:113, 24:115-24:137, 24:139-24:155N, 26:16, 33:19, 68:45 |
| 8 | 165 bp, 200 bp | 800 bp | A2, A3, A10, A11, A25, A26, A28,A32, A34, A66, A68, A69 | *01:51, 02:55, 03:24, 03:50, 11:10, 25:01:01-25:13, 26:01:01-26:06, 26:08-26:15, 26:17-26:18, 26:20-26:43:02, 26:45-26:56, 32:15, 34:01:01-34:09, 66:01-66:15, 68:01:01-68:66, 69:01 |
| 97 | 75 bp | 800 bp | A25, A32 | *25:01:01-25:13, 32:01:01-32:02, 32:04, 32:06-32:30 |
| 107 | 85 bp | 1070 bp | A2, A10, A26, A33 | *01:51, 02:146, 26:01:01-26:02, 26:04, 26:07:01-26:18, 26:20, 26:22-26:29, 26:31-26:43:02, 26:45-26:56, 33:13 |
| 117,8 | 80 bp, 175 bp, 500 bp | 1070 bp | A1, A11, A24, A26, A31, A34, A66 | *01:13, 01:28, 03:63, 03:88, 11:01:01-11:27, 11:29-11:80, 24:19, 24:44, 26:03:01-26:03:02, 26:06, 26:21, 31:03, 34:01:01-34:08, 66:01, 66:04-66:11, 66:13-66:15, 80:02 |
| 12 | 185 bp | 800 bp | A11, A25, A26, A31, A34, A43, A66 | *03:01:19, 11:11, 25:06, 26:09, 31:03-31:04, 34:01:01-34:09, 43:01, 66:02-66:03 |
| 13 | 175 bp, 225 bp | 1070 bp | A1, A2, A3, A25, A26, A34, A43, A66, A74 | *01:13, 02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:103, 02:135, 03:01:01:01-03:01:22, 03:02-03:07, 03:09-03:11N, 03:13-03:31, 03:33-03:35, 03:37-03:40, 03:42-03:56, 03:58, 03:60-03:71, 03:73-03:87, 03:90-03:106, 03:109-03:110, 03:112, 25:01:01-25:05, 25:07-25:13, 26:01:01-26:01:18, 26:02w, 26:03:01-26:03:02, 26:05-26:08, 26:10-26:33, 26:35-26:43:02, 26:45-26:56, 34:08, 43:01, 66:01, 66:04-66:15, 74:13 |
| 147 | 100 bp, 200 bp, 240 bp | 1070 bp | A29, A33 | *02:237, 03:95, 26:19, 26:22, 29:01:01:01-29:27, 31:03-31:04, 33:13, 34:04, 66:09 |
| 157,8 | 90 bp, 135 bp, 205 bp | 1070 bp | A1, A30 | *01:13, 01:28, 03:43, 03:82, 30:01:01-30:04:02, 30:06-30:20, 30:22-30:46, 31:35 |
| 16 | 240 bp, 370 bp, 395 bp | 1070 bp | A31, A32 | *02:237, 03:95, 29:14, 31:01:02-31:46, 32:05 |
| 17 | 140 bp, 180 bp | 1070 bp | A29, A32 | *03:43, 03:82, 29:13, 31:35, 32:01:01-32:03, 32:05-32:30, 36:02, 74:07 |
| 18 | 200 bp | 1070 bp | A33, A68 | *02:243, 33:01:01-33:01:05, 33:03:01-33:34, 68:29 |
| 19 | 160 bp, 200 bp | 800 bp | A74 | *29:19, 74:01-74:14N |
| 2011 | 220 bp, 245 bp | 800 bp | A2, A210, A25, A68 | *02:34-02:35:03, 02:46, 02:48, 02:56:01-02:56:02, 02:62, 02:70, 02:78, 02:103, 02:129, 25:05, 26:54, 68:01:01-68:66 |
| 21 | 240 bp, 375 bp | 800 bp | A2, A24, A26, A68, A69 | *02:55, 02:243, 24:82, 26:22, 33:22, 66:09, 68:29, 69:01 |
| 227,10 | 85 bp, 240 bp | 800 bp | A2, A36 | *02:34-02:35:03, 02:46, 02:48, 02:56:01-02:56:02, 02:62, 02:70, 02:78, 02:103, 02:129, 36:01-36:05 |
| 237,11 | 75 bp, 160 bp, 240 bp, 495 bp | 800 bp | A2, A26, A68, A80 | *02:55, 03:41, 03:63, 03:75, 03:88, 24:18, 26:03:01-26:03:02, 26:05-26:06, 26:21, 26:30, 33:24, 36:02, 68:05, 68:15, 68:20, 80:01 |
| 24 | 360 bp | 1070 bp | A1, A11, A36, A80 | *02:135, 03:01:19, 25:13, 26:30, 31:04, 34:09, 66:02-66:03, 66:12 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A low resolution typing.

In addition, wells number 2, 4, 5, 6, 8, 9, 12 and 19 to 23 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-A alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

4For several HLA-A alleles only partial 1st exon nucleotide sequences are available. We assume that unknown sequences are conserved within allelic groups.

Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

5A*23:14, 24:05, 24:13:02 and 24:24 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

6Primer mix 6 will amplify the B*18:27 allele.

7Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

8The primer pairs in wells 1, 2, 11 and 15 will in many samples give rise to two or three HLA-specific PCR fragments.

9Primer mix 3 may faintly amplify the A*30:04, 30:06, 30:17 and 30:29 alleles.

10Primer mix 22 might faintly amplify most A*11 alleles.

11Primer mixes 20 and 23 may have a tendency of primer dimer formation.

‘w’, might be weakly amplified.

1The sequence of the A*01:05N has been shown to be identical to A*01:04N.

The A*02:01:16 allele has been renamed to A*02:134.

The A*02:01:20 has been renamed to A*02:01:18.

The sequence of the A*02:23 allele has been shown to be identical to A*02:22:01.

The sequence of the A*02:98 allele has been shown to be identical to A*02:96.

The A*11:28 allele has been renamed to A*11:15:02. The sequence of the A*24:01 allele has been shown to be in error.

The sequence of the A*24:12 allele has been shown to be identical to A*24:08.

The A*24:16 allele has been renamed to A*31:08.

The A*24:65 allele has been renamed to A*24:13:02.

The A*26:44 allele has been renamed to A*26:43:02.

The sequence of the A*30:05 allele has been shown to be identical to A*30:04.

The A*30:21 allele has been renamed to A*30:11:02.

The sequence of the A*31:011 allele has been shown to be identical to A*31:01:02.

The sequence of the A*33:02 allele has been shown to be identical to A*33:03:01.

2The serological reactivity of all HLA-A alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

3Primer mixes 1, 2, 11 and 15 will in many samples give rise to two or three HLA-specific PCR fragments.

4The A*23:14, 24:05, 24:13:02 and 24:24 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

‘w’, may be weakly amplified.

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-B low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-B serol.3 | Amplified HLA-B alleles4,5 |
| --- | --- | --- | --- | --- |
| 256,10 | 110 bp | 800 bp | 7, 703, 40, 41, 42, 48, 61 | *07:02:01-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:47, 07:49N-07:50, 07:52, 07:54-07:59, 07:61-07:99, 07:101-07:117, 15:138, 37:07, 40:15-40:16, 40:23, 40:32, 40:98, 40:136, 40:158, 41:08, 42:05:01-42:05:02, 48:05, 48:08, 48:15, C*02:23, C*04:77 |
| 26 | 215 bp | 1070 bp | 8, 44 | *08:01:01-08:05, 08:07-08:25, 08:27-08:49, 08:51-08:61, 08:63-08:64, 08:66, 15:142, 15:180, 44:49, 51:68 |
| 27 | 140 bp, 235 bp | 1070 bp | 7, 8, 13, 15, 35, 4005, 44, 49, 61, 62, 77 | *07:20, 07:24, 07:60, 07:100, 08:21, 08:25, 13:01:01-13:04, 13:06-13:08Q, 13:10-13:23, 13:25-13:38, 13:40, 15:07:01-15:07:02, 15:36w, 15:45, 15:55, 15:68, 15:89w, 15:126, 15:205, 15:207, 35:05:01-35:05:02, 35:16-35:17, 35:22, 35:30, 35:51, 35:58, 35:72, 35:89, 35:97, 35:113-35:114, 35:125, 40:05, 40:71, 44:08w, 44:54, 44:57w, 44:60, 44:106, 44:110, 46:12, 46:20, 51:64, 53:14, 58:18, C*03:05, C*03:25, C*03:27 |
| 287,8 | 130 bp, 265 bp | 800 bp | 12, 13, 14, 17, 21, 35, 40, 41, 44, 45, 47, 49, 50, 60, 61, 64, 65 | *13:01:01-13:04, 13:06-13:13, 13:15-13:23, 13:25-13:40, 14:01:01-14:04, 14:07N, 14:09, 14:11-14:12, 14:14-14:19, 15:46, 15:53, 15:106, 18:44, 35:46-35:47, 35:63, 40:01:01-40:01:17, 40:07, 40:10:01-40:10:02, 40:14:01-40:16, 40:22N-40:23, 40:25, 40:30-40:34, 40:36, 40:38, 40:42-40:43, 40:45, 40:47-40:49, 40:51-40:55, 40:58-40:63, 40:65-40:67, 40:69, 40:72:01-40:73, 40:76-40:77, 40:79-40:81, 40:84, 40:87-40:88, 40:92, 40:100-40:102, 40:106, 40:108, 40:110, 40:112-40:114, 40:116-40:118N, 40:121, 40:123-40:126, 40:128-40:130, 40:132, 40:134-40:136, 40:139-40:141, 40:146-40:147, 40:150-40:156, 40:158, 41:01-41:16, 44:02:01:01-44:03:06, 44:03:08-44:05:02, 44:09-44:39, 44:41-44:43:02, 44:45-44:56N, 44:58N-44:80, 44:82, 44:84-44:102, 44:104-44:110, 44:112-44:115, 45:01-45:13, 47:01:01:01-47:07, 49:01:01-49:02, 49:04-49:15, 50:01:01-50:02, 50:04-50:06, 50:08-50:12, C*01:30 |
| 298 | 185 bp, 235 bp | 800 bp | 7, 16, 17, 27, 2708, 37, 38, 3901, 3902, 39, 64, 67 | *14:01:01-14:01:02, 14:07N-14:08, 14:10, 14:12, 14:14, 14:19, 27:01-27:06, 27:08-27:10, 27:12-27:13, 27:16-27:18, 27:20, 27:23, 27:26-27:27, 27:29, 27:31, 27:35-27:37, 27:39-27:42, 27:44-27:46, 27:48-27:61, 27:64N-27:69, 27:72-27:73, 37:02, 38:01:01-38:02:02, 38:03, 38:07-38:24, 38:26, 39:01:01:01-39:01:01:02L, 39:01:03-39:01:08, 39:01:10-39:02:01, 39:03, 39:05:01-39:09, 39:11, 39:14-39:15, 39:18, 39:19:02, 39:22, 39:24-39:44, 39:46-39:48, 39:50-39:62, 47:04-47:05, 48:21, 67:01:01, 67:03 |
| 30 | 190 bp | 800 bp | 14, 35, 38, 39, 65 | *07:28, 14:02:01-14:02:02, 14:02:04-14:02:05, 14:03-14:06:02, 14:09, 14:11, 14:13, 14:15-14:18, 15:77, 15:189, 35:26, 38:05, 39:04, 44:16, 44:37, 44:64:01-44:64:02, 44:91, 57:04, A*23:31, A*24:106, C*16:10 |
| 31 | 290 bp | 1070 bp | 15, 22, 62, 63, 71, 72, 75, 76, 77, 46 | *15:01:01:01-15:01:03, 15:01:06-15:02:05, 15:03:03-15:08, 15:10:02-15:11:03, 15:11:05-15:15, 15:17:01:01-15:17:02, 15:19-15:21, 15:24-15:28, 15:30-15:36, 15:38:01-15:40, 15:42-15:46, 15:48, 15:50, 15:55-15:58, 15:60, 15:63, 15:65-15:66, 15:70-15:71, 15:73, 15:75-15:79N, 15:81-15:89, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:107, 15:109-15:113, 15:116-15:118, 15:120-15:122, 15:125-15:126, 15:128-15:129, 15:135-15:150, 15:152, 15:154-15:155, 15:157, 15:159-15:160, 15:162-15:172, 15:174-15:175, 15:177-15:179, 15:181N-15:185, 15:187-15:196, 15:199, 15:201-15:209N, 46:01:01-46:26 |
| 32 | 165 bp, 220 bp, 330 bp | 1070 bp | 5, 8, 12, 21, 22, 37, 41, 42, 44, 45, 48, 51, 56, 57, 60, 62, 70, 71, 72, 82 | *08:01:01-08:05, 08:08N-08:12:03, 08:15-08:19N, 08:21-08:24, 08:26-08:27, 08:29-08:36, 08:38-08:39, 08:41-08:48, 08:50-08:54, 08:56-08:66, 15:03:01-15:03:03, 15:09-15:10:02, 15:18:01-15:18:04, 15:23, 15:29, 15:37, 15:46-15:47, 15:49, 15:51-15:54, 15:61-15:62, 15:64, 15:69, 15:72, 15:74, 15:80, 15:83, 15:90-15:91, 15:93, 15:98-15:99, 15:103, 15:106, 15:108, 15:114-15:115, 15:119, 15:123-15:124, 15:127, 15:131-15:134, 15:143, 15:151, 15:153, 15:156, 15:158, 15:161, 15:173, 15:176, 15:180, 15:186, 15:197-15:198, 15:200, 35:87, 37:12, 40:12, 40:136, 41:01-41:03:02, 41:05-41:09, 41:11-41:16, 42:01:01-42:02, 42:04-42:08, 42:10-42:14, 44:14-44:15, 44:18, 44:20, 44:62, 44:100, 45:01-45:13, 51:08, 51:20, 51:36, 51:44N, 51:97, 52:19, 55:20, 56:13, 57:09, 82:01-82:03, C*07:46 |
| 338,9,11 | 165 bp, 190 bp, 390 bp | 1070 bp | 5, 17, 21, 35, 51, 5102, 5103, 52, 56, 62, 63, 70, 78 | *15:01:02, 15:09, 15:16-15:17:02, 15:67, 15:95, 15:162, 15:168, 15:177, 15:196, 15:208, 35:01:10, 35:04:02, 40:26, 40:28, 44:62, 51:01:01-51:01:03, 51:01:05-51:02:03, 51:03-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03-51:24:04, 51:26-51:33, 51:37-51:41N, 51:43-51:44N, 51:46, 51:48-51:61, 51:63-51:80, 51:82-51:91, 51:94-51:98N, 51:100-51:106, 52:01:01-52:01:04, 52:01:06-52:13, 52:15-52:22, 56:05:01-56:06, 56:21, 58:08, 78:01-78:07 |
| 349 | 180 bp | 800 bp | 13, 22, 55, 60 | *07:78, 13:01:01-13:02:09, 13:07N-13:09, 13:11, 13:14-13:20, 13:22:01-13:23, 13:25, 13:27-13:34, 13:36-13:40, 40:48, 45:10, 49:07, 55:09, 55:22, 55:24 |
| 356 | 105 bp, 195 bp | 1070 bp | 8, 12, 27, 38, 39, 3902, 40, 44, 45, 48, 60, 70, 71, 72, 82 | *07:27, 07:50, 08:04, 08:17, 08:54, 15:03:01-15:03:03, 15:47, 15:49, 15:54, 15:61-15:62, 15:64, 15:68-15:69, 15:91, 15:98, 15:103, 15:123, 15:127, 15:131-15:132, 15:151, 15:156, 15:158, 15:173, 18:37, 27:18, 27:29, 38:03, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:23, 39:39, 39:49, 40:12, 40:46, 40:93, 42:11, 44:10, 44:15, 44:18, 44:40, 44:44, 45:01, 45:05-45:07, 45:11-45:13, 48:01:01-48:05, 48:07-48:24, 50:02, 52:16, 55:18, 82:01-82:03 |
| 36 | 280 bp | 1070 bp | 5, 7, 8, 13, 15, 16, 17, 18, 22, 35, 40, 42, 44, 46, 48, 49, 51, 53, 55, 56, 57, 58, 60, 61, 62, 63, 70, 71, 72, 75, 76, 77, 78 | *07:09, 07:11, 07:17, 08:28, 08:35, 08:37, 13:04, 13:10, 13:26, 15:01:01:01-15:01:04, 15:01:06-15:08, 15:11:01-15:16, 15:18:01-15:21, 15:23-15:29, 15:31-15:36, 15:38:01-15:40, 15:43-15:44, 15:46-15:47, 15:49-15:57, 15:60-15:62, 15:64-15:72, 15:74-15:76, 15:78:01-15:82, 15:84-15:85, 15:87-15:89, 15:91-15:98, 15:101-15:129, 15:131-15:132, 15:134-15:136, 15:138-15:149N, 15:151-15:161, 15:163-15:167, 15:169-15:176, 15:178-15:187, 15:189-15:195, 15:197-15:207, 15:209N, 18:01:01-18:15, 18:17N-18:25, 18:27-18:40, 18:42-18:52, 27:41, 35:01:01:01-35:01:22, 35:05:01-35:05:02, 35:07-35:08:04, 35:10-35:11:02, 35:14:01-35:17, 35:19-35:21, 35:23-35:30, 35:32, 35:35, 35:37, 35:40N-35:43:02, 35:45-35:54, 35:57-35:58, 35:61-35:64, 35:66-35:69, 35:71-35:72, 35:76-35:80, 35:82, 35:86, 35:89-35:94, 35:97, 35:99-35:105, 35:107-35:108:02, 35:110-35:126, 35:130N-35:135, 35:137-35:148, 39:07, 39:43, 40:03, 40:20, 40:38, 40:52, 40:59-40:60, 40:105, 40:158, 42:09, 44:17, 44:43:01-44:43:02, 45:09, 46:01:01-46:10, 46:12-46:17, 46:19-46:26, 48:02:01-48:02:02, 48:14, 48:23, 49:04-49:05, 51:37, 51:45, 51:63, 51:97, 53:01:01-53:03, 53:05-53:06, 53:08:01-53:18, 53:20-53:24, 54:06, 54:09, 54:14, 55:14, 55:23, 55:32, 56:03, 56:09, 56:18, 57:01:01-57:01:10, 57:06, 57:08, 57:10-57:11, 57:13-57:16, 57:18-57:27, 57:29-57:31, 57:33-57:37, 58:01:01-58:02, 58:04-58:07, 58:09-58:19, 58:21-58:26, 58:29-58:31N |
| 37 | 195 bp | 1070 bp | 27, 35, 37, 44, 58, 76 | *15:12, 15:14, 15:19, 15:91, 15:131, 15:161, 35:45, 35:71, 37:01:01-37:09, 37:12-37:13, 37:15-37:21, 37:23-37:25, 38:17, 44:17, 44:43:01-44:43:02, 45:09, 46:17, 53:22, 58:07 |
| 386 | 105 bp, 395 bp, 435 bp | 1070 bp | 5, 7, 15, 41, 42, 62, 63, 70, 71, 75, 77 | *07:04, 07:25, 15:09-15:10:02, 15:13, 15:16-15:18:04, 15:21, 15:23-15:24, 15:37, 15:44, 15:51-15:52, 15:66-15:67, 15:72, 15:80, 15:87, 15:90, 15:93, 15:95, 15:99, 15:108, 15:114-15:115, 15:119, 15:124, 15:133-15:134, 15:153, 15:157, 15:161-15:162, 15:168, 15:176-15:177, 15:186, 15:189, 15:196-15:198, 15:200, 15:208, 40:136, 41:08, 42:05:01-42:05:02 |
| 396 | 115 bp, 150 bp | 1070 bp | 18 | *18:01:01-18:15, 18:17N-18:28, 18:30-18:52 |
| 406 | 80 bp | 1070 bp | 7, 27, 2708, 44, 60 | *07:73, 27:01-27:21, 27:23-27:51, 27:53-27:66N, 27:68-27:73, 38:22, 40:46, 40:93, 44:40, 44:44 |
| 41 | 150 bp | 800 bp | 12, 13, 17, 18, 22, 27, 35, 37, 39, 44, 45, 47, 48, 51, 53, 56, 57, 58, 62, 70, 75, 77, 78 | *08:49, 13:01:01-13:01:05, 13:06-13:07N, 13:12-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:36, 13:39, 14:10, 15:02:01-15:02:05, 15:13, 15:20-15:21, 15:25:01-15:25:03, 15:36, 15:44, 15:62, 15:77, 15:80, 15:85, 15:88-15:89, 15:106, 15:112, 15:121, 15:139, 15:144, 15:154, 15:165, 15:170, 15:194, 15:204, 18:22, 27:19, 27:30, 35:01:01:01-35:04:03, 35:06-35:08:04, 35:10-35:17, 35:19-35:21, 35:23-35:30, 35:33-35:36, 35:38-35:42:02, 35:45-35:50, 35:52, 35:54-35:57, 35:59, 35:61-35:63, 35:65Q, 35:69-35:71, 35:74, 35:76-35:78, 35:80-35:85, 35:90-35:96, 35:98, 35:100-35:101:02, 35:103-35:113, 35:115-35:116, 35:120-35:126, 35:128-35:134N, 35:136-35:150, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:11, 37:13-37:25, 38:20, 39:42, 40:28, 44:02:01:01-44:14, 44:16-44:17, 44:19N, 44:21-44:30, 44:32-44:40, 44:42-44:46, 44:48-44:52N, 44:55-44:64:02, 44:66-44:98, 44:101-44:105, 44:107-44:115, 48:02:01-48:02:02, 51:04, 51:42, 51:46, 51:56, 53:01:01-53:13, 53:15-53:24, 55:14, 56:09, 56:11-56:12, 57:01:01-57:01:04, 57:01:06-57:10, 57:12, 57:14-57:20, 57:22-57:30, 57:32-57:37, 58:01:01-58:01:02, 58:01:04-58:01:08, 58:04-58:05, 58:09-58:15, 58:17N, 58:19, 58:21-58:24, 58:28-58:31N, 83:01, C*03:102 |
| 428 | 135 bp | 1070 bp | 8, 18, 22, 35, 39, 78 | *07:65, 08:32, 15:202, 18:01:01-18:08, 18:10-18:15, 18:17N-18:36, 18:38-18:47, 18:50-18:52, 35:01:01:01-35:03:02, 35:03:04-35:18, 35:20:01-35:24:02, 35:28-35:45, 35:48, 35:50-35:55, 35:57-35:62, 35:64-35:72, 35:74-35:151, 37:11, 39:19:01-39:19:02, 56:06, 78:01-78:05, 78:07 |
| 436,10 | 60 bp, 245 bp, 400 bp | 1070 bp | 5, 15, 44, 48, 60, 62, 63, 70, 71, 72, 75, 76, 77, 82 | *15:01:01:01-15:01:04, 15:01:06-15:21, 15:23-15:40, 15:42-15:58, 15:60-15:67, 15:69-15:99, 15:102-15:129, 15:131-15:209N, 40:12, 44:10, 82:01-82:03 |
| 449,10 | 210 bp | 800 bp | 27, 37 | *07:27, 08:49, 37:01:01-37:25, 38:17 |
| 45 | 170 bp | 800 bp | 16, 35, 38, 39, 3901, 3902, 67, 72 | *08:55, 15:69, 15:186, 35:35, 38:01:01-38:09, 38:11-38:21, 38:23-38:26, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:24, 39:26-39:29, 39:31, 39:35, 39:37-39:42, 39:44-39:46, 39:48-39:49, 39:51-39:62, 51:101, 58:20, 67:01:01-67:03 |
| 466,10 | 110 bp | 1070 bp | 38 | *38:01:01-38:02:02, 38:03, 38:05, 38:08-38:24, 38:26 |
| 47 | 395 bp | 1070 bp | 14, 16, 38, 39, 3901, 64, 65 | *07:69, 07:85, 08:65, 14:01:01-14:19, 38:01:01-38:02:03, 38:04-38:05, 38:08-38:25, 39:01:01:01-39:01:01:02L, 39:01:03-39:01:12, 39:03-39:07, 39:09, 39:11-39:12, 39:14-39:15, 39:18-39:19:02, 39:22, 39:24-39:38Q, 39:40N-39:48, 39:50-39:57, 39:59-39:62 |
| 48 | 160 bp, 425 bp | 1070 bp | 7, 703, 8, 14, 15, 16, 21, 39, 3901, 3902, 40, 41, 42, 45, 48, 50, 60, 61, 64, 65, 67, 70, 73, 81 | *07:02:01-07:26, 07:28-07:35, 07:37, 07:39-07:64, 07:66-07:80, 07:82-07:117, 08:01:01-08:01:13, 08:04-08:05, 08:07-08:35, 08:37-08:51, 08:53-08:66, 14:01:01-14:02:04, 14:02:06-14:19, 27:14, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:62, 40:01:01-40:01:17, 40:06:01:01-40:07, 40:09?-40:11:02?, 40:14:01?-40:15?, 40:16, 40:18?, 40:21?-40:22N?, 40:23, 40:24?-40:38?, 40:42?-40:43?, 40:44, 40:45?-40:46?, 40:48?-40:50?, 40:51, 40:52?, 40:53, 40:54?-40:58?, 40:59-40:60, 40:61?-40:65?, 40:66, 40:67?-40:69?, 40:70, 40:71?, 40:72:01-40:73, 40:74?, 40:75, 40:77, 40:78?, 40:79, 40:80?-40:82?, 40:83, 40:84?-40:85?, 40:86, 40:87?-40:88?, 40:90?-40:92?, 40:93, 40:95?, 40:96, 40:98?-40:102?, 40:103, 40:104?-40:108?, 40:109-40:110, 40:111?-40:116?, 40:118N?-40:121?, 40:123?, 40:124:01, 40:124:02?-40:126?, 40:127, 40:128?-40:130?, 40:131, 40:132?-40:134?, 40:136?-40:137?, 40:138-40:140, 40:141?, 40:145?, 40:146-40:148, 40:149?, 40:150, 40:151?, 40:152-40:153, 40:154?, 40:155N, 40:156?, 40:158?, 41:01-41:16, 42:01:01-42:02, 42:04-42:14, 45:01-45:13, 48:01:01-48:17, 48:19-48:24, 50:01:01-50:02, 50:04-50:12, 51:01:09, 51:10, 51:24:02-51:24:04, 54:11, 67:01:01-67:03, 73:01-73:02, 81:02 |
| 4913 | 180 bp | 1070 bp | 15, 40, 44, 48, 60, 62, 71 | *15:116, 15:124, 40:01:01-40:01:06, 40:01:08-40:01:17, 40:07, 40:10:01-40:10:02, 40:12, 40:21-40:23, 40:25, 40:30, 40:33-40:34, 40:36, 40:38, 40:42-40:43, 40:46-40:49, 40:51-40:52, 40:54-40:55, 40:59-40:63, 40:65-40:67, 40:69, 40:73-40:74, 40:76, 40:79, 40:81, 40:84, 40:87-40:88, 40:92, 40:100-40:102, 40:106, 40:108, 40:112-40:114, 40:116-40:118N, 40:123-40:126, 40:128, 40:130, 40:132, 40:134-40:135, 40:137-40:141, 40:146-40:147, 40:149-40:156, 44:31, 46:06, 48:03:01-48:03:02, 48:17, 48:23 |
| 50 | 290 bp | 1070 bp | 21, 4005, 41, 50, 60, 61 | *18:48, 40:02:01-40:06:03, 40:08, 40:09?-40:11:02?, 40:13, 40:14:01?-40:15?, 40:18?-40:19?, 40:20, 40:22N?, 40:24?-40:38?, 40:39, 40:42?-40:45?, 40:47?-40:50?, 40:52?-40:58?, 40:61?-40:65?, 40:67?-40:69?, 40:70, 40:71?, 40:72:02?, 40:74?-40:76?, 40:78?, 40:80?-40:82?, 40:84?-40:88?, 40:89, 40:90?-40:92?, 40:94, 40:95?-40:96?, 40:97, 40:98?-40:121?, 40:122, 40:123?, 40:124:02?-40:136?, 40:141?, 40:142N-40:144N, 40:145?, 40:148?-40:149?, 40:151?, 40:154?, 40:156?-40:158?, 47:01:01:01-47:07 |
| 516 | 105 bp | 1070 bp | 7, 8, 22, 41, 42 | *07:04, 07:25, 08:01:01-08:05, 08:07-08:12:03, 08:14-08:19N, 08:21-08:24, 08:26-08:39, 08:41-08:54, 08:56-08:66, 35:87, 40:136, 41:01-41:08, 41:10-41:16, 42:01:01-42:02, 42:04-42:07, 42:09-42:14, 55:20 |
| 52 | 325 bp | 1070 bp | 7, 22, 27, 2708, 35, 42, 45, 46, 54, 55, 56, 67, 73, 78, 81, 82 | *07:02:01-07:02:20, 07:04-07:07, 07:09-07:15, 07:17-07:26, 07:28-07:31, 07:33-07:36, 07:39-07:46, 07:47w, 07:48-07:49N, 07:51-07:68:02, 07:70-07:84, 07:86-07:117, 15:76, 15:101, 27:01-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:73, 35:76, 38:26, 42:01:01-42:02, 42:04-42:06, 42:08-42:10, 42:12-42:14, 44:90, 44:97, 45:06, 46:01:01-46:01:05, 46:02w, 46:03-46:26, 54:01-54:23, 55:01:01-55:05, 55:07-55:17, 55:19-55:48, 56:01:01-56:16, 56:18-56:22, 56:24-56:30, 67:01:01-67:03, 73:01-73:02, 81:01-81:05, 82:01-82:03, 83:01 |
| 536 | 115 bp, 195 bp, 225 bp, 260 bp | 1070 bp | 35, 37, 44, 47, 57 | *08:49, 08:60, 35:38, 35:45, 35:71, 35:115, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:25, 40:132, 42:13, 44:02:01:01-44:14, 44:16-44:17, 44:19N-44:36, 44:38-44:63, 44:65-44:115, 51:42, 53:22, 57:07, 57:09, 57:24, 83:01, C*15:25 |
| 54 | 215 bp | 800 bp | 5, 8, 12, 17, 21, 22, 41, 42, 44, 45, 51, 5102, 5103, 52, 53, 56, 62 | *08:09, 15:83, 41:01, 41:05-41:07, 41:09, 41:12, 41:14, 41:16, 42:04, 44:06, 44:15, 44:18, 44:20, 44:100, 45:01-45:13, 51:01:01, 51:01:03-51:01:08, 51:01:10-51:02:01, 51:02:03-51:04, 51:06-51:07:01, 51:08-51:09:01, 51:10-51:14, 51:16-51:24:04, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:88, 51:90-51:92, 51:94-51:102, 51:104-51:106, 52:01:02, 52:01:04, 52:02-52:03, 52:06:01-52:06:02, 52:09, 52:19, 52:21, 53:06, 55:20, 56:13, 58:08 |
| 5512 | 130 bp, 270 bp | 1070 bp | 17, 18, 22, 35, 37, 46, 51, 53, 62, 78 | *14:10, 15:57w, 18:22, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 37:04:01-37:04:02, 40:28, 46:01:01-46:26, 51:04, 51:46, 51:56, 53:02, 53:06, 57:14, 58:09, C*15:39 |
| 566,14 | 90 bp, 410 bp | 1070 bp | 15, 22, 27, 35, 47, 54, 55, 56, 78, 81, 82 | *07:65, 27:01w, 27:02-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:73, 35:76, 44:90, 44:97, 47:01:01:01-47:03, 47:06-47:07, 54:01-54:23, 55:01:01-55:05, 55:07-55:17, 55:19-55:48, 56:01:01-56:16, 56:18-56:22, 56:24-56:30, 81:01, 82:01-82:03, 83:01 |
| 576 | 90 bp, 175 bp | 1070 bp | 15, 27, 35, 48, 60, 62, 75, 77, 81 | *13:36, 15:02:01-15:02:05, 15:08, 15:11:01-15:11:05, 15:13, 15:15, 15:21, 15:31, 15:44, 15:55, 15:76w, 15:88-15:89, 15:112, 15:121, 15:139, 15:144, 15:148, 15:170, 15:189, 15:191, 15:194, 15:209N, 27:24, 35:46, 40:31, 40:45, 40:80, 48:01:01-48:01:03, 48:04, 48:06-48:07, 48:09, 48:11, 48:15-48:16, 48:18-48:20, 48:22, 48:24, 51:104, 55:35w, 81:01-81:05 |
| 58 | 145 bp, 430 bp | 1070 bp | 44, 49, 59, 61 | *40:13, 40:19, 40:109, 40:117, 44:18, 44:25, 44:50, 44:95, 49:01:01-49:01:02, 49:04-49:15, 54:12, 56:21, 59:01-59:05 |
| 596 | 120 bp, 210 bp | 800 bp | 5, 13, 15, 17, 22, 35, 45, 49, 50, 51, 5102, 5103, 52, 55, 56, 62, 63, 78 | *07:78, 07:84, 13:16, 13:31w, 15:04w, 15:16, 15:42, 15:67, 15:95, 15:137w, 40:95, 40:148, 49:01:01, 49:02-49:10, 49:12-49:15, 50:01:01-50:02, 50:04-50:11, 51:01:01-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:14, 51:16-51:24:04, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60, 51:61w, 51:63, 51:65-51:80, 51:82-51:106, 52:01:01-52:22, 54:20, 55:01:01-55:01:06, 55:03, 55:05, 55:09, 55:11, 55:15, 55:17, 55:21w, 55:24-55:25, 55:28-55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45, 56:05:01-56:06, 56:21, 56:25, 58:08, 78:01-78:07 |
| 60 | 430 bp | 1070 bp | 5, 17, 27, 44, 51, 5102, 5103, 52, 53, 57, 58, 61 | *27:02, 27:30, 27:53, 27:57, 27:62, 27:65N, 40:13, 40:19, 40:109, 40:117, 44:06, 44:25, 44:50, 44:95, 51:01:01-51:24:04, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:106, 52:01:01-52:19, 52:21-52:22, 53:01:01-53:02, 53:04-53:08:02, 53:10, 53:14-53:24, 57:01:01-57:11, 57:13-57:37, 58:01:01-58:02, 58:04-58:16, 58:18-58:31N |
| 61 | 145 bp | 1070 bp | 12, 21, 35, 40, 4005, 41, 45, 47, 50, 60, 61 | *15:46, 15:53, 15:106, 15:143, 18:48, 35:19, 35:47, 35:63, 40:01:01-40:11:02, 40:14:01-40:16, 40:18, 40:20, 40:22N-40:40, 40:42-40:45, 40:48-40:75, 40:77-40:92, 40:94-40:95, 40:97-40:108, 40:111-40:116, 40:118N-40:136, 40:138-40:156, 40:158, 41:01-41:04, 41:05?, 41:06-41:16, 44:09, 44:46, 44:75, 44:90, 45:01-45:13, 47:02, 47:03w, 50:01:01-50:02, 50:04-50:05, 50:07-50:12 |
| 62 | 300 bp | 1070 bp | 7, 12, 13, 15, 17, 18, 21, 27, 2708, 35, 37, 38, 39, 3902, 40, 4005, 41, 44, 45, 47, 48, 49, 50, 52, 60, 61, 62, 70, 72, 76 | *07:54, 08:17, 08:38, 08:54, 13:01:01-13:04, 13:06, 13:08Q-13:23, 13:25-13:35, 13:37-13:40, 15:01:01:01-15:01:04, 15:01:06-15:01:18, 15:01:20, 15:03:01-15:07:02, 15:12, 15:14, 15:19-15:20, 15:24-15:27:03, 15:30, 15:32-15:36, 15:38:01-15:40, 15:42-15:43, 15:45-15:50, 15:53-15:54, 15:56-15:58, 15:60-15:63, 15:65-15:66, 15:68-15:71, 15:73-15:75, 15:77-15:79N, 15:81-15:87, 15:91-15:92, 15:94N, 15:96-15:98, 15:101-15:107, 15:109-15:111N, 15:113, 15:116-15:118, 15:122-15:123, 15:125-15:129, 15:131-15:132, 15:135-15:138, 15:140-15:142, 15:145-15:147, 15:150-15:152, 15:154-15:160, 15:163-15:167, 15:169, 15:171-15:175, 15:178-15:179, 15:181N-15:185, 15:187-15:188, 15:190N, 15:192-15:193, 15:195, 15:199, 15:201-15:207, 18:12, 27:01-27:21, 27:24-27:73, 35:10, 35:13, 35:16, 35:28, 35:63, 35:69, 35:80, 37:01:01-37:07, 37:09-37:25, 38:03-38:04, 38:25, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:22-39:23, 39:49, 40:01:01-40:07, 40:09-40:16, 40:18-40:24, 40:26-40:40, 40:42-40:67, 40:70-40:105, 40:107-40:158, 41:01-41:16, 42:11, 44:02:01:01-44:05:02, 44:07-44:11, 44:13-44:115, 45:01-45:13, 47:01:01:01-47:07, 48:01:01-48:05, 48:07-48:15, 48:17-48:24, 49:01:01-49:15, 50:01:01-50:02, 50:04-50:12, 52:01:01-52:22, 53:17, 55:18, 55:34, 55:40, 78:05-78:06 |
| 638 | 160 bp, 330 bp | 1070 bp | 22, 44, 45, 47, 54 | *44:02:01:01-44:02:14, 44:06, 44:08-44:09, 44:11-44:12, 44:16, 44:19N-44:24, 44:27, 44:33-44:34, 44:41, 44:44, 44:48-44:49, 44:52N-44:53, 44:55, 44:59, 44:63, 44:66-44:68, 44:71-44:74, 44:80, 44:83-44:84, 44:86-44:87, 44:89-44:91, 44:93, 44:95, 44:97, 44:99-44:102, 44:104, 44:106, 44:112-44:113, 54:01-54:23, 83:01 |
| 64 | 180 bp, 210 bp | 1070 bp | 13, 15, 18, 35, 39, 40, 54, 55, 56, 59, 78 | *07:65, 13:06, 15:42, 15:86, 18:04, 35:42:02, 35:43:02, 35:60, 39:17, 40:01:11, 40:58, 45:08, 46:18, 54:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:16-54:23, 55:01:01-55:03, 55:05, 55:07, 55:10-55:12, 55:15-55:16, 55:18-55:19, 55:21, 55:25-55:26, 55:29-55:31, 55:33-55:45, 55:47-55:48, 56:05:01, 56:06, 56:10, 56:23, 59:01, 59:05, 78:01, 78:02:02-78:03, 78:07, C*15:02:04 |
| 6515 | 180 bp | 1070 bp | 12, 15, 21, 22, 44, 45, 49, 50, 51, 56, 61, 62, 82 | *13:03, 15:73, 40:71, 44:10, 44:15, 44:18, 45:01, 45:04-45:07, 45:11-45:13, 46:11, 49:01:01-49:03, 49:06, 49:08-49:15, 50:01:01-50:02, 50:04-50:08, 50:10-50:12, 51:15, 51:62, 51:106, 54:03, 56:01:01-56:02, 56:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:20, 56:24-56:30, 59:04, 82:01-82:03, C*03:12, C*03:19, C*03:102 |
| 666,10 | 90 bp, 240 bp | 800 bp | 57 | *55:14, 57:01:01-57:37 |
| 676,14 | 90 bp | 1070 bp | 5, 13, 15, 17, 39, 55, 56, 58, 60, 61, 62, 71 | *13:11, 15:18:03, 15:73, 39:17, 40:48, 40:71, 51:62, 56:02, 56:04, 56:10, 58:01:01-58:02, 58:04-58:31N, C*02:06w |
| 686,8 | 95 bp | 1070 bp | 8, 18, 22, 35, 37, 38, 39, 44, 51, 5102, 5103, 53, 78 | *07:65w, 08:32, 18:01:01-18:11, 18:13-18:15, 18:17N-18:36, 18:38-18:47, 18:49-18:52, 35:01:01:01-35:09:02, 35:11:01-35:12:02, 35:14:01-35:15, 35:17-35:18, 35:20:01-35:24:02, 35:27, 35:29-35:45, 35:48, 35:50-35:62, 35:64-35:68:02, 35:70-35:72, 35:74-35:75, 35:76w, 35:77-35:79, 35:81-35:151, 37:08, 38:06-38:07, 39:19:01-39:19:02, 44:06, 51:01:01-51:24:04, 51:26-51:46, 51:48-51:103, 51:105-51:106, 53:01:01-53:16, 53:18-53:24, 56:06w, 78:01-78:04, 78:07 |
| 696 | 115 bp, 150 bp | 1070 bp | 15, 18,22, 27, 35, 39, 46, 55, 62, 72, 73, 75, 76 | *07:100, 13:31, 15:01:01:01-15:01:04, 15:01:06-15:01:16, 15:01:18-15:01:20, 15:04, 15:07:01-15:08, 15:11:01-15:12, 15:14-15:15, 15:19, 15:24, 15:26N-15:28, 15:30, 15:32, 15:34-15:35, 15:38:01-15:38:02, 15:43, 15:45-15:46, 15:50, 15:53-15:54, 15:56-15:58, 15:60, 15:63, 15:66, 15:68, 15:70-15:71, 15:73, 15:75-15:77, 15:79N, 15:81-15:82, 15:85, 15:87, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:105, 15:109-15:111N, 15:113, 15:117-15:118, 15:120, 15:122, 15:125-15:126, 15:128-15:129, 15:135, 15:137, 15:140, 15:142-15:149N, 15:152, 15:154, 15:157, 15:159-15:160, 15:163-15:167, 15:169, 15:171-15:172, 15:174-15:175, 15:178, 15:180-15:184, 15:187, 15:189-15:193, 15:201-15:203, 15:205-15:207, 15:209N, 18:19, 27:25, 35:14:01-35:14:02, 35:43:01-35:44, 35:62, 35:67, 35:79, 35:86, 35:102, 35:117, 35:135, 39:18, 39:36, 46:01:01-46:02, 46:04-46:05, 46:07N, 46:09-46:10, 46:12, 46:14-46:17, 46:20, 46:22-46:24, 51:61, 52:21, 54:06, 55:21, 56:03, 73:01-73:02, A*68:56, C*06:20, C*12:50 |
| 7016 | 360 bp | 1070 bp | Bw4 |  |
| 71 | 350 bp | 1070 bp | Bw6 |  |
| 7210 | 285 bp | 800 bp | 7, 703, 8, 35, 40, 41, 42, 48, 53, 60, 61, 81 | *07:02:01-07:06, 07:08-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:52, 07:54-07:59, 07:61-07:76, 07:79-07:83, 07:85-07:99, 07:101-07:117, 08:01:01-08:05, 08:07-08:08N, 08:10-08:11, 08:13-08:15, 08:17-08:66, 35:66, 35:87, 37:07, 40:15-40:16, 40:30-40:32, 40:34, 40:45, 40:59, 40:80, 40:98, 40:137, 41:02:01-41:02:03, 41:04, 41:10-41:11, 41:13, 42:01:01-42:02, 42:05:01-42:07, 42:09-42:13, 48:01:01-48:01:03, 48:05-48:12, 48:14-48:20, 48:22, 53:15, 81:01-81:05 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may

be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 25 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B low resolution typing.

In addition, wells number 28 to 30, 34, 41, 44, 45, 54, 59, 66 and 72 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-B alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

4Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

The B*08:26, 08:50 and 08:62 and the B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:47 and 39:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*18:29 and the B*35:32, 35:37, 35:53N, 35:64, 35:68:01-35:68:02, 35:99 and 35:118-35:119 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*41:09 and the B*45:02 and 45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:04, 55:08, 55:13, 55:27 and 55:46 and the B*56:15, 56:19N and 56:22 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 alleles give rise to identical amplification patterns wih the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5The C*02:23 and C*04:77 alleles will be amplified by primer mix 25, the HLA-C*03:05, 03:25 and 03:27 alleles will be amplified by primer mix 27, the C*01:30 allele will be amplified by primer mix 28, the A*23:31 and A*24:106 and the C*16:10 alleles will be amplified by primer mix 30, the C*07:46 allele will be amplified by primer mix 32, primer mix 41 will amplify the C*03:102 allele, the C*15:25 allele will be amplified by primer mix 53, the C*15:39 allele will be amplified by primer mix 55, the C*15:02:04 allele will amplified by primer mix 64, the C*03:12, C*03:19 and C*03:102 alleles will amplified by primer mix 65, the C*02:06 allele will weakly amplified by primer mix 67 and the A*68:56 and the C*06:20 and C*12:50 alleles will amplified by primer mix 69.

6Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

7Primer mix 28 may yield less specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

8Primer mixes 28, 29, 33, 42, 63 and 68 give a lower yield of specific PCR product than the other HLA-B low resolution primer mixes.

9Primer mixes 33, 34 and 44 have a tendency to give rise to nonspecific amplifications, most pronounced in primer mix 33.

10Primer mixes 25, 43, 44, 46, 66 and 72 may give rise to a primer oligomer artifact.

11The B*57 and B*58 alleles might be faintly amplified by primer mix 33.

12In primer mix 55, the positive control band may be weaker than for other HLA-B low primer mixes. However, the amplification of the HLA-specific bands is of comparable strength to other HLA B low primer mixes.

13The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 49.

14Primer mixes 56 and 67 may generate a false positive band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

15The C*03:12 allele might be faintly amplified by primer mix 65.

16The Bw4-associated HLA-A specificities A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 70.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the B*07:01 allele has been shown to be in error.

The sequence of the B*08:06 allele has been shown to be identical to B*08:20

The sequence of the B*13:05 allele has been shown to be identical to B*13:04.

The B*13:24 allele has been renamed B*13:22:02

The B*15:01:05 allele has been corrected and renamed B*15:120

The B*15:22 allele has been renamed B*35:43.

The sequence of the B*15:41 allele has been shown to be identical to B*15:39.

The B*15:59 allele has been renamed B*35:44.

The sequence of the B*18:16 allele has been shown to be identical to B*18:14.

The sequence of the B*27:05:01 allele has been shown to be identical to B*27:05:02.

The sequence of the B*27:22 allele has been shown to be identical to the corrected B*27:06 sequence.

The B*35:73 allele has been renamed B*35:08:03.

The sequence of the B*39:012 allele has been shown to be identical to B*39:01:01:01.

The sequence of the B*39:21 allele has been shown to be identical to B*39:24.

The sequence of the B*40:17 allele has been shown to be identical to B*40:16.

The sequence of the B*40:41 allele has been shown to be identical to B*40:40.

The sequence of the B*42:03 allele has never been assigned.

The sequence of the B*44:01 allele has been shown to be identical to B*44:02.

The sequence of the B*50:03 allele has been shown to be identical to B*50:02.

The sequence of the B*51:25 allele has been shown to be identical to B*51:22.

The B*51:47 allele has been renamed B*51:09:02.

The sequence of the B*55:06 allele has been shown to be identical to B*55:04.

The sequence of the B*58:03 allele has never been assigned.

The B*79:01 allele has been renamed B*15:18:01.

The B*9530 allele has been renamed B*15:27:02.

2The B*08:26, 08:50 and 08:62 and the B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:47 and 39:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*18:29 and the B*35:32, 35:37, 35:53N, 35:64, 35:68:01-35:68:02, 35:99 and 35:118-35:119 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*41:09 and the B*45:02 and 45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:04, 55:08, 55:13 and 55:27 and the B*56:15, 56:19N and 56:22 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The serological reactivity of all HLA-B alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

‘ser’, serological HLA specificity.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Specificities and sizes of the PCR products of the 24 primer mixes used for DR low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DR serology3 | Amplified HLA-DRB alleles4 |
| --- | --- | --- | --- | --- |
| 736,8 | 200 bp, 255 bp | 515 bp | 1 | *01:01:01-01:02:05, 01:04-01:34 |
| 74 | 200 bp | 430 bp | 103 | *01:03 |
| 75 | 215 bp | 430 bp | 2, 15 | *15:01:01:01-15:49, 15:50Nw, 15:51 |
| 76 | 210 bp | 430 bp | 16 | *16:01:01-16:05:02, 16:07-16:17 |
| 775,6,11 | 120 bp, 220 bp | 430 bp | 3, 17, 18, 11 | *03:01:01:01-03:58, 11:07, 11:53, 15:25 |
| 785,6,11 | 80 bp, 210 bp | 430 bp | 3, 6, 17, 11, 13, 14 | *03:01:01:01-03:01:09, 03:04-03:06, 03:08-03:16, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:31, 03:33-03:34, 03:36-03:37, 03:43-03:48, 03:50-03:52, 03:54-03:58, 08:40, 11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-11:41, 11:48, 11:59, 11:63, 11:65:01-11:65:02, 11:68, 11:70, 11:73, 11:76, 11:79-11:80, 11:83, 11:85-11:87, 11:93, 13:01:01-13:04, 13:08, 13:10, 13:15-13:17, 13:19-13:20, 13:22-13:24, 13:27-13:29, 13:31-13:41, 13:43, 13:45, 13:48, 13:51-13:54, 13:57, 13:59, 13:61, 13:63-13:66:02, 13:68-13:76, 13:78-13:81, 13:83-13:85, 13:87-13:99, 13:101-13:102, 14:16, 14:19, 14:21, 14:82, 14:95 |
| 795,6 | 85 bp, 210 bp | 430 bp | 3, 6, 11, 13, 14, 1403, 18 | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 11:13:01w-11:13:02w, 11:26, 11:34, 13:15, 13:19, 13:26, 13:44, 13:53, 13:57, 13:85-13:86, 14:02-14:03:02, 14:06:01-14:06:02, 14:09, 14:12:01-14:13, 14:17-14:21, 14:24, 14:27, 14:29-14:30, 14:32:01w-14:32:02w, 14:33, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:65w, 14:67, 14:77-14:78, 14:80-14:81, 14:83, 14:85, 14:89, 14:94, 14:98 |
| 805,6 | 100 bp, 175 bp | 430 bp | 4 | *04:01:01-04:92 |
| 816 | 210 bp, 230 bp | 430 bp | 7, 13, 14 | *07:01:01:01-07:01:03, 07:03-07:19, 12:22, 13:17, 14:50 |
| 826 | 170 bp, 215 bp, 250 bp | 515 bp | 8, 12, 14 | *08:01:01-08:19, 08:21-08:41, 11:67, 12:04, 12:16, 12:22, 14:11, 14:15, 14:68, 14:93 |
| 835-7 | 85 bp, 135 bp, 180 bp | 430 bp | 3, 9, 11 | *03:08, 09:01:02-09:09, 11:07, 11:53 |
| 848 | 205 bp | 430 bp | 10 | *10:01:01-10:03 |
| 855,6 | 100 bp, 170 bp | 430 bp | 3, 8, 11, 14 | *03:08, 08:31, 08:41, 11:01:01-11:70, 11:72-11:98 |
| 865,6 | 85 bp, 105 bp | 430 bp | 12 | *08:32, 12:01:01-12:26 |
| 87 | 215 bp | 430 bp | 6, 8, 11, 13, 14, 1403 | *08:20-08:21, 11:01:01-11:04:06, 11:06:01-11:06:02, 11:08:01-11:12:02, 11:14:01-11:16, 11:18-11:21, 11:23-11:25, 11:27:01-11:33, 11:35-11:51, 11:54:01-11:54:02, 11:56-11:66, 11:68, 11:70, 11:72-11:81, 11:83-11:88, 11:90-11:97, 13:01:01-13:08, 13:10-13:16, 13:18-13:43, 13:45-13:85, 13:87-13:103, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:16, 14:19, 14:21-14:22, 14:25, 14:27, 14:40, 14:53, 14:63, 14:67, 14:69, 14:74, 14:77-14:78, 14:84-14:85, 14:98 |
| 886 | 195 bp, 215 bp | 430 bp | 6, 8, 11, 12, 13, 14 | *08:01:01-08:02:03, 08:04:01-08:09, 08:11, 08:16-08:17, 08:20-08:22, 08:24, 08:26, 08:28, 08:31, 08:39, 08:41, 11:01:01-11:06:02, 11:09-11:12:02, 11:14:01-11:16, 11:20-11:21, 11:23-11:25, 11:27:01-11:30, 11:32-11:33, 11:35-11:41, 11:43-11:44, 11:46-11:51, 11:54:01-11:56, 11:58:01-11:63, 11:65:01-11:70, 11:72, 11:74-11:78, 11:80-11:88, 11:90-11:97, 12:02:01-12:02:05, 12:13, 12:15-12:16, 12:18-12:21, 12:23, 12:26, 13:01:01-13:02:01, 13:02:03-13:02:04, 13:04-13:05:02, 13:07:01-13:09, 13:11:01-13:11:02, 13:14:01-13:24, 13:26-13:29, 13:31-13:32, 13:34-13:36, 13:38-13:43, 13:45-13:55, 13:57, 13:59, 13:61-13:65, 13:67-13:76, 13:78-13:80, 13:83-13:84, 13:87, 13:91-13:93, 13:96-13:100, 13:102-13:103, 14:15-14:16, 14:22, 14:24-14:25, 14:27, 14:37, 14:53, 14:73 |
| 8911 | 175 bp | 430 bp | 3, 6, 11, 13, 14, 1403, 17, 18 | *03:01:01:01-03:07, 03:09, 03:11:01-03:41, 03:43-03:45, 03:47-03:58, 08:20, 13:01:01-13:16, 13:18-13:42, 13:44, 13:46-13:66:02, 13:68-13:102, 14:02-14:03:02, 14:05:01-14:06:02, 14:09, 14:12:01-14:14, 14:17-14:21, 14:23:01, 14:23:03-14:24, 14:27, 14:29-14:30, 14:33, 14:36-14:37, 14:40-14:45, 14:47-14:48, 14:51, 14:56, 14:59, 14:63-14:65, 14:67, 14:77-14:78, 14:80-14:81, 14:83-14:85, 14:89, 14:91, 14:94-14:96, 14:98, 14:100 |
| 905,6,10 | 100 bp, 140 bp, 155 bp | 430 bp | 4, 6, 8,13, 14, 1404 | *04:62, 04:69, 04:73, 08:08, 11:69, 11:82, 13:45, 14:01:01-14:01:02, 14:04, 14:07:01-14:07:02, 14:10, 14:16, 14:22, 14:25-14:26, 14:28, 14:31-14:32:02, 14:35, 14:37-14:39, 14:49-14:50, 14:53-14:55, 14:57-14:58, 14:60-14:62, 14:68-14:71, 14:73-14:76, 14:79, 14:82, 14:86-14:88, 14:90, 14:93, 14:99, DRB4*01:03:01:02N |
| 915,6,9 | 110 bp, 135 bp, 170 bp | 430 bp | 3, 4, 6, 9, 11, 13,14,1404 | *03:10, 09:01:02-09:01:05, 09:01:07-09:02:02, 09:04-09:09, 11:13:01-11:13:02, 11:17, 11:52, 13:43, 14:01:01-14:02, 14:04-14:11, 14:13-14:14, 14:16-14:18, 14:19w, 14:20, 14:21w, 14:22-14:23:03, 14:26, 14:28-14:36, 14:38-14:39, 14:41, 14:43-14:52, 14:54-14:57, 14:59-14:62, 14:64-14:65, 14:68, 14:70-14:76, 14:79-14:83, 14:86-14:88, 14:90-14:97, 14:99-14:100, 15:27, 15:34 |
| 925,6 | 110 bp, 175 bp, 225 bp | 430 bp | 2, 3, 4, 6, 8, 11,13, 14, 1403, 1404, 16 | *03:10, 08:09, 08:20-08:21, 08:32, 08:35, 11:13:01-11:13:02, 11:17, 11:23, 11:25, 11:31, 11:45, 11:52, 11:55, 11:64, 11:89, 11:96, 13:13, 13:18, 13:43, 13:45, 13:47, 13:55, 14:01:01-14:01:03, 14:03:01-14:05:03, 14:07:01-14:08, 14:10-14:12:02, 14:14-14:16, 14:18, 14:22-14:23:03, 14:25-14:28, 14:31-14:32:02, 14:34-14:36, 14:38-14:40, 14:42-14:45, 14:49-14:50, 14:53-14:65, 14:67-14:79, 14:81-14:82, 14:84-14:93, 14:95-14:97, 14:99-14:100, 15:21w, 16:04w |
| 936,11 | 160 bp, 240 bp | 430 bp | 52 | DRB3*01:01:02:01-01:14, DRB3*02:01-02:25, DRB3*03:01:01-03:03 |
| 948,10 | 215 bp | 430 bp | 53 | DRB4*01:01:01:01-01:08 |
| 95 | 175 bp | 430 bp | 51 | DRB5*01:01:01-01:14, DRB5*02:02-02:05 |
| 9612 | - | - |  | Negative control |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 75, 80, 90, 91 and 92.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 73 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DR low resolution typing.

In addition, well number 82 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DRB alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site, the information available at the www.worldmarrow.org web site and the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

4For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP set are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

The DRB1*08:09 and the DRB1*14:15 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:20 and the DRB1*13:18, 13:47 and 13:55 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31, 08:41 and DRB1*11:67 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*13:13 and DRB1*14:84 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 73, 77 to 83, 85, 86, 88 and 90 to 93.

7Primer mix 83 may give rise to nonspecific amplifications.

8Primer mixes 73, 84 and 94 may have tendencies of primer oligomer formation.

9Primer mix 91 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

10The DRB4*01:03:01:02N allele is amplified by the primer pairs in well No. 90 and 94, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by these primer pairs.

11Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 in addition to primer mix 93.

12Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the DRB1*07:02 allele has been shown to be identical to DRB1*07:01:01:01.

The sequence of the DRB1*08:031 allele has been shown to be identical to DRB1*08:03:02.

The sequence of the DRB1*09:011 allele has been shown to be identical to DRB1*09:01:02.

The sequence of the DRB1*11:71 allele has been shown to be identical to DRB1*11:02:01.

The sequence of the DRB1*12:031 allele has been shown to be identical to DRB1*12:01:01.

The DRB1*14:66 allele has been renamed DRB1*14:32:02.

The sequence of the DRB1*16:06 allele has been shown to be identical to DRB1*16:05:01.

The sequence of the DRB3*01:01:01 allele has been shown to be identical to DRB3*01:01:02:01.

The DRB4*0101102N allele has been shown to be identical to DRB4*01:03:10:02N.

The sequence of the DRB5*02:01 allele has been shown to be identical to DRB5*02:02.

2The DRB1*08:09 and the DRB1*14:15 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:20 and the DRB1*13:18, 13:47 and 13:55 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31, 08:41 and DRB1*11:67 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The serological reactivity of all DRB alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170.

4Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 73, 77 to 83, 85, 86, 88 and 90 to 93.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘ser’, serological HLA specificity.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information for the primer matching region(s) is not known.

Change in revision R01 compared to R00:

- Primer mix 43 does not amplify the B*15:101 allele. This has been corrected in the Specificity and Interpretation tables.