PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A-B-DR Combi Tray

Prod. No:	101.701-24/24u

101.701-06/06u

Lot No:    	91M

Expiry Date:    	2014-May-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-A low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-A serology3 | Amplified HLA-A alleles4-6 |
| --- | --- | --- | --- | --- |
| 17,8 | 120 bp, 145bp, 225 bp | 800 bp | A1, A36 | *01:01:01:01-01:04N, 01:06-01:33, 01:35-01:101, 03:18, 11:94, 36:01-36:05 |
| 28 | 210 bp, 255 bp, 365 bp, 415 bp | 800 bp | A2,A19, A28, A203, A210, A3, A11, A9, A26, A30, A68 | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:22:02, 02:24:01-02:35:01, 02:35:03-02:47, 02:49-02:77, 02:78w, 02:79:01-02:97:02, 02:99, 02:101:01-02:128, 02:130-02:323, 02:325-02:326, 03:01:03w, 03:09w, 03:23:01w, 03:89w, 03:108w, 11:06w, 11:18w, 24:28w, 24:89w, 26:03:01-26:03:02w, 26:06w, 26:07:02w, 26:21w, 26:30w, 30:13w, 30:16w, 30:44w, 30:46w, 68:05w, 68:15w, 68:20w, 68:30w |
| 39 | 205 bp, 235 bp | 1070 bp | A1, A3, A11, A32, A34, A36 | *01:12, 01:19, 01:21, 03:01:01:01-03:17, 03:19-03:74, 03:76-03:94, 03:96-03:132, 11:25, 11:60, 24:92, 32:04, 33:49w, 34:02:01, 34:02:02w, 34:03-34:04, 34:07-34:09, 36:02 |
| 4 | 190 bp | 800 bp | A1, A3, A11, A30, A36, A68 | *01:01:01:01-01:01:22, 01:01:24-01:04N, 01:06-01:07, 01:09-01:11N, 01:13, 01:16N-01:18N, 01:20-01:29, 01:31N-01:33, 01:35-01:78, 01:80-01:98, 01:100-01:101, 02:78, 02:169, 03:12, 03:18, 03:88, 11:01:01-11:27, 11:29-11:106, 26:19, 30:08, 36:04, 68:13, 68:66 |
| 514 | 160 bp, 535 bp | 800 bp | A3, A9, A23, A24, A2403, A31, A32 | *01:95, 03:30, 23:01:01-23:46, 24:02:01:01-24:11N, 24:13:01-24:15, 24:17-24:64, 24:66-24:88, 24:90N-24:128, 24:130-24:182, 31:08, 32:05, 32:13 |
| 610 | 135 bp, 175 bp, 210 bp | 800 bp | A9, A23, A24, A29, A80 | *23:01:01-23:46, 24:05, 24:13:02, 24:24, 29:07, 31:29, 80:01-80:02, B*18:27 |
| 7 | 175 bp, 205 bp | 1070 bp | A2, A3, A9, A23, A24, A2403, A26 | *02:17:01w-02:17:02w, 23:14, 24:02:01:01-24:11N, 24:13:01-24:13:02, 24:17-24:50, 24:54-24:56, 24:58-24:63, 24:66-24:91, 24:93, 24:95-24:113, 24:115-24:137, 24:139-24:182, 26:16, 33:19, 68:45 |
| 8 | 165 bp, 200 bp | 800 bp | A2, A3, A10, A11, A25, A26, A28, A32, A34, A66, A68, A69 | *01:51, 02:55, 03:24, 03:50, 11:10, 25:01:01-25:16, 26:01:01-26:06, 26:08-26:15, 26:17-26:18, 26:20-26:43:02, 26:45-26:63, 26:65-26:69, 29:28, 32:15, 33:51, 34:01:01-34:09, 66:01-66:16, 68:01:01:01-68:83, 69:01 |
| 97 | 75 bp | 800 bp | A3, A25, A32 | *25:01:01-25:16, 32:01:01-32:02, 32:04, 32:06-32:36 |
| 107 | 85 bp | 1070 bp | A10, A26 | *01:51, 01:83, 02:146, 26:01:01-26:02, 26:04, 26:07:01-26:18, 26:20, 26:22-26:29, 26:31-26:43:02, 26:45-26:69, 33:13, 33:48 |
| 117,8 | 80 bp, 175 bp, 500 bp | 1070 bp | A1, A9, A10, A11, A26, A31, A34, A66 | *01:13, 01:28, 03:63, 03:88, 11:01:01-11:27, 11:29-11:106, 24:19, 24:44, 26:03:01-26:03:02, 26:06, 26:21, 31:03, 34:01:01-34:08, 66:01, 66:04-66:11, 66:13-66:15, 80:02 |
| 12 | 185 bp | 800 bp | A3, A10, A25, A26, A31, A34, A43, A66 | *02:309, 03:01:19, 11:11, 25:05-25:06, 26:09, 26:54, 31:03-31:04, 34:01:01-34:09, 43:01, 66:02-66:03, 66:16 |
| 13 | 175 bp, 225 bp | 1070 bp | A1, A2, A3, A10, A25, A26, A34, A43, A66 | *01:13, 02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:103, 02:135, 03:01:01:01-03:01:22, 03:01:24-03:07, 03:09-03:11N, 03:13-03:31, 03:33-03:35, 03:37-03:40, 03:42-03:56, 03:58, 03:60-03:71, 03:73-03:87, 03:90-03:106, 03:109-03:110, 03:112-03:132, 25:01:01-25:05, 25:07-25:16, 26:01:01-26:01:20, 26:01:22, 26:02w, 26:03:01-26:03:02, 26:05-26:08, 26:10-26:33, 26:35-26:43:02, 26:45-26:69, 30:55, 34:08, 43:01, 66:01, 66:04-66:15, 68:71, 74:13 |
| 147 | 100 bp, 200 bp, 240 bp | 1070 bp | A2, A29 | *02:237, 02:309, 03:95, 26:19, 26:22, 29:01:01:01-29:31, 31:03-31:04, 33:13, 33:48, 34:04, 66:09 |
| 157,8,12,13 | 90 bp, 135 bp, 205 bp | 1070 bp | A1, A30 | *01:13, 01:28, 03:43, 03:82, 30:01:01-30:04:02, 30:06-30:20, 30:22-30:55, 31:35 |
| 16 | 240 bp, 370 bp, 395 bp | 1070 bp | A2, A24, A31, A32 | *02:237, 03:95, 29:14, 31:01:02-31:54, 32:05 |
| 17 | 140 bp, 180 bp | 1070 bp | A32 | *01:95, 03:43, 03:82, 29:13, 31:35, 32:01:01-32:03, 32:05-32:36, 74:07 |
| 18 | 200 bp | 1070 bp | A33, A68 | *02:243, 33:01:01-33:01:06, 33:03:01-33:50, 68:29 |
| 1912 | 160 bp, 200 bp | 800 bp | A74 | *29:19, 74:01-74:15 |
| 2010 | 220 bp, 245 bp | 800 bp | A2, A210, A25, A68 | *02:34-02:35:03, 02:46, 02:48, 02:56:01-02:56:02, 02:62, 02:70, 02:78, 02:103, 02:129, 25:05, 26:54, 68:01:01:01-68:83 |
| 21 | 240 bp, 375 bp | 800 bp | A2, A26, A28, A68, A69 | *02:55, 02:243, 24:82, 26:22, 33:22, 66:09, 68:29, 69:01 |
| 227,11 | 85 bp, 240 bp | 800 bp | A2, A36 | *02:34-02:35:03, 02:46, 02:48, 02:56:01-02:56:02, 02:62, 02:70, 02:78, 02:103, 02:129, 36:01-36:05 |
| 237,10 | 75 bp, 160 bp, 240 bp, 495 bp | 800 bp | A2, A26, A28, A36, A68, A80 | *02:55, 03:41, 03:63, 03:75, 03:88, 24:18, 26:03:01-26:03:02, 26:05-26:06, 26:21, 26:30, 33:24, 36:02, 68:05, 68:15, 68:20, 80:01 |
| 2412 | 360 bp | 1070 bp | A10, A26, A31, A66 | *02:135, 02:309, 03:01:19, 25:13, 26:30, 26:65, 31:04, 34:09, 66:02-66:03, 66:12, 66:16 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A low resolution typing.

In addition, wells 2, 4, 5, 6, 8, 9, 12 and 19 to 23 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-A alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4For several HLA-A alleles only partial 1st exon nucleotide sequences are available. We assume that unknown sequences are conserved within allelic groups.

Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

5The A*23:14, 24:05, 24:13:02 and 24:24 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

6Primer mix 6 will amplify the B*18:27 allele.

7Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

8The primer pairs in wells 1, 2, 11 and 15 will in many samples give rise to two or three HLA-specific PCR fragments.

9Primer mix 3 may faintly amplify the A*30:04, 30:06, 30:17 and 30:29 alleles.

10Primer mixes 6, 20 and 23 may have a tendency of primer dimer formation.

11Primer mix 22 might faintly amplify most A*11 alleles.

12Primer mixes 15 and 24 may yield less specific PCR product than the other HLA-A low primer mixes.

13Primer mix 15 may have tendencies of unspecific amplifications.

14In primer mix 5, the specific PCR product of 535 base pairs may be difficult to distinguish from the internal control band. The alleles giving rise to a product of this size are the following: A*01:95, 23:09, 24:02:06, 24:02:27, 24:08, 24:24, 24:29 , 24:42, 24:67, 24:116, 24:137, 24:140,

‘w’, might be weakly amplified.

1The sequence of the A*0105N has been shown to be identical to A*01:04N.

The A*01:34N allele has been renamed A*01:01:38L.

The A*020116 allele has been renamed to A*02:134.

The A*020120 allele has been shown to be identical to A*02:01:18.

The sequence of the A*0223 allele has been shown to be identical to A*02:22:01.

The sequence of the A*0298 allele has been shown to be identical to A*02:96.

The A*02:100 allele has never been assigned.

The A*1128 allele has been renamed to A*11:15:02.

The sequence of the A*2401 allele has been shown to be in error.

The sequence of the A*2412 allele has been shown to be identical to A*24:08.

The A*2416 allele has been renamed to A*31:08.

The A*2465 allele has been renamed to A*24:13:02.

The A*26:44 allele has been renamed to A*26:43:02.

The sequence of the A*3005 allele has been shown to be identical to A*30:04.

The A*3021 allele has been renamed to A*30:11:02.

The sequence of the A*31011 allele has been shown to be identical to A*31:01:02.

The sequence of the A*3302 allele has been shown to be identical to A*33:03:01.

2The serological reactivity of all HLA-A alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

3Primer mixes 1, 2, 11 and 15 will in many samples give rise to two or three HLA-specific PCR fragments.

4The A*23:14 and the A*24:05, 24:13:02 and 24:24 alleles will give rise to identical amplification patterns. These four alleles can be separated by the respective high resolution SSP primer sets.

‘w’, may be weakly amplified.

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-B low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-B serol.3 | Amplified HLA-B alleles4,5 |
| --- | --- | --- | --- | --- |
| 256,10 | 110 bp | 800 bp | 7, 703, 40, 41, 42, 48, 61 | *07:02:01-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:47, 07:49N-07:50, 07:52, 07:54-07:59, 07:61-07:99, 07:101-07:121, 07:123-07:138, 15:138, 15:230, 37:07, 40:15-40:16, 40:23, 40:32, 40:98, 40:136, 40:158, 41:08, 42:05:01-42:05:02, 48:05, 48:08, 48:15, C*02:23, C*04:77 |
| 26 | 215 bp | 1070 bp | 8, 44 | *08:01:01-08:05, 08:07-08:25, 08:27-08:49, 08:51-08:61, 08:63-08:64, 08:66-08:77, 15:142, 15:180, 44:49, 51:68 |
| 27 | 140 bp, 235 bp | 1070 bp | 7, 8, 13, 15, 35, 4005, 44, 49, 61, 62, 77 | *07:20, 07:24, 07:60, 07:100, 07:131, 08:21, 08:25, 13:01:01-13:04, 13:06-13:08Q, 13:10-13:23, 13:25-13:38, 13:40-13:50, 15:07:01-15:07:02, 15:36w, 15:45, 15:55, 15:68, 15:89w, 15:126, 15:205, 15:207, 35:05:01-35:05:02, 35:16-35:17, 35:22, 35:30, 35:51, 35:58, 35:72, 35:89, 35:97, 35:113-35:114, 35:125, 35:164, 40:05, 40:71, 44:08w, 44:54, 44:57w, 44:60, 44:106, 44:110, 46:12, 46:20, 51:64, 53:14, 55:51, 58:18, C*03:05, C*03:25, C*03:27 |
| 287,8 | 130 bp, 265 bp | 800 bp | 12, 13, 14, 17, 21, 35, 40, 41, 44, 45, 47, 49, 50, 60, 61, 64, 65 | *13:01:01-13:04, 13:06-13:13, 13:15-13:23, 13:25-13:50, 14:01:01-14:04, 14:07N, 14:09, 14:11-14:12, 14:14-14:27, 15:46, 15:53, 15:106, 15:212, 18:44, 35:46-35:47, 35:63, 35:154, 40:01:01-40:01:23, 40:07, 40:10:01-40:10:02, 40:14:01-40:16, 40:22N-40:23, 40:25, 40:30-40:34, 40:36, 40:38, 40:42-40:43, 40:45, 40:47-40:49, 40:51-40:55, 40:58-40:63, 40:65-40:67, 40:69, 40:72:01-40:73, 40:76-40:77, 40:79-40:81, 40:84, 40:87:01-40:88, 40:92, 40:100-40:102, 40:106, 40:108, 40:110, 40:112-40:114, 40:116-40:118N, 40:121, 40:123-40:126, 40:128-40:130, 40:132, 40:134-40:136, 40:139-40:141, 40:146-40:147, 40:150-40:156, 40:158, 40:160, 40:163-40:164, 40:166, 40:168, 40:171-40:172, 41:01-41:19, 44:02:01:01-44:03:06, 44:03:08-44:05:03, 44:09-44:39, 44:41:01-44:43:02, 44:45-44:56N, 44:58N-44:80, 44:82, 44:84-44:102, 44:104-44:110, 44:112-44:129, 44:132-44:133, 45:01-45:13, 47:01:01:01-47:08, 49:01:01-49:02, 49:04-49:20, 50:01:01-50:02, 50:04-50:06, 50:08-50:14, 57:45, 57:51, C*01:30 |
| 298 | 185 bp, 235 bp | 800 bp | 7, 16, 17, 27, 2708, 37, 38, 3901, 3902, 39, 64, 67 | *14:01:01-14:01:02, 14:07N-14:08, 14:10, 14:12, 14:14, 14:19, 14:26, 27:01-27:05:15, 27:05:17-27:06, 27:08-27:10, 27:12-27:13, 27:16-27:18, 27:20, 27:23, 27:26-27:27, 27:29, 27:31, 27:35-27:37, 27:39-27:42, 27:44-27:46, 27:48-27:61, 27:64N-27:69, 27:72-27:75, 27:77-27:80, 27:82, 37:02, 38:01:01-38:02:02, 38:03, 38:07-38:24, 38:26-38:32, 38:34N-38:35, 39:01:01:01-39:01:01:02L, 39:01:03-39:01:08, 39:01:10-39:02:01, 39:03, 39:05:01-39:09, 39:11, 39:14-39:15, 39:18, 39:19:02, 39:22, 39:24:01-39:44, 39:46-39:48, 39:50-39:62, 39:64-39:67, 47:04-47:05, 48:21, 67:01:01, 67:03 |
| 30 | 190 bp | 800 bp | 14, 35, 38, 39, 65 | *07:28, 14:02:01-14:02:02, 14:02:04-14:02:05, 14:03-14:06:02, 14:09, 14:11, 14:13, 14:15-14:18, 14:20, 14:22-14:23, 14:25, 14:27, 15:77, 15:189, 15:233, 35:26, 38:05, 38:33, 39:04, 44:16, 44:37, 44:64:01-44:64:02, 44:91, 44:132, 52:26, 57:04, A*23:31, A*24:106, C*16:10 |
| 31 | 290 bp | 1070 bp | 15, 22, 62, 63, 71, 72, 75, 76, 77, 46 | *15:01:01:01-15:01:03, 15:01:06-15:02:05, 15:03:03-15:08, 15:10:02-15:11:03, 15:11:05-15:15, 15:17:01:01-15:17:02, 15:19-15:21, 15:24-15:28, 15:30-15:36, 15:38:01-15:40, 15:42-15:46, 15:48, 15:50, 15:55-15:58, 15:60, 15:63, 15:65-15:66, 15:70-15:71, 15:73, 15:75-15:79N, 15:81-15:89, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:107, 15:109-15:113, 15:116-15:118, 15:120-15:122, 15:125-15:126, 15:128-15:129, 15:135-15:150, 15:152, 15:154-15:155, 15:157, 15:159-15:160, 15:162-15:172, 15:174-15:175, 15:177-15:179, 15:181N-15:185, 15:187-15:196, 15:199, 15:201-15:209N, 15:211, 15:213-15:219, 15:223-15:225, 15:227-15:228, 15:230-15:234, 46:01:01-46:28 |
| 32 | 165 bp, 220 bp, 330 bp | 1070 bp | 5, 8, 12, 21, 22, 37, 41, 42, 44, 45, 48, 51, 56, 57, 60, 62, 70, 71, 72, 82 | *08:01:01-08:05, 08:08N-08:12:03, 08:15-08:19N, 08:21-08:24, 08:26-08:27, 08:29-08:36, 08:38-08:39, 08:41-08:48, 08:50-08:54, 08:56-08:69, 08:71-08:73, 08:75-08:76, 13:46, 15:03:01-15:03:03, 15:09-15:10:02, 15:18:01-15:18:04, 15:23, 15:29, 15:37, 15:46-15:47, 15:49, 15:51-15:54, 15:61-15:62, 15:64, 15:69, 15:72, 15:74, 15:80, 15:83, 15:90-15:91, 15:93, 15:98-15:99, 15:103, 15:106, 15:108, 15:114-15:115, 15:119, 15:123-15:124, 15:127, 15:131-15:134, 15:143, 15:151, 15:153, 15:156, 15:158, 15:161, 15:173, 15:176, 15:186, 15:197-15:198, 15:200, 15:210, 15:212, 15:220-15:221, 15:226N, 15:229, 35:87, 37:12, 38:30, 40:12, 40:136, 41:01-41:03:02, 41:05-41:09, 41:11-41:17, 41:19, 42:01:01-42:02, 42:04-42:08, 42:10-42:15, 44:14-44:15, 44:18, 44:20, 44:62, 44:100, 45:01-45:13, 51:08, 51:20, 51:36, 51:44N, 51:97, 52:19, 55:20, 56:13, 57:09, 82:01-82:03, C*07:46 |
| 338,9,11 | 165 bp, 190 bp, 390 bp | 1070 bp | 5, 17, 21, 35, 51, 5102, 5103, 52, 56, 62, 63, 70, 78 | *15:01:02, 15:09, 15:16:01-15:17:02, 15:67, 15:95, 15:162, 15:168, 15:177, 15:196, 15:208, 15:216, 15:222, 15:230, 35:01:10, 35:04:02, 40:26, 40:28, 44:62, 49:18, 50:14, 51:01:01-51:01:03, 51:01:05-51:02:03, 51:02:05-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03-51:24:04, 51:26-51:33, 51:37-51:41N, 51:43-51:44N, 51:46, 51:48-51:61, 51:63-51:80, 51:82-51:91, 51:94-51:98N, 51:100-51:122, 52:01:01:01-52:01:04, 52:01:06-52:13, 52:15-52:24, 52:26, 56:05:01-56:06, 56:21, 58:08:01-58:08:02, 78:01:01-78:07 |
| 34 | 180 bp | 800 bp | 13, 22, 55, 60 | *07:78, 13:01:01-13:02:12, 13:07N-13:09, 13:11, 13:14-13:20, 13:22:01-13:23, 13:25, 13:27-13:34, 13:36-13:45, 13:47, 13:49N-13:50, 40:48, 45:10, 49:07, 55:09, 55:22, 55:24 |
| 356 | 105 bp, 195 bp | 1070 bp | 8, 12, 27, 38, 39, 3902, 40, 44, 45, 48, 60, 70, 71, 72, 82 | *07:27, 07:50, 08:04, 08:17, 08:54, 15:03:01-15:03:03, 15:47, 15:49, 15:54, 15:61-15:62, 15:64, 15:68-15:69, 15:91, 15:98, 15:103, 15:123, 15:127, 15:131-15:132, 15:151, 15:156, 15:158, 15:173, 15:210, 15:220, 18:37, 27:18, 27:29, 37:28, 38:03, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:23, 39:39, 39:49, 40:12, 40:46, 40:93, 42:11, 44:10, 44:15, 44:18, 44:40, 44:44, 44:130, 45:01, 45:05-45:07, 45:11-45:13, 48:01:01-48:05, 48:07-48:25, 49:20, 50:02, 52:16, 55:18, 82:01-82:03 |
| 36 | 280 bp | 1070 bp | 5, 7, 8, 13, 15, 16, 17, 18, 22, 35, 40, 42, 44, 46, 48, 49, 51, 53, 55, 56, 57, 58, 60, 61, 62, 63, 70, 71, 72, 75, 76, 77, 78 | *07:09, 07:11, 07:17, 08:28, 08:35, 08:37, 13:04, 13:10, 13:26, 15:01:01:01-15:01:04, 15:01:06-15:08, 15:11:01-15:16:03, 15:18:01-15:21, 15:23-15:29, 15:31-15:36, 15:38:01-15:40, 15:43-15:44, 15:46-15:47, 15:49-15:57, 15:60-15:62, 15:64-15:72, 15:74-15:76, 15:78:01-15:82, 15:84-15:85, 15:87-15:89, 15:91-15:98, 15:101-15:129, 15:131-15:132, 15:134-15:136, 15:138-15:149N, 15:151-15:161, 15:163-15:167, 15:169-15:176, 15:178-15:187, 15:189-15:195, 15:197-15:207, 15:209N-15:215, 15:217-15:223, 15:225-15:228, 15:231-15:232, 15:234, 18:01:01-18:15, 18:17N-18:25, 18:27-18:40, 18:42-18:60, 18:62-18:66, 27:41, 35:01:01:01-35:01:29, 35:05:01-35:05:02, 35:07-35:08:04, 35:10-35:11:02, 35:14:01-35:17, 35:19-35:21, 35:23-35:30, 35:32, 35:35, 35:37, 35:40N-35:43:02, 35:45-35:54, 35:57-35:58, 35:61-35:64, 35:66-35:69, 35:71-35:72, 35:76-35:80, 35:82, 35:86, 35:89-35:94, 35:97, 35:99-35:105, 35:107-35:108:02, 35:110-35:126, 35:130N-35:135, 35:137-35:148, 35:158-35:159, 35:161, 35:165N-35:166, 35:168, 35:170-35:171, 35:173N-35:178, 39:07, 39:43, 40:03, 40:20, 40:38, 40:52, 40:59-40:60, 40:105, 40:158-40:159, 42:09, 44:17, 44:43:01-44:43:02, 45:09, 46:01:01-46:10, 46:12-46:17, 46:19-46:28, 48:02:01-48:02:02, 48:14, 48:23, 48:25, 49:04-49:05, 51:37, 51:45, 51:63, 51:97, 53:01:01-53:03, 53:05-53:06, 53:08:01-53:18, 53:20-53:26, 54:06, 54:09, 54:14, 55:14, 55:23, 55:32, 56:03, 56:09, 56:18, 56:31-56:32, 57:01:01-57:01:11, 57:06, 57:08, 57:10-57:11, 57:13-57:16, 57:18-57:27, 57:29-57:31, 57:33-57:38, 57:40-57:41, 57:43-57:45, 57:47-57:52, 58:01:01-58:02, 58:04-58:07, 58:09-58:19, 58:21-58:26, 58:29-58:33 |
| 37 | 195 bp | 1070 bp | 27, 35, 37, 44, 58, 76 | *15:12, 15:14, 15:19, 15:91, 15:131, 15:161, 18:54, 18:56, 35:45, 35:71, 37:01:01-37:09, 37:12-37:13, 37:15-37:21, 37:23-37:28, 37:30N-37:31, 38:17, 44:17, 44:43:01-44:43:02, 45:09, 46:17, 53:22, 58:07, A*24:174 |
| 386 | 105 bp, 395 bp, 435 bp | 1070 bp | 5, 7, 15, 41, 42, 62, 63, 70, 71, 75, 77 | *07:04, 07:25, 15:09-15:10:02, 15:13:01-15:13:02, 15:16:01-15:18:04, 15:21, 15:23-15:24, 15:37, 15:44, 15:51-15:52, 15:66-15:67, 15:72, 15:80, 15:87, 15:90, 15:93, 15:95, 15:99, 15:108, 15:114-15:115, 15:119, 15:124, 15:133-15:134, 15:153, 15:157, 15:161-15:162, 15:168, 15:176-15:177, 15:186, 15:189, 15:196-15:198, 15:200, 15:208, 15:216, 15:221-15:222, 15:226N, 15:229-15:230, 40:136, 41:08, 42:05:01-42:05:02 |
| 396 | 115 bp, 150 bp | 1070 bp | 18 | *18:01:01-18:15, 18:17N-18:28, 18:30-18:66 |
| 406 | 80 bp | 1070 bp | 7, 27, 2708, 44, 60 | *07:73, 27:01-27:21, 27:23-27:51, 27:53-27:66N, 27:68-27:74, 27:76-27:82, 38:22, 40:46, 40:93, 44:40, 44:44, 44:130 |
| 41 | 150 bp | 800 bp | 12, 13, 17, 18, 22, 27, 35, 37, 39, 44, 45, 47, 48, 51, 53, 56, 57, 58, 62, 70, 75, 77, 78 | *08:49, 13:01:01-13:01:06, 13:06-13:07N, 13:12-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:36, 13:39, 13:43, 13:50, 14:10, 15:02:01-15:02:05, 15:13:01-15:13:02, 15:20-15:21, 15:25:01-15:25:03, 15:36, 15:44, 15:62, 15:77, 15:80, 15:85, 15:88-15:89, 15:106, 15:112, 15:121, 15:139, 15:144, 15:154, 15:165, 15:170, 15:194, 15:204, 15:213-15:214, 15:223, 18:22, 27:19, 27:30, 35:01:01:01-35:04:03, 35:06-35:08:04, 35:10-35:17, 35:19-35:21, 35:23-35:30, 35:33-35:36, 35:38-35:42:02, 35:45-35:50, 35:52, 35:54-35:57, 35:59, 35:61-35:63, 35:65Q, 35:69-35:71, 35:74, 35:76-35:78, 35:80-35:85, 35:90-35:96, 35:98, 35:100-35:101:02, 35:103-35:113, 35:115-35:116, 35:120-35:126, 35:128-35:134N, 35:136-35:150, 35:152-35:166, 35:168-35:173N, 35:175-35:178, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:11, 37:13-37:31, 38:20, 39:42, 40:28, 44:02:01:01-44:14, 44:16-44:17, 44:19N, 44:21-44:30, 44:32-44:40, 44:42-44:46, 44:48-44:52N, 44:55-44:64:02, 44:66-44:98, 44:101-44:105, 44:107-44:134, 48:02:01-48:02:02, 48:25, 51:04, 51:42, 51:46, 51:56:01-51:56:02, 53:01:01-53:13, 53:15-53:26, 55:14, 56:09, 56:11-56:12, 57:01:01-57:01:04, 57:01:06-57:10, 57:12, 57:14-57:20, 57:22-57:30, 57:32-57:46, 57:48-57:52, 58:01:01-58:01:02, 58:01:04-58:01:09, 58:04-58:05, 58:09-58:15, 58:17N, 58:19, 58:21-58:24, 58:28-58:34, 83:01, C*03:102 |
| 42 | 135 bp | 1070 bp | 8, 18, 22, 35, 39, 78 | *07:65, 07:134, 08:32, 15:202, 18:01:01-18:08, 18:10-18:15, 18:17N-18:36, 18:38-18:47, 18:50-18:53, 18:55-18:65, 35:01:01:01-35:03:02, 35:03:04-35:18, 35:20:01-35:24:02, 35:28-35:45, 35:48, 35:50-35:55, 35:57-35:62, 35:64-35:72, 35:74-35:153, 35:155-35:166, 35:168-35:178, 37:11, 39:19:01-39:19:02, 56:06, 78:01:01-78:05, 78:07 |
| 436,10 | 60 bp, 245 bp, 400 bp | 1070 bp | 5, 15, 44, 48, 60, 62, 63, 70, 71, 72, 75, 76, 77, 82 | *15:01:01:01-15:01:04, 15:01:06-15:21, 15:23-15:40, 15:42-15:58, 15:60-15:67, 15:69-15:99, 15:102-15:129, 15:131-15:179, 15:181N-15:234, 40:12, 82:01-82:03 |
| 4410 | 210 bp | 800 bp | 27, 37 | *07:27, 08:49, 18:54, 37:01:01-37:31, 38:17 |
| 45 | 170 bp | 800 bp | 16, 35, 38, 39, 3901, 3902, 67, 72 | *08:55, 15:69, 15:186, 35:35, 38:01:01-38:09, 38:11-38:21, 38:23-38:29, 38:31-38:35, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:24:02, 39:26-39:29, 39:31, 39:35, 39:37-39:42, 39:44-39:46, 39:48-39:49, 39:51-39:67, 51:101, 58:20, 67:01:01-67:03, C*15:51 |
| 466 | 110 bp | 1070 bp | 38 | *38:01:01-38:02:02, 38:03, 38:05, 38:08-38:24, 38:26-38:34N |
| 47 | 395 bp | 1070 bp | 14, 16, 38, 39, 3901, 64, 65 | *07:69, 07:85, 08:65, 14:01:01-14:20, 14:22-14:27, 38:01:01-38:02:03, 38:04-38:05, 38:08-38:25, 38:27-38:35, 39:01:01:01-39:01:01:02L, 39:01:03-39:01:12, 39:03-39:07, 39:09, 39:11-39:12, 39:14-39:15, 39:18-39:19:02, 39:22, 39:24:01-39:38Q, 39:40N-39:48, 39:50-39:57, 39:59-39:62, 39:64-39:67 |
| 48 | 160 bp, 425 bp | 1070 bp | 7, 703, 8, 14, 15, 16, 21, 39, 3901, 3902, 40, 41, 42, 45, 48, 50, 60, 61, 64, 65, 67, 70, 73, 81 | *07:02:01-07:26, 07:28-07:35, 07:37, 07:39-07:64, 07:66-07:80, 07:82-07:139, 08:01:01-08:01:16, 08:04-08:05, 08:07-08:35, 08:37-08:51, 08:53-08:77, 14:01:01-14:02:04, 14:02:06-14:27, 27:14, 27:81, 39:01:01:01-39:01:01:02L, 39:01:03-39:20, 39:22-39:24:01, 39:25N-39:67, 40:01:01-40:01:23, 40:06:01:01-40:07, 40:09?-40:11:02?, 40:14:01?-40:15?, 40:16, 40:18?, 40:21?-40:22N?, 40:23, 40:24?-40:38?, 40:42?-40:43?, 40:44, 40:45?-40:46?, 40:48?-40:50?, 40:51, 40:52?, 40:53, 40:54?-40:58?, 40:59-40:60, 40:61?-40:65?, 40:66, 40:67?-40:69?, 40:70, 40:71?, 40:72:01-40:73, 40:74?, 40:75, 40:77, 40:78?, 40:79, 40:80?-40:82?, 40:83, 40:84?-40:85?, 40:86, 40:87:01?-40:88?, 40:90?-40:92?, 40:93, 40:95?, 40:96, 40:98?-40:102?, 40:103, 40:104?-40:108?, 40:109-40:110, 40:111?-40:116?, 40:118N?-40:121?, 40:123?, 40:124:01, 40:124:02?-40:126?, 40:127, 40:128?-40:130?, 40:131, 40:132?-40:134?, 40:136?-40:137?, 40:138-40:140, 40:141?, 40:145?, 40:146-40:148, 40:149?, 40:150, 40:151?, 40:152-40:153, 40:154?, 40:155N, 40:156?, 40:158?, 40:159, 40:160?, 40:161-40:162, 40:163?-40:164?, 40:165, 40:166?, 40:167, 40:168?-40:173?, 41:01-41:19, 42:01:01-42:02, 42:04-42:16, 45:01-45:13, 48:01:01-48:17, 48:19-48:25, 50:01:01-50:02, 50:04-50:14, 51:01:09, 51:01:21, 51:01:25, 51:10, 51:24:02-51:24:04, 54:11, 56:20:02, 67:01:01-67:03, 73:01-73:02, 81:02 |
| 4912 | 180 bp | 1070 bp | 15, 40, 44, 48, 60, 62, 71 | *15:116, 15:124, 40:01:01-40:01:06, 40:01:08-40:01:17, 40:01:19-40:01:23, 40:07, 40:10:01-40:10:02, 40:12, 40:21-40:23, 40:25, 40:30, 40:33-40:34, 40:36, 40:38, 40:42-40:43, 40:46-40:49, 40:51-40:52, 40:54-40:55, 40:59-40:63, 40:65-40:67, 40:69, 40:73-40:74, 40:76, 40:79, 40:81, 40:84, 40:87:01-40:88, 40:92, 40:100-40:102, 40:106, 40:108, 40:112-40:114, 40:116-40:118N, 40:123-40:126, 40:128, 40:130, 40:132, 40:134-40:135, 40:137-40:141, 40:146-40:147, 40:149-40:156, 40:160, 40:163, 40:166, 40:168, 40:170-40:172, 44:31, 46:06, 48:03:01-48:03:02, 48:17, 48:23 |
| 50 | 290 bp | 1070 bp | 21, 4005, 41, 50, 60, 61 | *18:48, 40:02:01-40:06:04, 40:08, 40:09?-40:11:02?, 40:13, 40:14:01?-40:15?, 40:18?-40:19?, 40:20, 40:22N?, 40:24?-40:38?, 40:39, 40:42?-40:45?, 40:47?-40:50?, 40:52?-40:58?, 40:61?-40:65?, 40:67?-40:69?, 40:70, 40:71?, 40:72:02?, 40:74?-40:76?, 40:78?, 40:80?-40:82?, 40:84?-40:88?, 40:89, 40:90?-40:92?, 40:94, 40:95?-40:96?, 40:97, 40:98?-40:121?, 40:122, 40:123?, 40:124:02?-40:136?, 40:141?, 40:142N-40:144N, 40:145?, 40:148?-40:149?, 40:151?, 40:154?, 40:156?-40:169?, 40:171?-40:173?, 47:01:01:01-47:08 |
| 516 | 105 bp | 1070 bp | 7, 8, 22, 41, 42 | *07:04, 07:25, 08:01:01-08:05, 08:07-08:12:03, 08:14-08:19N, 08:21-08:24, 08:26-08:39, 08:41-08:54, 08:56-08:69, 08:71-08:77, 35:87, 40:136, 41:01-41:08, 41:10-41:19, 42:01:01-42:02, 42:04-42:07, 42:09-42:16, 55:20 |
| 52 | 325 bp | 1070 bp | 7, 22, 27, 2708, 35, 42, 45, 46, 54, 55, 56, 67, 73, 78, 81, 82 | *07:02:01-07:02:28, 07:04-07:07, 07:09-07:15, 07:17-07:26, 07:28-07:31, 07:33-07:36, 07:39-07:46, 07:47w, 07:48-07:49N, 07:51-07:68:03, 07:70-07:84, 07:86-07:139, 14:21, 15:76, 15:101, 27:01-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:76, 27:78-27:82, 35:76, 38:26, 40:166, 42:01:01-42:02, 42:04-42:06, 42:08-42:10, 42:12-42:16, 44:90, 44:97, 45:06, 46:01:01-46:01:07, 46:02w, 46:03-46:28, 54:01:01-54:24, 55:01:01-55:05, 55:07-55:17, 55:19-55:53, 56:01:01-56:16, 56:18-56:22, 56:24-56:32, 67:01:01-67:03, 73:01-73:02, 81:01-81:05, 82:01-82:03, 83:01 |
| 536 | 115 bp, 195 bp, 225 bp, 260 bp | 1070 bp | 35, 37, 44, 47, 57 | *08:49, 08:60, 08:76, 35:38, 35:45, 35:71, 35:115, 35:169, 37:01:01-37:01:07, 37:03N-37:06, 37:08, 37:10-37:31, 40:132, 42:13, 44:02:01:01-44:14, 44:16-44:17, 44:19N-44:36, 44:38-44:63, 44:65-44:134, 51:42, 53:22, 57:07, 57:09, 57:24, 83:01, C*15:25 |
| 54 | 215 bp | 800 bp | 5, 8, 12, 17, 21, 22, 41, 42, 44, 45, 51, 5102, 5103, 52, 53, 56, 62 | *08:09, 13:46, 15:83, 41:01, 41:05-41:07, 41:09, 41:12, 41:14, 41:16-41:17, 42:04, 44:06, 44:15, 44:18, 44:20, 44:100, 45:01-45:13, 51:01:01, 51:01:03-51:01:08, 51:01:10-51:02:01, 51:02:03-51:04, 51:06:01-51:07:01, 51:08-51:09:01, 51:10-51:14, 51:16-51:24:04, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:88, 51:90-51:92, 51:94-51:102, 51:104-51:115, 51:117-51:118N, 51:120-51:122, 52:01:02, 52:01:04, 52:01:09-52:03, 52:06:01-52:06:02, 52:09, 52:19, 52:21, 53:06, 55:20, 56:13, 58:08:01 |
| 55 | 130 bp, 270 bp | 1070 bp | 17, 18, 22, 35, 37, 46, 51, 53, 62, 78 | *14:10, 15:57w, 18:22, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 37:04:01-37:04:02, 40:28, 46:01:01-46:28, 51:04, 51:46, 51:56:01-51:56:02, 53:02, 53:06, 57:14, 58:09, C*15:39 |
| 566,13 | 90 bp, 410 bp | 1070 bp | 15, 22, 27, 35, 47, 54, 55, 56, 78, 81, 82 | *07:65, 27:01w, 27:02:01-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:76, 27:78-27:82, 35:76, 44:90, 44:97, 47:01:01:01-47:03, 47:06-47:08, 54:01:01-54:24, 55:01:01-55:05, 55:07-55:17, 55:19-55:53, 56:01:01-56:16, 56:18-56:22, 56:24-56:32, 81:01, 82:01-82:03, 83:01 |
| 576 | 90 bp, 175 bp | 1070 bp | 15, 27, 35, 48, 60, 62, 75, 77, 81 | *13:36, 15:02:01-15:02:05, 15:08, 15:11:01-15:11:05, 15:13:01-15:13:02, 15:15, 15:21, 15:31, 15:44, 15:55, 15:76w, 15:88-15:89, 15:112, 15:121, 15:139, 15:144, 15:148, 15:170, 15:189, 15:191, 15:194, 15:209N, 15:213-15:215, 15:223, 27:24, 35:46, 40:31, 40:45, 40:80, 48:01:01-48:01:03, 48:04, 48:06-48:07, 48:09, 48:11, 48:15-48:16, 48:18-48:20, 48:22, 48:24, 81:01-81:05 |
| 58 | 145 bp, 430 bp | 1070 bp | 44, 49, 59, 61 | *40:13, 40:19, 40:109, 40:117, 44:18, 44:25, 44:50, 44:95, 49:01:01-49:01:03, 49:04-49:20, 51:112, 54:12, 56:21, 57:45, 57:51, 59:01:01:01-59:05 |
| 596 | 120 bp, 210 bp | 800 bp | 5, 13, 15, 17, 22, 35, 45, 49, 50, 51, 5102, 5103, 52, 55, 56, 62, 63, 78 | *07:78, 07:84, 13:16, 13:31w, 13:48, 15:04w, 15:16:01-15:16:03, 15:42, 15:67, 15:95, 15:137w, 40:95, 40:148, 40:161, 49:01:01, 49:01:03-49:10, 49:12-49:20, 50:01:01-50:02, 50:04-50:11, 50:13-50:14, 51:01:01-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:14, 51:16-51:24:04, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60, 51:61w, 51:63, 51:65-51:80, 51:82-51:122, 52:01:01:01-52:26, 54:20, 55:01:01-55:01:06, 55:03, 55:05, 55:09, 55:11, 55:15, 55:17, 55:21w, 55:24-55:25, 55:28-55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45, 55:52-55:53, 56:05:01-56:06, 56:21, 56:25, 58:08:01-58:08:02, 78:01:01-78:07 |
| 60 | 430 bp | 1070 bp | 5, 17, 27, 44, 51, 5102, 5103, 52, 53, 57, 58, 61 | *27:02:01-27:02:02, 27:30, 27:53, 27:57, 27:62, 27:65N, 27:75, 27:77, 40:13, 40:19, 40:109, 40:117, 44:06, 44:25, 44:50, 44:95, 51:01:01-51:24:04, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:122, 52:01:01:01-52:19, 52:21-52:26, 53:01:01-53:02, 53:04-53:08:02, 53:10, 53:14-53:26, 57:01:01-57:11, 57:13-57:52, 58:01:01-58:02, 58:04-58:16, 58:18-58:34 |
| 61 | 145 bp | 1070 bp | 12, 21, 35, 40, 4005, 41, 45, 47, 50, 60, 61 | *07:133, 15:46, 15:53, 15:106, 15:143, 15:212, 18:48, 35:19, 35:47, 35:63, 35:154, 40:01:01-40:11:02, 40:14:01-40:16, 40:18, 40:20, 40:22N-40:40, 40:42-40:45, 40:48-40:75, 40:77-40:92, 40:94-40:95, 40:97-40:108, 40:111-40:116, 40:118N-40:136, 40:138-40:156, 40:158-40:173, 41:01-41:04, 41:05?, 41:06-41:19, 44:09, 44:46, 44:75, 44:90, 44:129, 44:131, 45:01-45:13, 47:02, 47:03w, 50:01:01-50:02, 50:04-50:05, 50:07-50:14 |
| 62 | 300 bp | 1070 bp | 7, 12, 13, 15, 17, 18, 21, 27, 2708, 35, 37, 38, 39, 3902, 40, 4005, 41, 44, 45, 47, 48, 49, 50, 52, 60, 61, 62, 70, 72, 76 | *07:54, 07:123, 08:17, 08:38, 08:54, 13:01:01-13:04, 13:06, 13:08Q-13:23, 13:25-13:35, 13:37-13:50, 15:01:01:01-15:01:04, 15:01:06-15:01:18, 15:01:20-15:01:24, 15:03:01-15:07:02, 15:12, 15:14, 15:19-15:20, 15:24-15:27:03, 15:30, 15:32-15:36, 15:38:01-15:40, 15:42-15:43, 15:45-15:50, 15:53-15:54, 15:56-15:58, 15:60-15:63, 15:65-15:66, 15:68-15:71, 15:73-15:75, 15:77-15:79N, 15:81-15:87, 15:91-15:92, 15:94N, 15:96-15:98, 15:101-15:107, 15:109-15:111N, 15:113, 15:116-15:118, 15:122-15:123, 15:125-15:129, 15:131-15:132, 15:135-15:138, 15:140-15:142, 15:145-15:147, 15:150-15:152, 15:154-15:160, 15:163-15:167, 15:169, 15:171-15:175, 15:178-15:179, 15:181N-15:185, 15:187-15:188, 15:190N, 15:192-15:193, 15:195, 15:199, 15:201-15:207, 15:210-15:212, 15:217-15:220, 15:224-15:225, 15:227-15:228, 15:231-15:234, 18:12, 27:01-27:21, 27:24-27:82, 35:10, 35:13, 35:16, 35:28, 35:63, 35:69, 35:80, 37:01:01-37:07, 37:09-37:31, 38:03-38:04, 38:25, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:22-39:23, 39:49, 40:01:01-40:07, 40:09-40:16, 40:18-40:24, 40:26-40:40, 40:42-40:67, 40:70-40:105, 40:107-40:165, 40:167-40:173, 41:01-41:19, 42:11, 44:02:01:01-44:05:03, 44:07-44:11, 44:13-44:134, 45:01-45:13, 47:01:01:01-47:08, 48:01:01-48:05, 48:07-48:15, 48:17-48:25, 49:01:01-49:20, 50:01:01-50:02, 50:04-50:14, 52:01:01:01-52:26, 53:17:01-53:17:02, 55:18, 55:34, 55:40, 78:05-78:06 |
| 638 | 160 bp, 330 bp | 1070 bp | 22, 44, 45, 47, 54 | *44:02:01:01-44:02:21, 44:06, 44:08-44:09, 44:11-44:12, 44:16, 44:19N-44:24, 44:27:01-44:27:02, 44:33-44:34, 44:41:01-44:41:02, 44:44, 44:48-44:49, 44:52N-44:53, 44:55, 44:59, 44:63, 44:66-44:68, 44:71-44:74, 44:80, 44:83-44:84, 44:86-44:87, 44:89-44:91, 44:93, 44:95, 44:97, 44:99-44:102, 44:104, 44:106, 44:112-44:113, 44:116, 44:118-44:119, 44:121, 44:126-44:127, 44:131-44:132, 54:01:01, 54:02-54:24, 83:01 |
| 64 | 180 bp, 210 bp | 1070 bp | 13, 15, 18, 35, 39, 40, 54, 55, 56, 59, 78 | *07:65, 13:06, 15:42, 15:86, 15:224, 18:04, 35:42:02, 35:43:02, 35:60, 39:17, 39:63, 40:01:11, 40:58, 45:08, 46:18, 54:01:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:16-54:24, 55:01:01-55:03, 55:05, 55:07, 55:10-55:12, 55:15-55:16, 55:18-55:19, 55:21, 55:25-55:26, 55:29-55:31, 55:33-55:45, 55:47-55:48, 55:50, 55:52-55:53, 56:05:01, 56:06, 56:10, 56:23, 59:01:01:01-59:01:01:02, 59:05, 78:01:01-78:01:02, 78:02:02-78:03, 78:07, C*15:02:04 |
| 65 | 180 bp | 1070 bp | 12, 15, 21, 22, 44, 45, 49, 50, 51, 56, 61, 62, 82 | *13:03, 13:48, 15:73, 40:71, 44:10, 44:15, 44:18, 45:01, 45:04-45:07, 45:11-45:13, 46:11, 49:01:01-49:03, 49:06, 49:08-49:17, 49:19N-49:20, 50:01:01-50:02, 50:04-50:08, 50:10-50:13, 51:15, 51:62, 51:106, 52:25, 54:03, 56:01:01-56:02, 56:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:20:01-56:20:02, 56:24-56:30, 59:04, 82:01-82:03, C*03:12, C*03:19, C*03:102 |
| 666 | 90 bp, 240 bp | 800 bp | 57 | *55:14, 57:01:01-57:44, 57:46-57:50, 57:52 |
| 676,13 | 90 bp | 1070 bp | 5, 13, 15, 17, 39, 55, 56, 58, 60, 61, 62, 71 | *13:11, 15:18:03, 15:73, 15:224, 39:17, 40:48, 40:71, 51:62, 52:25, 56:02, 56:04, 56:10, 58:01:01-58:02, 58:04-58:34, C*02:06w, C*02:47w |
| 686 | 95 bp | 1070 bp | 8, 18, 22, 35, 37, 38, 39, 44, 51, 5102, 5103, 53, 78 | *07:65w, 07:134w, 08:32, 18:01:01-18:11, 18:13-18:15, 18:17N-18:36, 18:38-18:47, 18:49-18:65, 35:01:01:01-35:09:02, 35:11:01-35:12:03, 35:14:01-35:15, 35:17-35:18, 35:20:01-35:24:02, 35:27, 35:29-35:45, 35:48, 35:50-35:62, 35:64-35:68:02, 35:70-35:72, 35:74-35:75, 35:76w, 35:77-35:79, 35:81-35:153, 35:155-35:166, 35:168-35:178, 37:08, 38:06-38:07, 39:19:01-39:19:02, 44:06, 51:01:01-51:24:04, 51:26-51:46, 51:48-51:103, 51:105-51:111, 51:113-51:117, 51:119-51:122, 53:01:01-53:16, 53:18-53:26, 56:06w, 78:01:01-78:04, 78:07 |
| 696 | 115 bp, 150 bp | 1070 bp | 15, 18,22, 27, 35, 39, 46, 55, 62, 72, 73, 75, 76 | *07:100, 13:31, 13:41, 15:01:01:01-15:01:04, 15:01:06-15:01:16, 15:01:18-15:01:24, 15:04, 15:07:01-15:08, 15:11:01-15:12, 15:14-15:15, 15:19, 15:24, 15:26N-15:28, 15:30, 15:32, 15:34-15:35, 15:38:01-15:38:02, 15:43, 15:45-15:46, 15:50, 15:53-15:54, 15:56-15:58, 15:60, 15:63, 15:66, 15:68, 15:70-15:71, 15:73, 15:75-15:77, 15:79N, 15:81-15:82, 15:85, 15:87, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:105, 15:109-15:111N, 15:113, 15:117-15:118, 15:120, 15:122, 15:125-15:126, 15:128-15:129, 15:135, 15:137, 15:140, 15:142-15:149N, 15:152, 15:154, 15:157, 15:159-15:160, 15:163-15:167, 15:169, 15:171-15:172, 15:174-15:175, 15:178, 15:180-15:184, 15:187, 15:189-15:193, 15:201-15:203, 15:205-15:207, 15:209N, 15:211-15:212, 15:215, 15:217, 15:225, 15:227-15:228, 15:231-15:234, 18:19, 27:25, 27:75, 35:14:01-35:14:02, 35:43:01-35:44, 35:62, 35:67, 35:79, 35:86, 35:102, 35:117, 35:135, 35:163, 39:18, 39:36, 46:01:01-46:02, 46:04-46:05, 46:07N, 46:09-46:10, 46:12, 46:14-46:17, 46:20, 46:22-46:24, 46:27-46:28, 51:61, 52:21, 54:06, 55:21, 56:03, 73:01-73:02, A*26:68, A*68:56, C*06:20, C*12:50 |
| 7014 | 360 bp | 1070 bp | Bw4 |  |
| 71 | 350 bp | 1070 bp | Bw6 |  |
| 72 | 285 bp | 800 bp | 7, 703, 8, 35, 40, 41, 42, 48, 53, 60, 61, 81 | *07:02:01-07:06, 07:08-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:52, 07:54-07:59, 07:61-07:76, 07:79-07:83, 07:85-07:99, 07:101-07:121, 07:123-07:138, 08:01:01-08:05, 08:07-08:08N, 08:10-08:11, 08:13-08:15, 08:17-08:77, 35:66, 35:87, 37:07, 40:15-40:16, 40:30-40:32, 40:34, 40:45, 40:59, 40:80, 40:98, 40:137, 40:160, 41:02:01-41:02:04, 41:04, 41:10-41:11, 41:13, 41:18-41:19, 42:01:01-42:02, 42:05:01-42:07, 42:09-42:13, 42:15-42:16, 48:01:01-48:01:03, 48:05-48:12, 48:14-48:20, 48:22, 53:15, 81:01-81:05 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 25 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B low resolution typing.

In addition, wells number 28 to 30, 34, 41, 44, 45, 54, 59, 66 and 72 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-B alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4Nucleotide sequence information is available for only exons 2 and 3 of many HLA Class I alleles and for only exon 2 of many HLA Class II alleles and not for other exons or for the introns of these alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

The B*08:26, 08:50 and 08:62 and B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:47 and 39:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*18:29 and the B*35:32, 35:37, 35:53N, 35:64, 35:68:01-35:68:02, 35:99, 35:118-35:119 and 35:174 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*41:09 and the B*45:02 and 45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*54:01:02 and the B*55:01:07-55:02:06, 55:07, 55:10, 55:12, 55:16, 55:19, 55:26, 55:30, 55:35, 55:37, 55:39, 55:41-55:43, 55:47-55:48 and 55:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:04, 55:08, 55:13, 55:27, 55:46 and 55:49 and the B*56:15, 56:19N and 56:22 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 and 56:31-56:32 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5The C*02:23 and C*04:77 alleles will be amplified by primer mix 25, the C*03:05, 03:25 and 03:27 alleles will be amplified by primer mix 27, the C*01:30 allele will be amplified by primer mix 28, the A*23:31, A*24:106 and C*16:10 alleles will be amplified by primer mix 30, the C*07:46 allele will be amplified by primer mix 32, the A*24:174 allele will be amplified by primer mix 37, the C*03:102 allele will be amplified by primer mixes 41 and 65, the C*15:51 allele will amplified by primer mix 45, the C*15:25 allele will be amplified by primer mix 53, the C*15:39 allele will be amplified by primer mix 55, the C*15:02:04 allele will amplified by primer mix 64, the C*03:12 and 03:19 alleles will be amplified by primer mix 65, the C*02:06 and 02:47 alleles will be weakly amplified by primer mix 67 and the A*26:68, A*68:56, C*06:20 and C*12:50 alleles will be amplified by primer mix 69.

6Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

7Primer mix 28 may yield less specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

8Primer mixes 28, 29, 33 and 63 give a lower yield of specific PCR product than the other HLA-B low resolution primer mixes.

9Primer mix 33 has pronounced tendencies of giving rise to nonspecific amplifications.

10Primer mixes 25, 43 and 44 may give rise to a primer oligomer artifact.

11The B*57 and B*58 alleles might be faintly amplified by primer mix 33.

12The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 49.

13Primer mixes 56 and 67 may generate a false positive band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

14The Bw4-associated HLA-A specificities A9, A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 70.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the B*0701 allele has been shown to be in error.

The sequence of the B*08:06 allele has been shown to be identical to B*08:20.

The sequence of the B*1305 allele has been shown to be identical to B*13:04.

The B*1324 allele has been renamed B*13:22:02

The B*150105 allele has been corrected and renamed B*15:120

The B*1522 allele has been renamed B*35:43.

The sequence of the B*1541 allele has been shown to be identical to B*15:39.

The B*1559 allele has been renamed B*35:44.

The B*15:100 allele has never been assigned.

The sequence of the B*1816 allele has been shown to be identical to B*18:14.

The sequence of the B*27051 allele has been shown to be identical to B*27:05:02.

The sequence of the B*2722 allele has been shown to be identical to the corrected B*27:06 sequence.

The B*3573 allele has been renamed B*35:08:03.

The B*35:43:02 allele has been renamed B*35:185.

The sequence of the B*39012 allele has been shown to be identical to B*39:01:01:01.

The sequence of the B*3921 allele has been shown to be identical to B*39:24.

The sequence of the B*4017 allele has been shown to be identical to B*40:16.

The sequence of the B*4041 allele has been shown to be identical to B*40:40.

The sequence of the B*4203 allele has never been assigned.

The sequence of the B*4401 allele has been shown to be identical to B*44:02:01:01.

The sequence of the B*5003 allele has been shown to be identical to B*50:02.

The sequence of the B*5125 allele has been shown to be identical to B*51:22.

The B*5147 allele has been renamed B*51:09:02.

The sequence of the B*5506 allele has been shown to be identical to B*55:04.

The sequence of the B*5803 allele has never been assigned.

The B*7901 allele has been renamed B*15:18:01.

The B*9530 allele has been renamed B*15:27:02.

2The B*08:26, 08:50 and 08:62 and B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:47 and 39:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*18:29 and the B*35:32, 35:37, 35:53N, 35:64, 35:68:01-35:68:02, 35:99, 35:118-35:119 and 35:174 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*41:09 and the B*45:02 and 45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*54:01:02 and the B*55:01:07-55:02:06, 55:07, 55:10, 55:12, 55:16, 55:19, 55:26, 55:30, 55:35, 55:37, 55:39, 55:41-55:43, 55:47-55:48 and 55:50 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:04, 55:08, 55:13, 55:27, 55:46 and 55:49 and the B*56:15, 56:19N and 56:22 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The B*55:23 and 55:32 and the B*56:18 and 56:31-56:32 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The serological reactivity of all HLA-B alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

‘ser’, serological HLA specificity.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Specificities and sizes of the PCR products of the 24 primer mixes used for DR low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DR serology3 | Amplified HLA-DRB alleles4 |
| --- | --- | --- | --- | --- |
| 736,8 | 205 bp, 255 bp | 515 bp | 1 | *01:01:01-01:02:05, 01:04-01:37 |
| 74 | 200 bp | 430 bp | 103 | *01:03 |
| 756 | 200 bp, 215 bp | 430 bp | 2, 15 | *15:01:01:01-15:58 |
| 76 | 210 bp | 430 bp | 16 | *16:01:01-16:05:02, 16:07-16:18 |
| 775,6,7 | 120 bp, 220 bp | 430 bp | 3, 17, 18, 11 | *03:01:01:01-03:65, 11:07, 11:53, 11:103, 11:105, 11:107, 15:25 |
| 785,6,7 | 80 bp, 210 bp | 430 bp | 3, 6, 17, 11, 13, 14 | *03:01:01:01-03:01:14, 03:04:01-03:06, 03:08-03:16, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:31, 03:33-03:34, 03:36-03:37, 03:43-03:48, 03:50-03:52, 03:54-03:65, 08:40, 11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-11:41, 11:48, 11:59, 11:63, 11:65:01-11:65:02, 11:68, 11:70, 11:73, 11:76, 11:79-11:80, 11:83, 11:85-11:87, 11:93, 13:01:01-13:04, 13:08, 13:10, 13:15-13:17, 13:19-13:20, 13:22-13:24, 13:27-13:29, 13:31-13:41, 13:43, 13:45, 13:48, 13:51-13:54, 13:57, 13:59, 13:61:01-13:61:02, 13:63-13:66:02, 13:68-13:76, 13:78-13:81, 13:83-13:85, 13:87-13:99, 13:101-13:102, 13:104-13:107, 13:109, 13:111-13:117, 14:16, 14:19, 14:21, 14:82, 14:95, 14:109 |
| 795,6 | 85 bp, 210 bp | 430 bp | 3, 6, 11, 13, 14, 1403, 18 | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 11:13:01w-11:13:02w, 11:26, 11:34, 13:15, 13:19, 13:26, 13:44, 13:53, 13:57, 13:85-13:86, 13:104, 14:02-14:03:02, 14:06:01-14:06:02, 14:09, 14:12:01-14:13, 14:17-14:21, 14:24, 14:27, 14:29-14:30, 14:32:01w-14:32:02w, 14:33, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:65w, 14:67, 14:77-14:78, 14:80-14:81, 14:83, 14:85, 14:89, 14:94, 14:98, 14:102, 14:106, 14:108-14:109, 14:115 |
| 805,6 | 100 bp, 175 bp | 430 bp | 3, 4 | *04:01:01-04:102 |
| 816 | 210 bp, 235 bp | 430 bp | 7, 13, 14 | *07:01:01:01-07:01:04, 07:03-07:21, 12:22, 13:17, 13:116, 14:50 |
| 826 | 170 bp, 215 bp, 250 bp | 515 bp | 8, 12, 14 | *08:01:01-08:19, 08:21-08:48, 11:67, 12:04, 12:16, 12:22, 14:11, 14:15, 14:68, 14:93 |
| 835,6 | 90 bp, 135 bp, 180 bp | 430 bp | 3, 9, 11 | *03:08, 03:65, 09:01:02-09:16, 11:07, 11:53, 11:103, 11:105, 11:107 |
| 84 | 205 bp | 430 bp | 10 | *10:01:01-10:03 |
| 855,6 | 100 bp, 170 bp | 430 bp | 3, 8, 11, 14 | *03:08, 03:65, 08:31, 08:41, 11:01:01-11:70, 11:72-11:113 |
| 865,6 | 85 bp, 105 bp | 430 bp | 12 | *08:32, 12:01:01-12:34 |
| 878 | 215 bp | 430 bp | 6, 8, 11, 13, 14, 1403 | *08:20-08:21, 11:01:01-11:04:08, 11:06:01-11:06:02, 11:08:01-11:12:02, 11:14:01-11:16, 11:18-11:21, 11:23-11:25, 11:27:01-11:33, 11:35-11:51, 11:54:01-11:54:02, 11:56-11:66, 11:68, 11:70, 11:72-11:81, 11:83-11:88, 11:90-11:97, 11:99-11:102, 11:106, 11:108-11:113, 13:01:01-13:08, 13:10-13:16, 13:18-13:43, 13:45-13:85, 13:87-13:115, 13:117-13:119, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:16, 14:19, 14:21-14:22, 14:25, 14:27, 14:40, 14:53, 14:63, 14:67, 14:69, 14:74, 14:77-14:78, 14:84-14:85, 14:98, 14:102, 14:105, 14:109, 14:115-14:116, DRB3*02:27 |
| 886,8 | 195 bp, 225 bp | 430 bp | 6, 8, 11, 12, 13, 14 | *08:01:01-08:02:04, 08:04:01-08:09, 08:11, 08:16-08:17, 08:20-08:22, 08:24, 08:26, 08:28, 08:31, 08:39, 08:41-08:44, 11:01:01-11:06:02, 11:09-11:12:02, 11:14:01-11:16, 11:20-11:21, 11:23-11:25, 11:27:01-11:30, 11:32-11:33, 11:35-11:41, 11:43-11:44, 11:46:01-11:51, 11:54:01-11:56, 11:58:01-11:63, 11:65:01-11:70, 11:72, 11:74:01-11:78, 11:80-11:88, 11:90-11:97, 11:99-11:102, 11:106, 11:108-11:113, 12:02:01-12:02:05, 12:13, 12:15-12:16, 12:18-12:21, 12:23, 12:26-12:27, 12:31N-12:33, 13:01:01-13:02:01, 13:02:03-13:02:05, 13:04-13:05:02, 13:07:01-13:09, 13:11:01-13:11:02, 13:14:01-13:24, 13:26-13:29, 13:31-13:32, 13:34-13:36, 13:38-13:43, 13:45-13:55, 13:57, 13:59, 13:61:01-13:65, 13:67-13:76, 13:78-13:80, 13:83-13:84, 13:87, 13:91-13:93, 13:96:01-13:100, 13:102-13:109, 13:111-13:114, 13:116-13:117, 14:15-14:16, 14:22, 14:24-14:25, 14:27, 14:37, 14:53, 14:73, 14:105 |
| 897 | 175 bp | 430 bp | 3, 6, 11, 13, 14, 1403, 17, 18 | *03:01:01:01-03:07, 03:09, 03:11:01-03:41, 03:43-03:45, 03:47-03:63, 08:20, 13:01:01-13:16, 13:18-13:42, 13:44, 13:46-13:66:02, 13:68-13:102, 13:104-13:115, 13:117-13:119, 14:02-14:03:02, 14:05:01-14:06:02, 14:09, 14:12:01-14:14, 14:17-14:21, 14:23:01, 14:23:03-14:24, 14:27, 14:29-14:30, 14:33, 14:36-14:37, 14:40-14:45, 14:47-14:48, 14:51, 14:56, 14:59, 14:63-14:65, 14:67, 14:77-14:78, 14:80-14:81, 14:83-14:85, 14:89, 14:91, 14:94-14:96, 14:98, 14:100, 14:102-14:103, 14:106, 14:108-14:109, 14:115-14:116 |
| 905,6 | 100 bp, 140 bp, 155 bp | 430 bp | 4, 6, 8,13, 14, 1404 | *04:62, 04:69, 04:73, 08:08, 11:69, 11:82, 13:45, 14:01:01-14:01:02, 14:04, 14:07:01-14:07:02, 14:10, 14:16, 14:22, 14:25-14:26, 14:28, 14:31-14:32:02, 14:35, 14:37-14:39, 14:49-14:50, 14:53-14:55, 14:57-14:58, 14:60-14:62, 14:68-14:71, 14:73-14:76, 14:79, 14:82, 14:86-14:88, 14:90, 14:93, 14:99, 14:101, 14:104-14:105, 14:107, 14:110-14:114, 14:117, DRB4*01:03:01:02N |
| 915,6, 9 | 110 bp, 135 bp, 170 bp | 430 bp | 3, 4, 6, 9, 11, 13,14,1404 | *03:10, 09:01:02-09:01:05, 09:01:07-09:02:02, 09:04-09:16, 11:13:01-11:13:02, 11:17, 11:52, 13:43, 14:01:01-14:02, 14:04-14:11, 14:13-14:14, 14:16-14:18, 14:19w, 14:20, 14:21w, 14:22-14:23:03, 14:26, 14:28-14:36, 14:38-14:39, 14:41, 14:43-14:52, 14:54-14:57, 14:59-14:62, 14:64-14:65, 14:68, 14:70-14:76, 14:79-14:83, 14:86-14:88, 14:90-14:97, 14:99-14:101, 14:103-14:108, 14:109w, 14:110-14:114, 14:117, 15:27, 15:34 |
| 925,6,8,10 | 110 bp, 175 bp, 225 bp | 430 bp | 2, 3, 4, 6, 8, 11,13, 14, 1403, 1404, 16 | *03:10, 08:09, 08:20-08:21, 08:32, 08:35, 11:13:01-11:13:02, 11:17, 11:23, 11:25, 11:31, 11:45, 11:52, 11:55, 11:64, 11:89, 11:96, 13:13, 13:18, 13:43, 13:45, 13:47, 13:55, 13:119, 14:01:01-14:01:03, 14:03:01-14:05:03, 14:07:01-14:08, 14:10-14:12:02, 14:14-14:16, 14:18, 14:22-14:23:03, 14:25-14:28, 14:31-14:32:02, 14:34-14:36, 14:38-14:40, 14:42-14:45, 14:49-14:50, 14:53-14:65, 14:67-14:79, 14:81-14:82, 14:84-14:93, 14:95-14:97, 14:99-14:105, 14:107, 14:110-14:117, 15:21w, 16:04w, 16:18w |
| 936,7 | 160 bp, 240 bp | 430 bp | 52 | DRB3*01:01:02:01-01:15, DRB3*02:01-02:28, DRB3*03:01:01-03:03 |
| 9410 | 215 bp | 430 bp | 53 | DRB4*01:01:01:01-01:08 |
| 95 | 175 bp | 430 bp | 51 | DRB5*01:01:01-01:14, DRB5*02:02-02:05 |
| 9611 | - | - |  | Negative Control |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 75 and 90 to 92.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 73 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DR low resolution typing.

In addition, well number 82 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological split of all DRB1 alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP set are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

The DRB1*08:09 and the DRB1*14:15 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:20 and the DRB1*13:18, 13:47 and 13:55 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31, 08:41 and DRB1*11:67 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*13:13 and 13:119 and DRB1*14:84 and 14:116 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 73, 75, 77 to 83, 85, 86, 88 and 90 to 93.

7Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 in addition to primer mix 93.

8Primer mix 73, 87, 88 and 92 may give rise a primer oligomer formation.

9Primer mix 91 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

10The DRB4*01:03:01:02N allele is amplified by the primer pair in well No. 94, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer pair.

11Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘w’, might be weakly amplified.

1The sequence of the DRB1*0702 allele has been shown to be identical to DRB1*07:01:01:01.

The sequence of the DRB1*080301 allele has been shown to be identical to DRB1*08:03:02.

The sequence of the DRB1*09011 allele has been shown to be identical to DRB1*09:01:02.

The sequence of the DRB1*1171 allele has been shown to be identical to DRB1*11:02:01.

The sequence of the DRB1*12031 allele has been shown to be identical to DRB1*12:01:01.

The DRB1*1466 allele has been renamed DRB1*14:32:02.

The sequence of the DRB1*1606 allele has been shown to be identical to DRB1*16:05:01.

The sequence of the DRB3*010101 allele has been shown to be identical to DRB3*01:01:02:01.

The DRB4*0101102N allele has been renamed DRB4*01:03:10:02N.

The sequence of the DRB5*0201 allele has been shown to be identical to DRB5*02:02.

Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 in addition to primer mix 93.

2The DRB1*08:09 and the DRB1*14:15 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:20 and the DRB1*13:18, 13:47 and 13:55 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31, 08:41 and DRB1*11:67 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*13:13 and 13:119 and DRB1*14:84 and 14:116 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The serological reactivity of all DRB alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 73, 75, 77 to 83, 85, 86, 88 and 90 to 93.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘ser’, serological HLA specificity.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information for the primer matching region(s) is not known.

Change in revision R01 compared to R00:

- The HLA-A*68:13 allele is amplified by primer mix 4. This has been corrected in the specificity and interpretation tables.

Change in revision R02 compared to R01:

- Primer mix 43 does not amplify the B*15:101 allele. This has been corrected in the Specificity and Interpretation tables.

Change in revision R03 compared to R02:

- The HLA-A*03:01:03, 03:09, 03:23:01, 03:89, 03:108, 11:06, 11:18, 24:28, 24:89, 26:03:01-26:03:02, 26:06, 26:07:02, 26:21, 26:30, 30:13, 30:16, 30:44, 30:46, 68:05, 68:15, 68:20 and 68:30 alleles are weakly amplified by primer mix 2. This has been corrected in the Specificity and Interpretation Tables.

- The DR serology has been corrected in the Specificity Table.

Changes in revision R04 compared to R03:

- In primer mix 5, the specific PCR product of 535 base pairs may be difficult to distinguish from the internal control band. A foot note has been added to the Specificity Table.

- Primer mix 17 does not amplify the HLA-A*36:02 allele. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R05 compared to R04:

- Primer mix 43 does not amplify the B*44:10 allele. This has been corrected in the Specificity and Interpretation Tables.