Gel Documentation Form and Worksheet

HLA-A-B-DR SSP Combi Tray	Lot No: 73X	Expiry Date: 2017-May-01

(101.701-24/06 – 24u/06u)

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):__________

Test  Date:_______________

Tested By:

Review Date:_________________

Reviewed By:_________________

Interpretation:___________	Failed lanes: ___________                                Comments:_________________

Gel Picture

PHOTO DOCUMENT

HLA-A low  resolution

HLA-B low  resolution

DR low  resolution

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-A low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

The primer pairs in wells 1, 2, 11, 14 and 15 will in many samples give rise to two or three HLA-specific PCR fragments.

Primer mix 3 may faintly amplify the A*30:04:01-30:04:02, 30:06, 30:17, 30:29, 30:46 and 30:77 alleles.

Primer mix 6 may give rise to a lower yield of A*23 alleles than the other A low primer mixes.

Primer mix 9 may weakly amplify the A*34 alleles.

Primer mix 22 might faintly amplify most A*11 alleles.

Primer mixes 6, 10, 14, 20 and 23 may have a tendency of giving rise to primer oligomer formation.

Primer mixes 6, 15 and 24 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.

Primer mixes 15, 20 may have tendencies of unspecific amplifications.

Primer mix 21 may generate a false positive band of about 560 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings.

HLA-B low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 28 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

Primer mixes 28, 29, 33, 42, 53 and 63 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes.

Primer mix 33 has pronounced tendencies of unspecific amplifications.

Primer mixes 25, 43, 44, 50, 53 and 69 have a tendency to giving rise to primer oligomer formation, most pronounced in primer mix 53

The B*57 and B*58 alleles might be faintly amplified by primer mix 33.

The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 49.

Primer mixes 56 and 67 may generate a false positive band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

Primer mix 63 may give rise to a lower yield of B*54 alleles than the other B low primer mixes.

The Bw4-associated HLA-A specificities A9, A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 70.

DR low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 73, 75, 77 to 83, 85, 86, 88 and 90 to 93.

Primer mix 73, 80, 87, 88 and 91 to 94 may have a tendency to giving rise to primer oligomer formation.

Primer mix 91 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 in addition to primer mix 93.

The DRB4*01:03:01:02N allele is amplified by primer mix 94, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer mix.

Primer mixes 75, 76, 80, 81 and 92 may have tendencies of unspecific amplifications.

Due to sharing of sequence motifs in codon 38 and 47, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 and DRB3*01:23 in primer mix 77, in addition to primer mix 93.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

HLA-A low resolution Interpretation Table

1HLA-A, HLA-B and HLA-DRB alleles listed on the IMGT/HLA web page 2014-January-17, release 3.15.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-A alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4The A*01:136 and the A*11:94 and A*11:112 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mixes 1 and 11.

The A*24:14-24:15, 24:51-24:53, 24:57, 24:64, 24:94, 24:114, 24:138, 24:188, 24:222N and 24:228 and the C*04:01:03 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mixes 5.

5The HLA-A alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns.

These alleles can be separated by the respective high resolution SSP primer sets

| Alleles | Alleles |
| --- | --- |
| A*03:88, A*11:130 | A*23:14:01-23:14:02, A*24:24, 24:71 |
| A*03:152, A*24:92 | A*31:08, A*33:53 |

‘w’, might be weakly amplified.

HLA-B low resolution Interpretation Table

1HLA-A, HLA-B and HLA-DRB alleles listed on the IMGT/HLA web page 2014-January-17, release 3.15.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-B alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

page http://hla.alleles.org/alleles/deleted.html.

4The HLA-B alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns.

These alleles can be separated by the respective high resolution SSP primer sets.

| Alleles | Alleles |
| --- | --- |
| B*07:202, B*81:02 | B*41:09, 41:29, B*45:02-45:03 |
| B*08:26, 08:50, 08:62, 08:85, 08:94, B*42:07 | B*53:30,B* 57:45, 57:51, 57:69 |
| B*13:21, 13:35, 13:59, 13:71, B*44:135, 44:158 | B*55:04, 55:08, 55:13, 55:27, 55:46, 55:49, B*56:01:05, 56:15, 56:19N, 56:22 |
| B*14:08:01, B*39:25N, 39:30, 39:32-39:34, 39:47, 39:50, 39:74, 39:82 | B*55:23, 55:32, B*56:18, 56:31-56:32 |
| B*18:29, 18:72, 18:92, B*35:32:01-35:32:02, 35:37, 35:53N, 35:64:01-35:64:02, 35:68:01-35:68:02, 35:99, 35:118-35:119, 35:174 | B*57:67, 58:36 |

‘?’, nucleotide sequence information not available for the primer matching sequence.

‘w’, might be weakly amplified.

DR low resolution Interpretation Table

1HLA-A, HLA-B, HLA-DRB alleles listed on the IMGT/HLA web page 2014-January-17, release 3.15.0, www.ebi.ac.uk/imgt/hla.

The DRB4*02:01N and DRB4*03:01N null alleles will not be amplified by the DR low resolution primer set.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all DRB alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4The HLA-DRB1, -DRB3, -DRB4 and –DRB5  alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns:

Alleles

DRB1*08:20, DRB1*13:18, 13:47, 13:55, 13:158, 13:164

DRB1*08:31, 08:41, DRB1*11:67

DRB1*13:13, 13:119, 13:154, 13:156, DRB1*14:84, 14:116, 14:144

‘w’, might be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 77 does not amplify the DRB1*03:11:01, *11:53 and *15:100 alleles. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- Due to sharing of sequence motifs in codon 38 and 47, DRB3*01:23 will also be amplified in primer mix 77, in addition to primer mix 93. A footnote has been changed in the Specificity Table.

Changes in revision R03 compared to R02:

- Primer mix 34 amplifies the C*02:06 and C*15:42 alleles and primer mix 65 amplifies the C*15:15 allele. This has been corrected in the Specificity and Interpretation Tables.

- Primer mix 21 may generate a false positive band of about 560 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings. A footnote has been added in the Specificity Table.