PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A-B-DR-DQ Combi Tray

Prod. No:	101.708-24/24u

Lot No:    	70S

Expiry Date:    	2016-March-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 21 primer mixes used for HLA-A low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-A serology3 | Amplified HLA-A alleles4,5,6 |
| --- | --- | --- | --- | --- |
| 16,7 | 120 bp, 145 bp, 225 bp | 800 bp | A1, A36 | *01:01:01:01-01:04N, 01:06-01:33, 01:35-01:130, 03:18, 03:135, 11:94, 11:112, 36:01-36:05 |
| 27 | 175 bp, 215 bp, 255 bp, 365 bp, 415 bp | 800 bp | A2, A203, A210, A19, A28 | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:01:81, 02:01:83-02:22:02, 02:24:01-02:35:01, 02:35:03-02:47, 02:49-02:77, 02:78w, 02:79:01-02:97:02, 02:99, 02:101:01-02:128, 02:130-02:414, 02:416-02:417 |
| 311 | 205 bp, 235 bp | 1070 bp | A1, A3, A11, A32, A34, A36 | *01:12, 01:19, 01:21, 01:126, 02:338, 03:01:01:01-03:17:02, 03:19-03:74, 03:76-03:94, 03:96-03:134, 03:136-03:166, 11:25, 11:60, 11:130, 24:92, 32:04, 34:02:01-34:04, 34:07-34:10N, 36:02 |
| 4 | 190 bp | 800 bp | A1, A2, A3, A11, A26, A30, A36, A68 | *01:01:01:01-01:01:22, 01:01:24-01:01:47, 01:01:49-01:04N, 01:06-01:07, 01:09-01:11N, 01:13, 01:16N-01:18N, 01:20-01:29, 01:31N-01:33, 01:35-01:78, 01:80-01:98, 01:100-01:130, 02:78, 02:169, 03:12, 03:18, 03:88, 03:135, 11:01:01-11:27, 11:29-11:52Q, 11:54-11:139, 11:141-11:142, 26:19, 26:72, 30:08, 36:04, 68:13:01, 68:66, C*03:174, C*08:69, C*16:56 |
| 58,12 | 135 bp, 200 bp | 800 bp | A9, A23, A24, A29, A80 | *23:01:01-23:56, 23:58, 24:24, 24:71, 29:07, 29:49, 31:29, 80:01:01:01-80:03, B*18:27 |
| 6 | 175 bp, 205 bp | 1070 bp | A2, A23, A24, A26, A33 | *02:17:01w-02:17:02w, 11:139, 23:14:01-23:14:02, 24:02:01:01-24:11N, 24:13:01-24:13:02, 24:17-24:50, 24:54-24:56, 24:58-24:63, 24:66-24:91, 24:93, 24:95-24:113, 24:115-24:137, 24:139-24:187, 24:189-24:210, 24:212-24:221, 24:223-24:227, 24:229-24:232N, 26:16, 33:19, 68:45 |
| 7 | 165 bp, 200 bp | 800 bp | A2/A28, A3, A10, A11, A25, A26, A32, A34, A66, A68, A69 | *01:51, 02:55, 03:24, 03:50, 11:10, 25:01:01-25:16, 25:18-25:22, 26:01:01-26:06, 26:08-26:15, 26:17-26:18, 26:20-26:43:02, 26:45-26:63, 26:65-26:71N, 26:73-26:85, 29:28, 32:15, 33:51, 34:01:01-34:11, 66:01-66:19, 68:01:01:01-68:100, 69:01 |
| 86,9 | 75 bp | 800 bp | A3, A25, A32 | *25:01:01-25:22, 32:01:01-32:01:13, 32:01:15-32:02, 32:04, 32:06-32:37, 32:39-32:57, B*07:81, B*08:52, B*18:67, B*38:41, B*53:05, B*53:16 |
| 96 | 80 bp, 240 bp | 800 bp | A10, A26, A43 | *01:43, 01:51, 11:17, 11:40, 26:01:01-26:02, 26:04, 26:07:01-26:20, 26:22-26:29, 26:31-26:43:02, 26:45-26:77, 26:79-26:85, 33:13, 33:48, 43:01, 68:84 |
| 106,7,8 | 80 bp, 175 bp, 500 bp | 1070 bp | A1, A11, A24, A26, A31, A34, A66 | *01:13, 01:28, 03:63, 03:88, 11:01:01-11:27, 11:29-11:52Q, 11:54-11:139, 11:141-11:142, 24:19, 24:44, 26:03:01-26:03:02, 26:06, 26:21, 26:78, 34:01:01-34:08, 34:10N-34:11, 66:01, 66:04-66:11, 66:13-66:14, 66:17-66:19, 80:02 |
| 116 | 125 bp, 185 bp | 800 bp | A3, A10, A25, A26, A31, A34, A43, A66 | *02:309, 03:01:19, 11:11, 25:05-25:06, 26:09, 26:54, 31:03-31:04, 34:01:01-34:11, 43:01, 66:02-66:03, 66:16 |
| 12 | 175 bp, 225 bp | 1070 bp | A1, A2, A3, A10, A25, A26, A34, A43, A66 | *01:01:56w, 01:13, 02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:103, 02:135, 03:01:01:01-03:01:22, 03:01:24-03:07, 03:09-03:11N, 03:13-03:31, 03:33-03:35, 03:37-03:40, 03:42-03:56, 03:58, 03:60-03:71, 03:73-03:87, 03:90-03:106, 03:109-03:110, 03:112-03:141, 03:143-03:151, 03:153-03:166, 25:01:01-25:05, 25:07-25:22, 26:01:01-26:01:20, 26:01:22-26:01:27, 26:02w, 26:03:01-26:03:02, 26:05-26:08, 26:10-26:28, 26:29w, 26:30-26:33, 26:35-26:43:02, 26:45-26:48, 26:49w, 26:50-26:72, 26:74-26:77, 26:79-26:85, 30:55, 34:08, 43:01, 66:01, 66:04-66:09, 66:10w, 66:11-66:15, 66:17-66:19, 68:71, 74:13 |
| 136,7,10,12 | 80 bp, 115 bp, 200 bp, 240 bp, 460 bp | 1070 bp | A26, A29, A31, A34, | *02:237, 02:309, 03:95, 26:19, 26:22, 29:01:01:01-29:53, 31:03-31:04, 32:42, 34:04, 66:09 |
| 146,12 | 90 bp, 135 bp, 205 bp | 1070 bp | A1, A30 | *01:13, 01:28, 03:43, 03:82, 11:113, 30:01:01-30:04:02, 30:06-30:20, 30:22-30:74, 31:35 |
| 15 | 240 bp, 380 bp | 1070 bp | A24, A31, A32 | *02:237, 03:95, 29:14, 31:01:02-31:74, 32:05, 33:53 |
| 16 | 140 bp, 180 bp | 1070 bp | A29, A32 | *01:95, 03:43, 03:82, 29:13, 29:39, 31:35, 32:01:01-32:03, 32:05-32:57, 74:07 |
| 17 | 200 bp, 390 bp | 1070 bp | A32, A33, A68, A74 | *02:243, 29:48, 32:15, 33:01:01-33:01:07, 33:03:01-33:37, 33:39-33:52, 33:54-33:68, 68:29, 74:04 |
| 18 | 340 bp, 375 bp | 800 bp | A2, A19, A68, A74 | *01:121, 02:65, 02:407, 68:25, 74:01-74:06, 74:08-74:17 |
| 1912 | 210 bp, 240 bp | 800 bp | A2, A210, A25, A30, A68 | *02:34-02:35:03, 02:46, 02:48, 02:56:01-02:56:02, 02:62, 02:70, 02:78, 02:103, 02:129, 25:05, 26:54, 68:01:01:01-68:100 |
| 206 | 65 bp, 240 bp, 375 bp | 800 bp | A2, A26, A28, A32, A66, A68, A69 | *02:55, 02:149, 02:243, 02:309, 23:57, 24:82, 26:22, 31:41, 32:06, 33:22, 66:06, 66:09, 68:29, 69:01 |
| 216,12 | 75 bp, 160 bp, 240 bp, 495 bp | 800 bp | A2, A26, A68, A80 | *02:55, 03:41, 03:63, 03:75, 03:88, 11:130, 24:18, 24:204, 24:213, 26:03:01-26:03:02, 26:05-26:06, 26:21, 26:30, 26:78, 33:24, 36:02, 68:05, 68:15, 68:20, 80:01:01:01-80:01:01:02, 80:03 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A low resolution SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A low resolution typing. In addition, wells number 2, 4, 5, 7, 8, 9, 11 and 18 to 21 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-A alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

5Due to the sharing of sequence motifs between HLA-Class I alleles, primer mix 4 will amplify the C*03:174, C*08:69 and C*16:56 alleles, primer mix 5 will amplify the B*18:27 allele and primer mix 8 will amplify the B*07:81, B*08:52, B*18:67, B*38:41, B*53:05 and B*53:16 alleles.

6HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

7The primer pairs in vials 1, 2, 9 and 13 will in many samples give rise to two or three HLA-specific PCR fragments.

8Primer mix 5 may give rise to a lower yield of A*23 alleles than the other A low primer mixes.

9Primer mix 8 may weakly amplify the A*34 alleles.

10Primer mix 13 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.

11Primer mix 3 may faintly amplify the A*30:04:01-30:04:02, 30:06, 30:17 and 30:29 alleles.

12Primer mixes 5, 13, 14, 19 and 21 may have a tendency of giving rise to primer oligomer formation.

‘w’, might be weakly amplified.

1The sequence of the A*0105N has been shown to be identical to A*01:04N.

The A*01:34N allele has been renamed A*01:01:38L

The A*020116 allele has been renamed to A*02:134.

The A*020120 has been shown to be identical to A*02:01:18.

The sequence of the A*02:01:82 allele has been shown to be identical to A*02:01:84.

The sequence of the A*0223 allele has been shown to be identical to A*02:22:01.

The sequence of the A*0298 allele has been shown to be identical to A*02:96.

The A*02:100 allele has never been assigned.

The A*1128 allele has been renamed to A*11:15:02.

The A*11:53 allele has been shown to be identical to the corrected A*11:02:01.

The sequence of the A*2401 allele has been shown to be in error.

The sequence of the A*2412 allele has been shown to be identical to A*24:08.

The A*2416 allele has been renamed to A*31:08.

The A*2465 allele has been renamed to A*24:13:02.

The A*24:211 allele has been renamed to A*24:135:02.

The A*26:44 allele has been renamed to A*26:43:02.

The sequence of the A*3005 allele has been shown to be identical to A*30:04.

The A*3021 allele has been renamed to A*30:11:02.

The sequence of the A*31011 allele has been shown to be identical to A*31:01:02.

The sequence of the A*3302 allele has been shown to be identical to A*33:03:01.

The sequence of the A*33:38 allele has been shown to be identical to A*33:44

2The serological reactivity of all HLA-A alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

3The A*36 alleles will give rise to identical amplification patterns as a number of A*01 alleles. These alleles can be separated by the A*01 and A*36 high resolution SSP primer sets.

4The A*03:01:23, 03:08, 03:32, 03:36N, 03:57, 03:59, 03:72, 03:89, 03:107-03:108, 03:111, 03:142 and 03:152 and the A*24:92 alleles will give rise to identical amplification patterns. These alleles can be separated by the respective high resolution SSP primer sets.

5The A*03:88 and the A*11:130 alleles will give rise to identical amplification patterns. These alleles can be separated by the respective high resolution SSP primer sets.

6The A*23:01:01-23:13, 23:15-23:56 and A* 23:58 and the B*18:27 alleles will give rise to identical amplification patterns. These alleles can be separated by the respective high resolution SSP primer sets.

7The A*23:14:01-23:14:02 and the A*24:24 and 24:71 alleles will give rise to identical amplification patterns. These alleles can be separated by the respective high resolution SSP primer sets.

8The A*31:01:02-31:02, 31:05-31:28, 31:30-31:34, 31:36-31:40 and 31:42-31:74, and the A*33:53 alleles will give rise to identical amplification patterns. These alleles can be separated by the respective high resolution SSP primer sets

9The primer pairs in vials 1, 2, 10 and 14 will in many samples give rise to two or three HLA-specific PCR fragments.

‘w’, may be weakly amplified.

10The A*32:01:14, 32:03, 32:38 and the A*74:07 alleles will give rise to identical amplification patterns. These alleles can be separated by the respective high resolution SSP primer sets.

Specificities and sizes of the PCR products of the 43 primer mixes used for HLA-B low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | HLA-B serol.3 | Amplified HLA-B alleles4,5 |
| --- | --- | --- | --- | --- |
| 2210 | 285 bp | 800 bp | 7, 15, 37, 40, 41, 42, 48 | *07:02:01-07:06, 07:08-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:52, 07:54-07:59, 07:61-07:76, 07:79-07:83, 07:85-07:99, 07:101-07:121, 07:123-07:138, 07:140-07:164, 07:166-07:190, 08:01:01-08:05, 08:07-08:08N, 08:10-08:11, 08:13-08:15, 08:17-08:83, 08:85-08:88, 08:90-08:102, 08:104-08:106, 08:108, 35:66, 35:87, 37:07, 40:15-40:16, 40:30-40:32, 40:34, 40:45, 40:59, 40:80, 40:98, 40:137, 40:160:01-40:160:02, 41:02:01-41:02:05, 41:04, 41:10-41:11, 41:13, 41:18-41:19, 41:23-41:24, 41:27, 42:01:01-42:02, 42:05:01-42:07, 42:09-42:13, 42:15-42:20, 44:150, 44:166, 48:01:01-48:01:04, 48:05-48:12, 48:14-48:20, 48:22, 48:27-48:30, 53:15, 81:01-81:06 |
| 23 | 215 bp | 1070 bp | 8, 15, 44 | *08:01:01-08:05, 08:07-08:25, 08:27-08:49, 08:51-08:61, 08:63-08:64, 08:66-08:84, 08:86N-08:93, 08:95-08:108, 15:142, 15:180, 35:218, 44:49, 44:156, 49:25, 51:68 |
| 24 | 140 bp, 235 bp | 1070 bp | 7, 8, 13, 15, 35, 40, 4005, 44, 46, 53, 57, 58 | *07:20, 07:24, 07:60, 07:100, 07:131, 08:21, 08:25, 13:01:01-13:04, 13:06-13:08, 13:10-13:23, 13:25-13:38, 13:40-13:57, 13:59-13:66, 15:07:01-15:07:02, 15:36w, 15:45, 15:55, 15:68, 15:89w, 15:126, 15:207, 35:05:01-35:05:03, 35:16-35:17, 35:22, 35:30, 35:51, 35:58, 35:72, 35:89, 35:97, 35:113-35:114, 35:125, 35:164, 35:199, 40:05, 40:71, 40:174, 44:08w, 44:54, 44:57w, 44:60, 44:106, 44:110, 44:135, 44:158, 46:12, 51:64, 51:148, 53:14, 55:51, 58:18, C*03:05, C*03:25, C*03:27, C*03:143, C*03:167 |
| 257,8 | 130 bp, 265 bp | 800 bp | 13, 14, 15, 35, 40, 41, 44, 45, 47, 49, 50 | *13:01:01-13:04, 13:06-13:13, 13:15-13:23, 13:25-13:57, 13:59-13:67, 14:01:01-14:02:08, 14:03-14:04, 14:07N, 14:09, 14:11-14:12, 14:14-14:33, 15:46, 15:53, 15:106, 15:212, 18:44:01-18:44:02, 35:46-35:47, 35:63, 35:154, 35:207, 35:217, 40:01:01-40:01:29, 40:07, 40:10:01-40:10:02, 40:14:01-40:16, 40:22N-40:23, 40:25, 40:30-40:34, 40:36, 40:38, 40:42-40:43, 40:45, 40:47-40:49, 40:51-40:55, 40:58-40:63, 40:65-40:67, 40:69, 40:72:01-40:73, 40:76-40:77, 40:79-40:81, 40:84, 40:87:01-40:88, 40:92, 40:100-40:102, 40:106, 40:108, 40:110, 40:112-40:114, 40:116-40:118N, 40:121, 40:123-40:126, 40:128-40:130, 40:132, 40:134-40:136, 40:139-40:141, 40:146-40:147, 40:150-40:156, 40:158, 40:160:01-40:160:02, 40:163-40:164, 40:166, 40:168, 40:171-40:172, 40:175, 40:178-40:179, 40:182-40:183, 40:186-40:188, 40:191-40:199, 40:204, 40:207-40:208, 40:212-40:213, 40:215-40:219, 40:221-40:223, 40:227-40:228, 40:231, 40:233-40:238, 41:01-41:27, 44:02:01:01-44:03:06, 44:03:08-44:03:15, 44:03:17-44:03:18, 44:04-44:05:04, 44:09-44:39, 44:41:01-44:43:02, 44:45-44:56N, 44:58N-44:80, 44:82, 44:84-44:102, 44:104-44:110, 44:112-44:129, 44:132-44:133, 44:135-44:144, 44:146-44:155, 44:157-44:175, 45:01-45:14, 47:01:01:01-47:08, 49:01:01-49:02, 49:04-49:24, 50:01:01-50:02, 50:04-50:06, 50:08-50:20, 50:31-50:32, 53:30, 57:45, 57:51, C*01:30, C*12:87 |
| 268 | 185 bp, 235 bp | 800 bp | 14, 16, 27, 37, 38, 39, 47, 64, 67 | *07:02:32, 14:01:01-14:01:03, 14:07N-14:08, 14:10, 14:12, 14:14, 14:19, 14:26, 14:32, 27:01-27:05:15, 27:05:17-27:05:19, 27:05:21-27:06, 27:08-27:10, 27:12-27:13, 27:16-27:18, 27:20, 27:23, 27:26-27:27, 27:29, 27:31, 27:35-27:37, 27:39-27:42, 27:44-27:46, 27:48-27:61, 27:64N-27:69, 27:72-27:75, 27:77-27:80, 27:82-27:105, 37:02, 38:01:01-38:02:02, 38:02:04-38:03, 38:07-38:24, 38:26-38:32, 38:34N-38:43, 39:01:01:01-39:01:01:03, 39:01:03-39:01:08, 39:01:10-39:02:01, 39:03, 39:05:01-39:09, 39:11, 39:14-39:15, 39:18, 39:19:02, 39:22, 39:24:01-39:39:01, 39:40:01N-39:44, 39:46-39:48, 39:50-39:62, 39:64-39:71, 39:73-39:75, 39:77, 39:79-39:82, 47:04-47:05, 48:21, 48:26, 67:01:01, 67:03 |
| 279 | 180 bp | 800 bp | 14, 35, 38, 39, 65 | *07:28, 08:87, 14:02:01-14:02:02, 14:02:04-14:02:05, 14:02:07-14:06:02, 14:09, 14:11, 14:13, 14:15-14:18, 14:20, 14:22-14:23, 14:25, 14:27, 14:29-14:31, 14:33, 15:77, 15:189, 15:233, 35:26, 38:05, 38:33, 39:04, 44:16, 44:37:01-44:37:02, 44:64:01-44:64:02, 44:91, 44:132, 44:150, 52:26, 57:04, A*23:31, A*24:106, C*07:231, C*16:10 |
| 28 | 290 bp | 1070 bp | 15, 62, 63, 75, 76, 77, 46 | *15:01:01:01-15:01:03, 15:01:06-15:02:06, 15:03:03, 15:04-15:08, 15:10:02-15:11:03, 15:11:05-15:15, 15:17:01:01-15:17:02, 15:19-15:21, 15:24:01-15:28, 15:30-15:36, 15:38:01-15:40, 15:42-15:46, 15:48, 15:50, 15:55-15:58, 15:60, 15:63, 15:65-15:66, 15:70-15:71, 15:73, 15:75-15:79N, 15:81-15:89, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:107, 15:109-15:113, 15:116-15:118, 15:120-15:122, 15:125-15:126, 15:128-15:129, 15:135-15:150, 15:152, 15:154-15:155, 15:157, 15:159-15:160, 15:162-15:172, 15:174-15:175, 15:177-15:179, 15:181N-15:185, 15:187-15:196, 15:199, 15:201-15:209N, 15:211, 15:213-15:219, 15:223-15:225, 15:227-15:228, 15:230-15:234, 15:236-15:237, 15:239-15:241, 15:244-15:251, 15:256-15:262N, 15:264-15:265, 15:267-15:273, 46:01:01-46:38, C*01:73, C*07:02:30, C*08:16:02 |
| 29 | 165 bp, 220 bp, 285 bp, 330 bp | 1070 bp | 8, 15, 35, 37, 40, 41, 42, 44, 45, 51, 56, 57, 70, 71, 72 | *08:01:01-08:05, 08:08N-08:13, 08:15-08:27, 08:29-08:36, 08:38-08:48, 08:50-08:73, 08:75-08:76, 08:78, 08:80-08:86N, 08:88, 08:90-08:106, 13:46, 15:03:01-15:03:04, 15:09-15:10:02, 15:18:01-15:18:05, 15:23, 15:29, 15:37, 15:46-15:47:02, 15:49, 15:51-15:54, 15:61-15:62, 15:64:01-15:64:02, 15:69, 15:72, 15:74, 15:80, 15:83, 15:90-15:91, 15:93, 15:98-15:99, 15:103, 15:106, 15:108, 15:114-15:115, 15:119, 15:123-15:124, 15:127, 15:131-15:134, 15:143, 15:151, 15:153, 15:156, 15:158, 15:161, 15:173, 15:176, 15:186, 15:197-15:198, 15:200, 15:210, 15:212, 15:220-15:221, 15:226N, 15:229, 15:235, 15:238, 15:242-15:243, 15:252-15:253, 15:255, 15:259, 15:263, 15:266, 15:274-15:275, 35:87, 37:12, 38:30, 40:12, 40:136, 40:231, 41:01-41:03:02, 41:05-41:09, 41:11-41:17, 41:19-41:27, 42:01:01-42:02, 42:04-42:08, 42:10-42:15, 42:17-42:19, 44:14-44:15, 44:18, 44:20, 44:62, 44:100, 44:166, 45:01-45:14, 51:08, 51:20, 51:36, 51:44N, 51:97, 51:141, 52:19, 55:20, 55:56, 56:13, 57:09, 82:01-82:03, C*07:46 |
| 308,9,11 | 165 bp, 190 bp, 390 bp | 1070 bp | 15, 35, 40, 44, 51, 52, 56, 58, 62, 63, 70, 78 | *13:62, 15:01:02, 15:09, 15:16:01-15:17:02, 15:67, 15:95, 15:162, 15:168, 15:177, 15:196, 15:208, 15:216, 15:222, 15:230, 15:243, 15:254, 15:268, 15:273, 35:01:10, 35:04:02, 40:26, 40:28, 44:62, 49:18, 50:14, 51:01:01-51:01:03, 51:01:05-51:02:03, 51:02:05-51:09:02, 51:11N-51:12, 51:13:02-51:14, 51:16-51:23, 51:24:03-51:24:04, 51:26-51:33, 51:37-51:41N, 51:43-51:44N, 51:46, 51:48-51:61:02, 51:63-51:80, 51:82-51:91, 51:94-51:98N, 51:100-51:130, 51:132-51:142, 51:145-51:147, 51:149N-51:152, 52:01:01:01-52:01:04, 52:01:06-52:01:14, 52:01:16-52:13, 52:15-52:24, 52:26-52:31, 53:28, 56:05:01-56:06, 56:21, 58:08:01-58:08:02, 78:01:01-78:07 |
| 31 | 180 bp | 800 bp | 13, 15, 40, 55, 62, 70, 71, 72, 75 | *07:78, 13:01:01-13:02:15, 13:07N-13:09, 13:11, 13:14-13:20, 13:22:01-13:23, 13:25, 13:27-13:34, 13:36-13:45, 13:47, 13:49N-13:50, 13:52, 13:54-13:58, 13:60-13:61, 13:63N-13:67, 40:48, 45:10, 49:07, 50:31, 54:26, 55:09, 55:22, 55:24 |
| 326 | 105 bp, 195 bp | 1070 bp | 7, 8, 15, 27, 38, 39, 40, 44, 45, 48, 55, 72 | *07:27, 07:50, 08:04, 08:17, 08:54, 15:03:01-15:03:04, 15:47:01-15:47:02, 15:49, 15:54, 15:61-15:62, 15:64:01-15:64:02, 15:68-15:69, 15:91, 15:98, 15:103, 15:123, 15:127, 15:131-15:132, 15:151, 15:156, 15:158, 15:173, 15:210, 15:220, 15:235, 15:242, 15:253, 15:259, 15:266, 15:274, 18:37, 27:18, 27:29, 37:28, 38:03, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:23, 39:39:01-39:39:02, 39:49, 40:12, 40:46, 40:93, 42:11, 44:10, 44:15, 44:18, 44:40, 44:44, 44:130, 44:140, 44:156, 45:01, 45:05-45:07, 45:11-45:14, 48:01:01-48:05, 48:07-48:30, 49:20, 50:02, 52:16, 55:18, 82:01-82:03 |
| 33 | 195 bp | 1070 bp | 15, 35, 37, 38, 44, 45, 58, 76 | *15:12, 15:14, 15:19, 15:91, 15:131, 15:161, 15:270, 18:54, 18:56, 35:45, 35:71, 37:01:01-37:09, 37:12-37:13, 37:15-37:21, 37:23-37:28, 37:30N-37:33N, 37:35-37:36, 38:17, 44:17, 44:43:01-44:43:02, 44:144, 45:09, 46:17, 53:22, 58:07, A*24:174 |
| 346,10 | 105 bp, 395 bp, 435 bp | 1070 bp | 7, 15, 41, 42, 59, 62, 63, 77 | *07:04, 07:25, 07:146, 15:09-15:10:02, 15:13:01-15:13:02, 15:16:01-15:18:05, 15:21, 15:23-15:24:02, 15:37, 15:44, 15:51-15:52, 15:66-15:67, 15:72, 15:80, 15:87, 15:90, 15:93, 15:95, 15:99, 15:108, 15:114-15:115, 15:119, 15:124, 15:133-15:134, 15:153, 15:157, 15:161-15:162, 15:168, 15:176-15:177, 15:186, 15:189, 15:196-15:198, 15:200, 15:208, 15:216, 15:221-15:222, 15:226N, 15:229-15:230, 15:238, 15:243, 15:252, 15:254, 15:263, 15:268, 15:273, 15:275, 40:136, 40:231, 41:08, 42:05:01-42:05:02 |
| 356,10 | 115 bp, 150 bp | 1070 bp | 18 | *18:01:01:01-18:15, 18:17N-18:28, 18:30-18:71, 18:73-18:89, 48:28 |
| 366 | 80 bp | 1070 bp | 27, 40, 44 | *07:73, 08:88, 08:102, 27:01-27:21, 27:23-27:51, 27:53-27:66N, 27:68-27:74, 27:76-27:105, 38:22, 40:46, 40:93, 44:40, 44:44, 44:130, 49:26 |
| 37 | 150 bp | 800 bp | 13, 15, 18, 27, 35, 37, 39, 40, 44, 48, 51, 53, 55, 56, 57, 58, 62, 75, 77 | *08:49, 13:01:01-13:01:06, 13:06-13:07N, 13:12:01-13:13, 13:17, 13:20, 13:22:01-13:23, 13:25-13:26, 13:28-13:29, 13:36, 13:39, 13:43, 13:50-13:52, 13:57, 13:60-13:61, 13:63N, 14:10, 15:02:01-15:02:06, 15:13:01-15:13:02, 15:20-15:21, 15:25:01-15:25:03, 15:36, 15:44, 15:62, 15:77, 15:80, 15:85, 15:88-15:89, 15:106, 15:112, 15:121, 15:139, 15:144, 15:154, 15:165, 15:170, 15:194, 15:204, 15:213-15:214, 15:223, 15:240, 15:250, 15:265, 15:271, 18:22, 18:69, 27:19, 27:30, 35:01:01:01-35:04:03, 35:06-35:08:06, 35:10-35:17, 35:19-35:21, 35:23-35:30, 35:33-35:36, 35:38-35:42:02, 35:45-35:50, 35:52, 35:54-35:57, 35:59, 35:61:01-35:63, 35:65Q, 35:69-35:71, 35:74, 35:76-35:78, 35:80-35:85, 35:90-35:96, 35:98, 35:100-35:101:02, 35:103-35:113, 35:115-35:116, 35:120-35:126, 35:128-35:134N, 35:136-35:150, 35:152-35:173N, 35:175-35:184, 35:186-35:198, 35:200-35:204, 35:206-35:229, 37:01:01-37:01:07, 37:01:09, 37:03N-37:06, 37:08, 37:10-37:11, 37:13-37:37, 38:20, 39:42, 40:28, 44:02:01:01-44:14, 44:16-44:17, 44:19N, 44:21-44:30, 44:32-44:40, 44:42-44:46, 44:48-44:52N, 44:55-44:64:02, 44:66-44:98, 44:101-44:105, 44:107-44:134, 44:136-44:137, 44:139-44:157, 44:159-44:165, 44:167-44:175, 46:33, 48:02:01-48:02:03, 48:25, 50:17, 51:04, 51:42, 51:46, 51:56:01-51:56:02, 51:139, 53:01:01-53:13, 53:15-53:30, 55:14, 56:09, 56:11-56:12, 57:01:01-57:01:04, 57:01:06-57:10, 57:12, 57:14-57:20, 57:22-57:30, 57:32-57:46, 57:48-57:64, 58:01:01-58:01:02, 58:01:04-58:01:12, 58:04-58:05, 58:09-58:15, 58:17N, 58:19, 58:21-58:24, 58:28-58:29, 58:31N-58:37, 58:39N-58:40, 58:42, 83:01, C*03:102 |
| 38 | 140 bp | 1070 bp | 7, 8, 18, 35, 37, 39, 56, 78 | *07:65, 07:134, 08:32, 15:202, 15:239, 18:01:01:01-18:08, 18:10-18:15, 18:17N-18:36, 18:38-18:47, 18:50-18:53, 18:55-18:65, 18:68-18:78, 18:80-18:89, 35:01:01:01-35:03:02, 35:03:04-35:18, 35:20:01-35:24:02, 35:28-35:45, 35:48, 35:50-35:55, 35:57-35:62, 35:64:01-35:72, 35:74-35:153, 35:155-35:197, 35:199-35:202, 35:204-35:206, 35:208-35:209, 35:211-35:216N, 35:218-35:229, 37:11, 37:37, 39:19:01-39:19:02, 40:185, 40:209, 56:06, 78:01:01-78:05, 78:07 |
| 396,10 | 60 bp, 245 bp, 400 bp | 1070 bp | 15, 62, 63, 70, 71, 72, 75, 76, 77, 40, 44 | *15:01:01:01-15:01:04, 15:01:06-15:21, 15:23-15:40, 15:42-15:58, 15:60-15:67, 15:69-15:99, 15:102-15:129, 15:131-15:179, 15:181N-15:235, 15:237-15:254, 15:256-15:275, 40:12, 82:01-82:03, C*03:179 |
| 409,10 | 210 bp | 800 bp | 7, 8, 15, 18, 37, 38 | *07:27, 08:49, 18:54, 37:01:01-37:37, 38:17, C*07:302 |
| 418 | 170 bp | 800 bp | 15, 35, 38, 39, 67 | *08:55, 15:69, 15:186, 35:35, 38:01:01-38:09, 38:11-38:21, 38:23-38:29, 38:31-38:35, 38:38-38:43, 39:01:01:01-39:01:01:03, 39:01:03-39:20, 39:22-39:24:02, 39:26-39:29, 39:31, 39:35, 39:37-39:42, 39:44-39:46, 39:48-39:49, 39:51-39:73, 39:75-39:81, 40:184, 51:101, 58:20, 67:01:01-67:03, C*15:51 |
| 426,10 | 110 bp | 1070 bp | 38 | *38:01:01-38:02:02, 38:02:04-38:03, 38:05, 38:08-38:24, 38:26-38:34N, 38:36-38:43, 49:26, 52:01:10 |
| 43 | 395 bp | 1070 bp | 14, 38, 39, 64, 65 | *07:69, 07:85, 07:180, 08:65, 14:01:01-14:20, 14:22-14:33, 38:01:01-38:02:05, 38:04-38:05, 38:08-38:25, 38:27-38:43, 39:01:01:01-39:01:01:03, 39:01:03-39:01:16, 39:03-39:07, 39:09, 39:11-39:12, 39:14-39:15, 39:18-39:19:02, 39:22, 39:24:01-39:38Q, 39:40:01N-39:48, 39:50-39:57, 39:59-39:62, 39:64-39:82, 49:26 |
| 44 | 160 bp, 425 bp | 1070 bp | 7, 8, 14, 39, 40, 41, 42, 45, 48, 50, 60, 64, 65, 67, 73 | *07:02:01-07:26, 07:28-07:35, 07:37, 07:39-07:64, 07:66-07:80, 07:82-07:148, 07:150-07:178, 07:181N-07:190, 08:01:01-08:01:20, 08:04-08:05, 08:07-08:35, 08:37-08:51, 08:53:01-08:77, 08:79-08:108, 14:01:01-14:02:04, 14:02:06-14:02:07, 14:02:09-14:29, 14:31-14:33, 27:14, 27:81, 39:01:01:01-39:01:01:03, 39:01:03-39:20, 39:22-39:24:01, 39:25N-39:82, 40:01:01-40:01:29, 40:06:01:01-40:06:04, 40:06:06-40:07, 40:09?, 40:11:01?-40:11:02?, 40:14:01?-40:15?, 40:16, 40:18?, 40:21?-40:22N?, 40:23, 40:24?-40:38?, 40:42?-40:43?, 40:44, 40:45?-40:46?, 40:48?-40:50?, 40:51, 40:52?, 40:53, 40:54?-40:58?, 40:59-40:60, 40:61?-40:65?, 40:66, 40:67?-40:69?, 40:70, 40:71?, 40:72:01-40:73, 40:74?, 40:75, 40:77, 40:78?, 40:79, 40:80?-40:82?, 40:83, 40:84?-40:85?, 40:86, 40:87:01?-40:88?, 40:90?-40:92?, 40:93, 40:95?, 40:96, 40:98?-40:102?, 40:103, 40:104?-40:108?, 40:109-40:110, 40:111?-40:116?, 40:118N?-40:121?, 40:123?, 40:124:01, 40:124:02?-40:126?, 40:127, 40:128?-40:130?, 40:131, 40:132?-40:134?, 40:136?-40:137?, 40:138-40:140, 40:141?, 40:145?, 40:146-40:148, 40:149?, 40:150, 40:151?, 40:152-40:153, 40:154?, 40:155N, 40:156?, 40:158?, 40:159, 40:160:01?-40:160:02?, 40:161-40:162, 40:163?-40:164?, 40:165, 40:166?, 40:167, 40:168?-40:171?, 40:172, 40:173?-40:176?, 40:177, 40:178?-40:185?, 40:186, 40:187?, 40:189?, 40:190, 40:191?-40:200?, 40:202?-40:213?, 40:215?-40:228?, 40:230, 40:231?-40:238?, 41:01-41:27, 42:01:01-42:02, 42:04-42:20, 45:01-45:14, 48:01:01-48:17, 48:19-48:30, 50:01:01-50:02, 50:04-50:20, 50:31-50:32, 51:01:09, 51:01:21, 51:01:25, 51:10, 51:24:02-51:24:04, 54:11, 56:20:02, 67:01:01-67:03, 73:01-73:02, 81:02 |
| 4513 | 180 bp | 1070 bp | 15, 40, 44, 46, 48, 60 | *15:116, 15:124, 40:01:01-40:01:06, 40:01:08-40:01:17, 40:01:19-40:01:29, 40:07, 40:10:01-40:10:02, 40:12, 40:21-40:23, 40:25, 40:30, 40:33-40:34, 40:36, 40:38, 40:42-40:43, 40:46-40:49, 40:51-40:52, 40:54-40:55, 40:59-40:63, 40:65-40:67, 40:69, 40:73-40:74, 40:76, 40:79, 40:81, 40:84, 40:87:01-40:88, 40:92, 40:100-40:102, 40:106, 40:108, 40:112-40:114, 40:116-40:118N, 40:123-40:126, 40:128, 40:130, 40:132, 40:134-40:135, 40:137-40:141, 40:146-40:147, 40:149-40:156, 40:160:01-40:160:02, 40:163, 40:166, 40:168, 40:170-40:172, 40:175, 40:178-40:179, 40:182-40:183, 40:185-40:188, 40:191, 40:193-40:196, 40:198-40:199, 40:204, 40:207-40:208, 40:210, 40:212-40:213, 40:215-40:218, 40:221-40:223, 40:227-40:228, 40:231, 40:234-40:236, 40:238, 44:31, 44:169, 46:06, 48:03:01-48:03:02, 48:17, 48:23, 57:59, C*03:129, C*17:16 |
| 466 | 105 bp | 1070 bp | 7, 8, 35, 41, 42, 55 | *07:04, 07:25, 07:146, 08:01:01-08:05, 08:07-08:12:03, 08:14-08:19N, 08:21-08:24, 08:26-08:39, 08:41-08:54, 08:56:01-08:69, 08:71-08:78, 08:80-08:102, 08:104-08:108, 35:87, 40:136, 40:231, 41:01-41:08, 41:10-41:27, 42:01:01-42:02, 42:04-42:07, 42:09-42:18, 42:20, 44:166, 55:20 |
| 47 | 325 bp | 800 bp | 7, 15, 27, 35, 42, 45, 46, 54, 55, 56, 67, 73, 81 | *07:02:01-07:02:38, 07:04-07:07, 07:09-07:15, 07:17-07:26, 07:28-07:31, 07:33:01-07:36, 07:39-07:46, 07:47w, 07:48-07:49N, 07:51-07:68:03, 07:70-07:84, 07:86-07:179, 07:181N-07:190, 14:21, 15:76, 15:101, 15:255, 27:01-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:76, 27:78-27:91, 27:93-27:105, 35:76, 35:186, 35:229, 38:26, 40:166, 42:01:01-42:02, 42:04-42:06, 42:08-42:10, 42:12-42:20, 44:90, 44:97, 45:06, 46:01:01-46:01:09, 46:02w, 46:03-46:38, 54:01:01-54:28, 55:01:01-55:05, 55:07-55:17, 55:19-55:58, 56:01:01-56:16, 56:18-56:22, 56:24-56:32, 56:34-56:39, 67:01:01-67:03, 73:01-73:02, 81:01-81:06, 82:01-82:03, 83:01 |
| 486 | 80 bp, 115 bp, 160 bp, 195 bp, 225 bp, 260 bp | 1070 bp | 35, 37, 44, 47, 51, 57 | *08:49, 08:60, 08:76, 08:87, 35:38, 35:45, 35:71, 35:115, 35:169, 37:01:01-37:01:07, 37:01:09, 37:03N-37:06, 37:08, 37:10-37:37, 40:132, 40:178, 40:231, 42:13, 44:02:01:01-44:03:16, 44:03:18-44:14, 44:16-44:17, 44:19N-44:36, 44:38-44:63, 44:65-44:134, 44:136-44:149N, 44:151-44:157, 44:159-44:165, 44:167-44:175, 51:42, 53:22, 57:07, 57:09, 57:24, 83:01, C*15:25 |
| 49 | 130 bp, 270 bp | 800 bp | 18, 35, 37, 40, 46, 51, 53, 57, 58 | *14:10, 15:57w, 18:22, 18:69, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 35:188, 35:190, 37:04:01-37:04:02, 40:28, 46:01:01-46:38, 51:04, 51:46, 51:56:01-51:56:02, 51:139, 53:02, 53:06, 53:28, 57:14, 58:09, C*01:73, C*15:39 |
| 506,14 | 90 bp, 410 bp | 1070 bp | 7, 27, 35, 47, 54, 55, 56, 81 | *07:65, 27:01w, 27:02:01-27:11, 27:13-27:15, 27:17, 27:19-27:21, 27:24-27:28, 27:30-27:38, 27:40-27:58, 27:60-27:76, 27:78-27:91, 27:93-27:105, 35:76, 35:186, 35:229, 44:90, 44:97, 47:01:01:01-47:03, 47:06-47:08, 54:01:01-54:28, 55:01:01-55:05, 55:07-55:17, 55:19-55:58, 56:01:01-56:16, 56:18-56:22, 56:24-56:32, 56:34-56:39, 81:01, 82:01-82:03, 83:01 |
| 516 | 90 bp, 175 bp | 1070 bp | 15, 27, 35, 40, 48, 62, 75, 77, 81 | *07:174, 13:36, 15:02:01-15:02:06, 15:08, 15:11:01-15:11:05, 15:13:01-15:13:02, 15:15, 15:21, 15:31, 15:44, 15:55, 15:76w, 15:88-15:89, 15:112, 15:121, 15:139, 15:144, 15:148, 15:170, 15:189, 15:191, 15:194, 15:209N, 15:213-15:215, 15:223, 15:265, 27:24, 35:46, 40:31, 40:45, 40:80, 48:01:01-48:01:04, 48:04, 48:06-48:07, 48:09, 48:11, 48:15-48:16, 48:18-48:20, 48:22, 48:24, 48:27-48:30, 56:39w, 81:01-81:06 |
| 52 | 145 bp, 430 bp | 1070 bp | 40, 44, 49 | *40:13, 40:19, 40:109, 40:117, 44:18, 44:25, 44:50, 44:95, 49:01:01-49:01:04, 49:04-49:25, 51:112, 53:30, 54:12, 56:21, 57:45, 57:51, 59:01:01:01-59:05 |
| 536 | 120 bp, 210 bp | 800 bp | 13, 15, 45, 49, 50, 51, 52, 55, 56, 58, 62, 63, 78 | *07:78, 07:84, 13:16, 13:31w, 13:48, 13:62, 15:04w, 15:16:01-15:16:03, 15:42, 15:67, 15:95, 15:137w, 15:254, 40:95, 40:148, 40:161, 46:32w, 49:01:01, 49:01:03-49:10, 49:12-49:26, 50:01:01-50:02, 50:04-50:11, 50:13-50:16, 50:18-50:20, 50:31-50:32, 51:01:01-51:03, 51:05, 51:07:01-51:09:02, 51:11N-51:14, 51:16-51:24:04, 51:26-51:41N, 51:43-51:44N, 51:48-51:55, 51:57-51:58, 51:60-51:61:02, 51:63, 51:65-51:80, 51:82-51:134, 51:136-51:138, 51:140-51:147, 51:149N-51:152, 52:01:01:01-52:31, 54:20, 55:01:01-55:01:06, 55:01:08-55:01:10, 55:03, 55:05, 55:09, 55:11, 55:15, 55:17, 55:21w, 55:24-55:25, 55:28-55:29, 55:31, 55:33, 55:36, 55:38, 55:40, 55:44-55:45, 55:52-55:55N, 55:58, 56:05:01-56:06, 56:21, 56:25, 56:37, 58:08:01-58:08:02, 78:01:01-78:07 |
| 54 | 430 bp | 1070 bp | 27, 40, 44, 45, 51, 52, 53, 57, 58, 59 | *18:67, 27:02:01-27:02:02, 27:30, 27:53, 27:57, 27:62, 27:65N, 27:75, 27:77, 27:83, 27:95, 27:102, 37:34, 40:13, 40:19, 40:109, 40:117, 44:06, 44:25, 44:50, 44:95, 51:01:01-51:01:39, 51:01:41-51:24:04, 51:26-51:46, 51:48-51:53, 51:55-51:77, 51:79-51:152, 52:01:01:01-52:19, 52:21-52:31, 53:01:01-53:02, 53:04-53:08:02, 53:10, 53:14-53:30, 57:01:01-57:11, 57:13-57:64, 58:01:01-58:02, 58:04-58:16:01, 58:18-58:29, 58:31N-58:40, 58:42 |
| 558 | 145 bp | 1070 bp | 15, 35, 40, 41, 44, 45, 47, 50, 72 | *07:133, 15:46, 15:53, 15:106, 15:143, 15:212, 18:48, 35:19, 35:47, 35:63, 35:154, 35:217, 39:69, 40:01:01-40:11:02, 40:14:01-40:16, 40:18, 40:20, 40:22N-40:40, 40:42-40:45, 40:48-40:75, 40:77-40:92, 40:94-40:95, 40:97-40:108, 40:111-40:116, 40:118N-40:136, 40:138-40:156, 40:158-40:184, 40:186-40:187, 40:189-40:200, 40:202-40:207, 40:210-40:238, 41:01-41:04, 41:05?, 41:06-41:25, 41:27, 44:09, 44:46, 44:75, 44:90, 44:129, 44:131, 45:01-45:14, 47:02, 47:03w, 50:01:01-50:02, 50:04-50:05, 50:07-50:20, 50:31-50:32 |
| 56 | 300 bp | 1070 bp | 7, 8, 13, 15, 18, 27, 35, 37, 38, 39, 40, 41, 44, 45, 46, 47, 48, 49, 50, 52, 55, 57, 58, 62, 72, 76, 78 | *07:54, 07:123, 08:17, 08:38, 08:54, 08:101, 13:01:01-13:04, 13:06, 13:08-13:23, 13:25-13:35, 13:37-13:67, 15:01:01:01-15:01:04, 15:01:06-15:01:18, 15:01:20-15:01:32, 15:03:01-15:07:02, 15:12, 15:14, 15:19-15:20, 15:24:01-15:27:03, 15:30, 15:32-15:36, 15:38:01-15:40, 15:42-15:43, 15:45-15:50, 15:53-15:54, 15:56-15:58, 15:60-15:63, 15:65-15:66, 15:68-15:71, 15:73-15:75, 15:77-15:79N, 15:81-15:87, 15:91-15:92, 15:94N, 15:96-15:98, 15:101-15:107, 15:109-15:111N, 15:113, 15:116-15:118, 15:122-15:123, 15:125-15:129, 15:131-15:132, 15:135-15:138, 15:140-15:142, 15:145-15:147, 15:150-15:152, 15:154-15:160, 15:163-15:167, 15:169, 15:171-15:175, 15:178-15:179, 15:181N-15:185, 15:187-15:188, 15:190N, 15:192-15:193, 15:195, 15:199, 15:201-15:207, 15:210-15:212, 15:217-15:220, 15:224-15:225, 15:227-15:228, 15:231-15:237, 15:239-15:245Q, 15:247-15:251, 15:253, 15:255-15:262N, 15:264, 15:266-15:267, 15:269-15:272N, 15:274, 18:12, 27:01-27:21, 27:24-27:91, 27:93-27:105, 35:10, 35:13, 35:16, 35:28, 35:63, 35:69, 35:80, 35:207, 35:226, 37:01:01-37:07, 37:09-37:37, 38:03-38:04, 38:25, 39:02:01-39:02:02, 39:08, 39:13:01-39:13:02, 39:22-39:23, 39:49, 40:01:01-40:07, 40:09-40:16, 40:18-40:24, 40:26-40:40, 40:42-40:67, 40:70-40:105, 40:107-40:165, 40:167-40:231, 40:233-40:238, 41:01-41:27, 42:11, 44:02:01:01-44:02:21, 44:02:23-44:05:04, 44:07-44:11, 44:13-44:170, 44:172-44:175, 45:01-45:14, 47:01:01:01-47:08, 48:01:01-48:05, 48:07-48:15, 48:17-48:30, 49:01:01-49:25, 50:01:01-50:02, 50:04-50:14, 50:16-50:20, 50:31-50:32, 52:01:01:01-52:31, 53:17:01-53:17:02, 53:28, 55:18, 55:34, 55:40, 78:05-78:06 |
| 578,9,12 | 160 bp, 330 bp | 1070 bp | 15, 37, 44, 54 | *08:87, 44:02:01:01-44:02:29, 44:06, 44:08-44:09, 44:11-44:12, 44:16, 44:19N-44:24, 44:27:01-44:27:02, 44:33-44:34:02, 44:41:01-44:41:02, 44:44, 44:48-44:49, 44:52N-44:53, 44:55, 44:59:01-44:59:02, 44:63, 44:66-44:68, 44:71-44:74, 44:80, 44:83-44:84, 44:86-44:87, 44:89-44:91, 44:93, 44:95, 44:97, 44:99-44:102, 44:104, 44:106, 44:112-44:113, 44:116, 44:118-44:119, 44:121, 44:126:01-44:127, 44:131-44:132, 44:137-44:138Q, 44:142, 44:145, 44:148-44:149N, 44:151-44:152, 44:162, 44:168-44:173, 54:01:01, 54:01:03-54:28, 83:01 |
| 58 | 180 bp, 210 bp | 1070 bp | 7, 13, 15, 18, 35, 39, 40, 45, 54, 55, 56, 59, 78 | *07:65, 13:06, 13:53, 15:42, 15:86, 15:224, 18:04, 35:08:05, 35:42:02, 35:60, 35:185, 39:17, 39:63, 40:01:11, 40:58, 45:08, 46:18, 54:01:01-54:02, 54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:16-54:25, 54:27-54:28, 55:01:01-55:03, 55:05, 55:07, 55:10-55:12, 55:15-55:16, 55:18-55:19, 55:21, 55:25-55:26, 55:29-55:31, 55:33-55:45, 55:47-55:48, 55:50, 55:52-55:57, 56:05:01, 56:06, 56:10, 56:23, 59:01:01:01-59:01:01:02, 59:05, 78:01:01-78:01:02, 78:02:02-78:03, 78:07, C*15:02:04 |
| 59 | 180 bp | 1070 bp | 13, 15, 40, 44, 45, 46, 49, 50, 51, 54, 56 | *13:03, 13:48, 15:73, 40:71, 44:10, 44:15, 44:18, 44:140, 45:01, 45:04-45:07, 45:11-45:14, 46:11, 49:01:01-49:03, 49:06, 49:08-49:17, 49:19N-49:26, 50:01:01-50:02, 50:04-50:08, 50:10-50:13, 50:15-50:19, 50:32, 51:15, 51:62, 51:106, 52:25, 54:03, 56:01:01-56:01:04, 56:02, 56:04, 56:07-56:08, 56:13-56:14, 56:16-56:17, 56:20:01-56:20:02, 56:24-56:30, 56:33-56:39, 59:04, 82:01-82:03, C*03:19, C*03:102 |
| 606,10 | 90 bp, 240 bp | 800 bp | 55, 57 | *55:14, 57:01:01-57:44, 57:46-57:50, 57:52-57:64, 58:36, C*06:72 |
| 616 | 95 bp | 1070 bp | 7, 8, 18, 35, 37, 38, 39, 44, 51, 52, 53, 78 | *07:65w, 07:134w, 08:32, 18:01:01:01-18:11, 18:13-18:15, 18:17N-18:36, 18:38-18:47, 18:49-18:65, 18:67-18:89, 35:01:01:01-35:08:04, 35:08:06-35:09:03, 35:11:01-35:12:03, 35:14:01-35:15, 35:17-35:18, 35:20:01-35:24:02, 35:27, 35:29:01-35:45, 35:48, 35:50-35:62, 35:64:01-35:68:02, 35:70-35:72, 35:74-35:75, 35:76w, 35:77-35:79, 35:81-35:153, 35:155-35:185, 35:187-35:190, 35:192-35:197, 35:199-35:206, 35:208-35:209, 35:211-35:216N, 35:218-35:225, 35:227-35:229, 37:08, 38:06-38:07, 39:19:01-39:19:02, 44:06, 51:01:01-51:24:04, 51:26-51:46, 51:48-51:103, 51:105-51:111, 51:113-51:117, 51:119-51:146, 51:148-51:152, 53:01:01-53:16, 53:18-53:27, 53:29, 56:06w, 78:01:01-78:04, 78:07 |
| 626,8,10 | 115 bp, 150 bp | 1070 bp | 15, 18, 27, 35, 39, 46, 54, 55, 56, 62, 73, 75, 76 | *07:100, 13:31, 13:41, 15:01:01:01-15:01:04, 15:01:06-15:01:16, 15:01:18-15:01:30, 15:01:32, 15:04, 15:07:01-15:08, 15:11:01-15:12, 15:14-15:15, 15:19, 15:24:01-15:24:02, 15:26N-15:28, 15:30, 15:32, 15:34-15:35, 15:38:01-15:38:02, 15:43, 15:45-15:46, 15:50, 15:53-15:54, 15:56-15:58, 15:60, 15:63, 15:66, 15:68, 15:70-15:71, 15:73, 15:75-15:77, 15:79N, 15:81-15:82, 15:85, 15:87, 15:92, 15:94N, 15:96-15:97, 15:101-15:102, 15:104-15:105, 15:109-15:111N, 15:113, 15:117-15:118, 15:120, 15:122, 15:125-15:126, 15:128-15:129, 15:135, 15:137, 15:140, 15:142-15:149N, 15:152, 15:154, 15:157, 15:159-15:160, 15:163-15:167, 15:169, 15:171-15:172, 15:174-15:175, 15:178, 15:180-15:184, 15:187, 15:189-15:193, 15:201-15:203, 15:205-15:207, 15:209N, 15:211-15:212, 15:215, 15:217, 15:225, 15:227-15:228, 15:231-15:234, 15:236, 15:239, 15:241, 15:244-15:247, 15:251, 15:256-15:262N, 15:264, 15:267, 15:269-15:270, 15:272N, 18:19, 27:25, 27:75, 35:14:01-35:14:02, 35:43:01-35:44, 35:62, 35:67, 35:79, 35:86, 35:102, 35:117, 35:135, 35:163, 35:185, 35:213, 39:18, 39:36, 44:146, 46:01:01-46:02, 46:04-46:05, 46:07N, 46:09-46:10, 46:12, 46:14-46:17, 46:20, 46:22-46:24, 46:26-46:32, 46:34-46:38, 51:61:01-51:61:02, 52:21, 54:06, 55:21, 56:03, 73:01-73:02, A*26:68, A*68:56, C*02:56, C*06:20, C*12:50 |
| 6315 | 360 bp | 1070 bp | Bw4 |  |
| 64 | 350 bp | 1070 bp | Bw6 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B low resolution SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 22 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band.

In addition, wells number 25, 26, 27, 31, 37, 40, 41, 47, 49, 53 and 60 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all HLA-B alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

5The HLA-C*03:05, C*03:25, C*03:27, C*03:143 and C*03:167 alleles will be amplified by primer mix  24, the C*01:30 and C*12:87 alleles will be amplified by primer mix 25, the A*23:31, A*24:106, C*07:231 and C*16:10 alleles will be amplified by primer mix 27, the C*07:02:30 and C*08:16:02 alleles will be amplified by primer mix 28, the C*01:73, allele will be amplified by primer mixes 28 and 49,  the C*07:46 allele will be amplified by primer mix 29, the A*24:174 allele will be amplified by primer mix 33, the C*03:102 allele will be amplified by primer mixes 37 and 59, the C*03:179 allele will be amplified by primer mix 39, the C*07:302 allele will be amplified by primer mix 40, the C*15:51 allele will amplified by primer mix 41, the C*03:129 and C*17:16 alleles will be amplified by primer mix 45, the C*15:25 allele will be amplified by primer mix 48, the C*15:39 alleles will be amplified by primer mix 49, the C*15:02:04 allele will amplified by primer mix 58, the C*03:19 allele will be amplified by primer mix 59, the C*03:102 alleles will be amplified by primer mix 65, the C*06:72 allele will be amplified by primer mix 60 and the A*26:68, A*68:56, C*02:56, C*06:20 and  C*12:50 alleles will be amplified by primer mix 62.

6HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

7Primer mix 25 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

8Primer mixes 25, 26, 30, 41, 55, 57 and 62 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes.

9Primer mixes 27, 30, 40 and 57 may have tendencies of unspecific amplifications, most pronounced in primer mix 30.

10Primer mixes 22, 34, 35, 39, 40, 42, 60 and 62 have a tendency to giving rise to primer oligomer formation.

11The B*15, B*46, B*57, B*58 and C*03 alleles may be faintly amplified by primer mix 30.

12Primer mix 57 may give rise to a lower yield of B*54 alleles than the other B low primer mixes.

13The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 45.

14Primer mix 50 might generate a false band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

15The Bw4-associated HLA-A specificities A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 63.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the B*0701 allele has been shown to be in error.

The B*08:06 allele has been shown to be identical to B*08:20.

The sequence of the B*13:05 allele has been shown to be identical to B*13:04.

The B*13:24 allele has been renamed B*13:22:02

The B*15:01:05 allele has been corrected and renamed B*15:120

The B*15:22 allele has been renamed B*35:43.

The sequence of the B*15:41 allele has been shown to be identical to B*15:39.

The B*1559 allele has been renamed B*35:44.

The B*15:100 allele has never been assigned.

The sequence of the B*1816 allele has been shown to be identical to B*18:14.

The sequence of the B*27051 allele has been shown to be identical to B*27:05:02.

The sequence of the B*2722 allele has been shown to be identical to the corrected B*27:06 sequence.

The B*3573 allele has been renamed B*35:08:03.

The B*35:43:02 allele has been renamed B*35:185.

The sequence of the B*39012 allele has been shown to be identical to B*39:01:01:01.

The sequence of the B*3921 allele has been shown to be identical to B*39:24.

The sequence of the B*4017 allele has been shown to be identical to B*40:16.

The sequence of the B*4041 allele has been shown to be identical to B*40:40.

The sequence of the B*4203 allele has never been assigned.

The sequence of the B*4401 allele has been shown to be identical to B*44:02:01:01.

The sequence of the B*5003 allele has been shown to be identical to B*50:02.

The sequence of the B*5125 allele has been shown to be identical to B*51:22.

The B*5147 allele has been renamed B*51:09:02.

The sequence of the B*5506 allele has been shown to be identical to B*55:04.

The sequence of the B*5803 allele has never been assigned.

The sequence of the B*58:30 allele has been shown to be identical to B*58:01:01.

The B*7901 allele has been renamed B*15:18:01.

The B*9530 allele has been renamed B*15:27:02.

2The B*07:174 and the 81:02 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The B*08:26, 08:50, 08:62, 08:85 and 08:94 and the B*42:07 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

4The B*13:04, 13:10, 13:21, 13:35 and 13:59 and the 44:135 and 44:158 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5The B*14:08 and the B*39:25N, 39:30, 39:32-39:34, 39:43, 39:47, 39:50, 39:74 and 39:82 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

6The B*18:29 and 18:72 and the B*35:09:01-35:09:03, 35:18, 35:31-35:32:02, 35:37, 35:53N, 35:64:01-35:64:02, 35:68:01-35:68:02, 35:75, 35:88, 35:99, 35:118-35:119, 35:127, 35:151, 35:174 and 35:205 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

7The B*41:09 and the B*45:02-45:03 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

8The B*51:104, 51:118N and 51:147 and the B*58:08:01-58:08:02 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

9The B*51:148 and the B*53:14 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

10The B*53:30 and the  B*57:45 and 57:51 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets

11The B*54:01:02, 55:01:07, 55:02:01-55:02:06, 55:07, 55:10, 55:12, 55:16, 55:19, 55:26, 55:30, 55:35, 55:37, 55:39, 55:41-55:43, 55:47-55:48, 55:50 and 55:57 and the B*56:10 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

12The B*55:04, 55:08, 55:13, 55:23, 55:27, 55:32, 55:46, 55:49 and the B*56:01:05, 56:15, 56:18-56:19N, 56:22 and 56:31-56:32 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

13The B*57:01:01-57:01:04, 57:01:06-57:03:02, 57:05-57:06, 57:08, 57:10, 57:15-57:20, 57:22-57:23, 57:25-57:30, 57:32-57:44, 57:46, 57:48-57:50, 57:52-57:58 and 57:60-57:64 and the B*58:36 alleles give rise to identical amplification patterns with the HLA-B low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

14The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

‘ser’, serological HLA specificity.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Specificities and sizes of the PCR products of the 23 primer mixes used for DR low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DR serology3 | Amplified HLA-DRB alleles4 |
| --- | --- | --- | --- | --- |
| 656,7 | 205 bp, 255 bp | 515 bp | 1 | *01:01:01-01:02:08, 01:04-01:38, 01:40N-01:58 |
| 66 | 200 bp | 430 bp | 103 | *01:03, 01:39N, 01:42 |
| 67 | 200 bp, 220 bp | 430 bp | 15 | *15:01:01:01-15:94 |
| 689 | 210 bp | 430 bp | 16 | *16:01:01-16:05:02, 16:07-16:23 |
| 695,6,11 | 120 bp, 220 bp | 430 bp | 3, 17, 18, 11 | *03:01:01:01-03:75, 03:77-03:84, 03:86-03:90, 11:07, 11:53, 11:103, 11:105, 11:107, 11:125, 15:25 |
| 705,11 | 80 bp, 210 bp | 430 bp | 3, 17, 11, 13, 14 | *03:01:01:01-03:01:22, 03:04:01-03:06, 03:08-03:16, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:31, 03:33-03:34, 03:36-03:37, 03:43-03:48, 03:50-03:52, 03:54-03:68N, 03:70-03:73, 03:75-03:84, 03:86, 03:89, 08:40, 11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-11:41, 11:48, 11:59, 11:63, 11:65:01-11:65:02, 11:68, 11:70, 11:73, 11:76, 11:79-11:80, 11:83, 11:85-11:87, 11:93, 11:118, 11:122, 11:124, 11:127, 11:131-11:132, 11:135, 11:138-11:139, 11:142, 13:01:01-13:04, 13:08, 13:10, 13:15-13:17, 13:19-13:20, 13:22-13:24, 13:27-13:29, 13:31-13:41, 13:43, 13:45, 13:48, 13:51-13:54, 13:57, 13:59, 13:61:01-13:61:02, 13:63-13:66:02, 13:68-13:76, 13:78-13:81, 13:83-13:85, 13:87-13:99, 13:101-13:102, 13:104-13:107, 13:109, 13:111-13:117, 13:120-13:131, 13:133, 13:135, 13:137N-13:145, 13:147-13:149, 13:151-13:153, 13:155, 13:159, 14:16, 14:19, 14:21, 14:82, 14:95, 14:109, 14:132, 14:137N |
| 715 | 85 bp, 210 bp | 430 bp | 3, 18, 11, 13, 14 | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 03:74, 03:88, 03:90, 11:13:01w-11:13:02w, 11:26, 11:34, 13:15, 13:19, 13:26:01-13:26:02, 13:44, 13:53, 13:57, 13:85-13:86, 13:104, 14:02-14:03:02, 14:06:01-14:06:02, 14:09, 14:12:01-14:13, 14:17-14:21, 14:24, 14:27, 14:29-14:30, 14:32:01w-14:32:02w, 14:33, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:65w, 14:67, 14:77-14:78, 14:80-14:81, 14:83, 14:85, 14:89, 14:94, 14:98, 14:102, 14:106, 14:108-14:109, 14:115, 14:119, 14:121, 14:135 |
| 725,6,7 | 100 bp, 175 bp | 430 bp | 3, 4 | *04:01:01-04:05:11, 04:05:13-04:158N |
| 73 | 210 bp, 235 bp | 430 bp | 7, 13, 14 | *07:01:01:01-07:01:04, 07:03-07:24, 12:22, 13:17, 13:116, 14:50 |
| 746 | 170 bp, 215 bp, 250 bp | 515 bp | 8, 12, 14 | *08:01:01-08:19, 08:21-08:53, 11:67, 12:04, 12:16:01-12:16:02, 12:22, 12:39, 14:11, 14:15, 14:68, 14:93 |
| 755,6 | 90 bp, 135 bp, 180 bp | 430 bp | 3, 9, 11 | *03:08, 03:65, 09:01:02-09:20, 11:07, 11:53, 11:103, 11:105, 11:107, 11:125 |
| 767 | 175 bp | 430 bp | 10 | *03:76, 10:01:01-10:05, 11:59, 11:80, 11:83, 11:87, 11:135, 11:142, 13:27, 13:41, 13:71, 13:129 |
| 775,6 | 100 bp, 170 bp | 430 bp | 11, 3, 8 | *03:08, 03:65, 08:31, 08:41, 11:01:01-11:70, 11:72-11:145 |
| 785,6,9 | 85 bp, 105 bp | 430 bp | 12, 8 | *08:32, 08:53, 12:01:01-12:39 |
| 797,8 | 215 bp | 430 bp | 6, 8, 11, 13, 14, 1403 | *03:76, 08:20-08:21, 11:01:01-11:04:11, 11:06:01-11:06:03, 11:08:01-11:12:02, 11:14:01-11:16, 11:18-11:21, 11:23:01-11:25, 11:27:01-11:33, 11:35-11:51, 11:54:01-11:54:02, 11:56-11:66, 11:68, 11:70, 11:72-11:81, 11:83-11:88, 11:90-11:97, 11:99-11:102:02, 11:106, 11:108-11:124, 11:126-11:135, 11:137-11:142, 11:144-11:145, 13:01:01-13:02:01, 13:02:03-13:08, 13:10-13:16, 13:18-13:43, 13:45-13:85, 13:87-13:115, 13:117-13:128, 13:130-13:145, 13:147-13:161, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:16, 14:19, 14:21-14:22, 14:25, 14:27, 14:40, 14:53, 14:63, 14:67, 14:69, 14:74, 14:77-14:78, 14:84-14:85, 14:98, 14:102, 14:105, 14:109, 14:115-14:116, 14:128, 14:135, 14:137N, DRB3*02:27 |
| 806,7,8 | 195 bp, 225 bp | 430 bp | 13, 8, 11, 12, 14 | *08:01:01-08:02:04, 08:04:01-08:09, 08:11, 08:16-08:17, 08:20-08:22, 08:24, 08:26, 08:28, 08:31, 08:39, 08:41-08:44, 08:50, 08:52, 11:01:01-11:01:17, 11:02:01-11:06:03, 11:09-11:12:02, 11:14:01-11:16, 11:20-11:21, 11:23:01-11:25, 11:27:01-11:30, 11:32-11:33, 11:35-11:41, 11:43-11:44, 11:46:01-11:51, 11:54:01-11:56, 11:58:01-11:63, 11:65:01-11:70, 11:72, 11:74:01-11:78, 11:80-11:88, 11:90-11:97, 11:99-11:102:02, 11:106, 11:108-11:118, 11:120-11:124, 11:126-11:129, 11:133-11:135, 11:137-11:142, 11:144-11:145, 12:02:01-12:02:05, 12:13, 12:15-12:16:02, 12:18-12:21, 12:23, 12:26-12:27, 12:31N-12:33, 12:37, 13:01:01-13:02:01, 13:02:03-13:02:08, 13:04-13:05:02, 13:07:01-13:09, 13:11:01-13:11:02, 13:14:01-13:24, 13:26:01-13:29, 13:31-13:32, 13:34-13:36, 13:38-13:43, 13:45-13:55, 13:57, 13:59, 13:61:01-13:65, 13:67-13:76, 13:78-13:80, 13:83-13:84, 13:87, 13:91-13:93, 13:96:01-13:100, 13:102-13:109, 13:111-13:114, 13:116-13:117, 13:121, 13:123-13:132, 13:135-13:136, 13:138-13:150, 13:153, 13:155, 13:158-13:160, 14:15-14:16, 14:22, 14:24-14:25, 14:27, 14:37, 14:53, 14:73, 14:105, 14:128 |
| 8111 | 175 bp | 430 bp | 3, 13, 14, 8 | *03:01:01:01-03:07, 03:09, 03:11:01-03:41, 03:43-03:45, 03:47-03:63, 03:66-03:68N, 03:70-03:84, 03:86, 03:88-03:90, 08:20, 13:01:01-13:16, 13:18-13:42, 13:44, 13:46-13:66:02, 13:68-13:102, 13:104-13:115, 13:117-13:121, 13:123-13:158, 13:161, 14:02-14:03:02, 14:05:01-14:06:02, 14:09, 14:12:01-14:14, 14:17-14:21, 14:23:01, 14:23:03-14:24, 14:27, 14:29-14:30, 14:33, 14:36-14:37, 14:40-14:45, 14:47-14:48, 14:51, 14:56, 14:59, 14:63-14:65, 14:67, 14:77-14:78, 14:80-14:81, 14:83-14:85, 14:89, 14:91, 14:94-14:96, 14:98, 14:100, 14:102-14:103, 14:106, 14:108-14:109, 14:115-14:116, 14:121, 14:123, 14:127:01-14:127:02, 14:134-14:136 |
| 825 | 100 bp, 140 bp, 155 bp | 430 bp | 14, 4, 8,13 | *04:62, 04:69, 04:73, 04:105:01-04:105:02, 04:122, 04:146, 08:08, 11:69, 11:82, 13:45, 14:01:01-14:01:02, 14:01:04, 14:04, 14:07:01-14:07:02, 14:10, 14:16, 14:22, 14:25-14:26, 14:28, 14:31-14:32:03, 14:35, 14:37-14:39, 14:49-14:50, 14:53-14:54:01, 14:55, 14:57-14:58, 14:60-14:62, 14:68-14:71, 14:73-14:76, 14:79, 14:82, 14:86-14:88, 14:90, 14:93, 14:99, 14:101, 14:104-14:105, 14:107, 14:110-14:114, 14:117-14:120, 14:122, 14:124-14:125, 14:128-14:129, 14:131, 14:137N-14:140, DRB4*01:03:01:02N |
| 835,6,7,10 | 110 bp, 135 bp, 170 bp | 430 bp | 14, 3, 9, 11, 12, 13, 15 | *03:10, 09:01:02-09:01:05, 09:01:07-09:02:02, 09:04-09:20, 11:13:01-11:13:02, 11:17, 11:52, 13:43, 13:159, 14:01:01-14:02, 14:04-14:11, 14:13-14:14, 14:16-14:18, 14:19w, 14:20, 14:21w, 14:22-14:23:04, 14:26, 14:28-14:36, 14:38-14:39, 14:41, 14:43-14:52, 14:54:01-14:57, 14:59-14:62, 14:64-14:65, 14:68, 14:70-14:76, 14:79-14:83, 14:86-14:88, 14:90-14:97, 14:99-14:101, 14:103-14:108, 14:109w, 14:110-14:114, 14:117-14:127:02, 14:129-14:134, 14:137N-14:140, 15:27, 15:34, 15:66 |
| 845,6 | 110 bp, 175 bp, 225 bp | 430 bp | 14, 3, 8, 11, 13, 15, 16 | *03:10, 08:09, 08:20-08:21, 08:32, 08:35, 08:36:02, 08:53, 11:13:01-11:13:02, 11:17, 11:23:01, 11:25, 11:31, 11:45, 11:52, 11:55, 11:64, 11:89, 11:96, 11:119, 13:13, 13:18, 13:43, 13:45, 13:47, 13:55, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158-13:159, 14:01:01-14:01:04, 14:03:01-14:05:04, 14:07:01-14:08, 14:10-14:12:02, 14:14-14:16, 14:18, 14:22-14:23:04, 14:25-14:28, 14:31-14:32:03, 14:34-14:36, 14:38-14:40, 14:42-14:45, 14:49-14:50, 14:53-14:65, 14:67-14:79, 14:81-14:82, 14:84-14:93, 14:95-14:97, 14:99-14:105, 14:107, 14:110-14:120, 14:122-14:140, 15:21w, 16:04w, 16:18w |
| 856,11 | 160 bp, 240 bp | 430 bp | 52 | DRB3*01:01:02:01-01:15, DRB3*02:01-02:29N, DRB3*03:01:01-03:03 |
| 867,12 | 215 bp | 430 bp | 53 | DRB4*01:01:01:01-01:08 |
| 87 | 175 bp | 430 bp | 51 | DRB5*01:01:01-01:14, DRB5*02:02-02:06 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 67, 72, 82, 83 and 84.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 65 contains the primer pair giving rise to the longer, 515 bp, internal positive control band. In addition, well number 74 contains the primer pair giving rise to the longer, 515 bp, internal positive control band.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DRB alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4For several DRB1 alleles 1st and/or 3rd exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 65 and 69, 72, 74, 75, 77, 78, 80 and 83 to 85.

7Primer mixes 65, 72, 76, 79, 80, 83 and 86 have a tendency to giving rise to primer oligomer formation.

8Primer mixes 79 and 80 may give rise to a lower yield of HLA-specific PCR product than the other DR low resolution primer mixes.

9Primer mixes 68 and 78 may have tendencies of unspecific amplifications.

10Primer mix 83 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

11Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 69, 70 and 81 in addition to primer mix 85.

12The DRB4*01:03:01:02N allele is amplified by primer mix 86, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer mix.

‘w’, might be weakly amplified.

1The DRB1*04:05:12 allele has been shown to be identical to DRB1*04:05:14.

The DRB1*0702 allele has been shown to be identical to DRB1*07:01:01:01.

The DRB1*08031 allele has been shown to be identical to DRB1*08:03:02.

The DRB1*09011 allele has been shown to be identical to DRB1*09:01:02.

The DRB1*03:11:02 allele has been renamed DRB1*03:81.

The DRB1*1171 allele has been shown to be identical to DRB1*11:02:01.

The DRB1*12031 allele has been shown to be identical to DRB1*12:01:01.

The DRB1*1466 allele has been shown to be identical to DRB1*14:23:02.

The DRB1*1606 allele has been shown to be identical to DRB1*16:05:01.

The DRB3*010101 allele has been shown to be identical to DRB3*01:01:02:01.

The DRB4*0101102N allele has been renamed DRB4*01:03:01:02N.

The DRB5*0201 allele has been shown to be identical to DRB5*02:02.

The DRB4*02:01N and DRB4*03:01N null alleles will not be amplified by the DR low resolution primer set.

2The DRB1*08:09 and the DRB1*14:15 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The DRB1*08:20 and the DRB1*13:18, 13:47, 13:55 and 13:158 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

4The DRB1*08:31, 08:41 and DRB1*11:67 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

5The DRB1*13:13, 13:119, 13:154, 13:156 and the DRB1*14:84 and 14:116 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

6The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

‘ser’, serological HLA specificity.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information for the primer matching region(s) is not known.

Specificities and sizes of the PCR products of the 8 primer mixes used for DQ low resolution SSP typing

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DQ serology3 | Amplified DQB1 alleles4 |
| --- | --- | --- | --- | --- |
| 88 | 225 bp | 515 bp | 5 | *05:01:01:01-05:26 |
| 895 | 220 bp, 275 bp | 515 bp | 1, 5, 6 | *06:01:01-06:84 |
| 90 | 210 bp | 430 bp | 2 | *02:01:01-02:29 |
| 916 | 220 bp | 515 bp | 3, 7 | *03:01:01:01-03:01:16, 03:04, 03:09-03:10, 03:13-03:14:02, 03:16, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36, 03:42, 03:44, 03:46-03:60, 03:69, 03:71 |
| 926 | 130 bp | 515 bp | 6, 8 | *03:02:01-03:02:07, 03:05:01-03:05:04, 03:07-03:08, 03:11, 03:18, 03:32, 03:37, 03:45, 03:61, 03:63-03:64, 03:66N-03:68, 03:70, 06:29 |
| 936,7 | 135 bp | 515 bp | 2, 3, 9 | *02:03, 03:03:02:01-03:03:06, 03:06, 03:12, 03:15, 03:20, 03:25-03:26, 03:30-03:31, 03:33-03:34, 03:38-03:41, 03:43, 03:65, 04:03:01-04:03:02, 06:51, 06:66 |
| 947,8 | 85 bp, 185 bp | 515 bp | 3, 7, 8, 9 | *03:01:01:01-03:72 |
| 956 | 160 bp | 430 bp | 4 | *04:01:01-04:09, 04:11-04:12 |
| 969 |  |  |  | Negative control |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQ low resolution SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 88 contains the primer pair giving rise to the longer, 515 bp, internal positive control band. In addition, wells number 89 and 91 to 94 contain the primer pair giving rise to the longer, 515 bp, internal positive control band.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological split of the DQB1*05:05-05:26, DQB1*06:06 to 06:07 alleles, the DQB1*06:10, 06:13, 06:15-06:24 and 06:27 to 06:84, the DQB1*02:04-02:29, the DQB1*03:07-03:09 and 03:11- 03:72 alleles and the DQB1:04:0301-04:12 alleles is not known. In this table we have used the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170, and also inferred the serological grouping from the naming of the sequence-defined allele.

4For several DQB1 alleles 1st and/or 3rd exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

5Primer mix 89 has a tendency to giving rise to primer oligomer formation.

6Primer mixes 91, 92, 93 and 95 may give rise to a lower yield of HLA-specific PCR product than the other DQ low resolution primer mixes.

7Primer mixes 93 and 94 may have tendencies of unspecific amplifications.

8The primer pair in well 94 will in some samples give rise to two HLA-specific PCR fragments and may give rise to a lower yield of HLA-specific PCR product than the other DQ low primer mixes.

9Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1The sequence of the DQB1*03031 allele has been shown to be identical to DQB1*03:03:02.

2The serological split of the DQB1*05:05-05:26, DQB1*06:06 to 06:07 alleles, the DQB1*06:10, 06:13, 06:15-06:24 and 06:27 to 06:84, the DQB1*02:04-02:29, the DQB1*03:07-03:09 and 03:11- 03:72 alleles and the DQB1:04:0301-04:12 alleles is not known. In this table we have used the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170, and also inferred the serological grouping from the naming of the sequence-defined allele.

3Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

Changes in revision R01 compared to R00:

- Primer mix 39 does not amplify the B*44:10 allele. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- Primer mix 18 does not amplify the A*74:07 allele, and primer mix 10 does not amplify the A*31:03 and A*66:15 alleles. This has been corrected in the Specificity and Interpretation tables.

The A*32:01:14, 32:03, 32:38 and the A*74:07 alleles give rise to identical amplification patterns with this lot of the HLA-A low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

Changes in revision R03 compared to R02:

- Primer mix 68 may have tendencies of unspecific amplifications. A footnote has been added in the Specificity Table.