Gel Documentation Form and Worksheet

HLA-A-B-DR-DQ Combi Tray	Lot No: 65Y	Expiry Date: 2017-December-01

(101.708-24/24u)

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):__________

Test  Date:_______________

Tested By:

Review Date:_________________

Reviewed By:_________________

Interpretation:___________	Failed lanes: ___________                                Comments:_________________

Gel Picture

PHOTO DOCUMENT

HLA-A low  resolution

HLA-B low  resolution

HLA-DR low  resolution

HLA-DQ low  resolution

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-A low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

The primer pairs in vials 1, 2, 9, 10 and 13 will in many samples give rise to two or three HLA-specific PCR fragments.

Primer mix 5 may give rise to a lower yield of A*23 alleles than the other A low primer mixes.

Primer mix 8 may weakly amplify the A*34 alleles.

Primer mix 13 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.

Primer mix 3 may faintly amplify the A*30:04:01-30:04:02, 30:06, 30:17, 30:29, 30:46 and 30:77 alleles.

Primer mixes 5, 9, 13, 14, 19 and 21 may have a tendency of giving rise to primer oligomer formation.

Primer mix 19 may have tendency of unspecific amplification.

Primer mix 20 may generate a false positive band of about 560 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings.

HLA-B low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 25 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes in B*40, B*41, B*45, B*49 and B*50 alleles.

Primer mixes 25, 26, 30, 41, 55, 57 and 62 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B low resolution primer mixes.

Primer mixes 27, 30, 40 and 57 may have tendencies of unspecific amplifications, most pronounced in primer mix 30.

Primer mixes 22, 24, 34, 35, 39, 40, 42, 60 and 62 have a tendency to giving rise to primer oligomer formation.

The B*15, B*46, B*57, B*58 and C*03 alleles may be faintly amplified by primer mix 30.

Primer mix 57 may give rise to a lower yield of B*54 alleles than the other B low primer mixes.

The C*17:01 to C*17:04 alleles might be faintly amplified by primer mix 45.

Primer mix 50 might generate a false band of about 800 base pairs. This band should be disregarded when interpreting HLA-B low resolution typings.

The Bw4-associated HLA-A specificities A23, A24, A2403, A25 and A32 are not amplified by the primer pair in primer mix 63.

DR low resolution primer set

HLA-specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 65 and 69, 72, 74, 75, 77, 78, 80 and 83 to 85.

Primer mixes 65, 69, 72, 76, 79, 80, 82, 83 and 86 have a tendency to giving rise to primer oligomer formation.

Primer mixes 79 and 80 may give rise to a lower yield of HLA-specific PCR product than the other DR low resolution primer mixes.

Primer mixes 67, 68, 73 and 78 may have tendencies of unspecific amplifications.

Primer mix 83 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

Due to sharing of sequence motifs in codon 38 and 47, DRB3*01:14 will also be amplified in primer mixes 69, 70 and 81, and DRB3*01:23 in primer mix 69, in addition to primer mix 85.

The DRB4*01:03:01:02N allele is amplified by primer mix 86, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer mix.

DQ low resolution primer set

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 89 has a tendency to giving rise to primer oligomer formation.

Primer mixes 91, 92, 93 and 95 may give rise to a lower yield of HLA-specific PCR product than the other DQ low resolution primer mixes.

Primer mixes 93 and 94 may have tendencies of unspecific amplifications.

The primer pair in well 94 will in some samples give rise to two HLA-specific PCR fragments and may give rise to a lower yield of HLA-specific PCR product than the other DQ low primer mixes.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

HLA-A low resolution Interpretation Table

1HLA-A, HLA-B, HLA-DRB and HLA-DQB1 alleles listed on the IMGT/HLA web page 2015-January-19, release 3.19.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-A alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4The A*36 alleles will give rise to identical amplification patterns as a number of A*01 alleles. These alleles can be separated by the A*01 and A*36 high resolution SSP primer sets.

5The A*01:136 and the 11:94, 11:112 and 11:211 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mixes 1 and 10.

6The HLA-A alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns:

| Alleles | Alleles |
| --- | --- |
| A*03:01:23, 03:08, 03:32, 03:36N, 03:57, 03:59, 03:72, 03:89, 03:107-03:108, 03:111, 03:142, 03:152, 03:172-03:173, 03:176, 03:178N, 03:198-03:199, 03:203, 03:205, A*24:92 | A*23:57, 23:66, A*24:94, 24:138, 24:188, 24:228 |
| A*03:88, A*11:130 | A*31:01:02:01-31:02, 31:05-31:28, 31:30-31:34, 31:36-31:84, 31:86-31:89, A* 33:53 |
| A*23:01:01-23:01:12, 23:01:14-23:01:16, 23:03:01-23:13, 23:15-23:56, 23:58-23:65, 23:67-23:69, B*18:27 | A*32:01:18, B*07:81, B*08:52, B*18:67, B*38:41, B*51:185, B*53:05, 53:16, 53:33 |
| A*23:14:01-23:14:02, A*24:24, 24:71 |

‘w’, might be weakly amplified.

HLA-B low resolution Interpretation Table

1HLA-A, HLA-B, HLA-DRB and HLA-DQB1 alleles listed on the IMGT/HLA web page 2015-January-19, release 3.19.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-B alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4The HLA-B alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns:

| Alleles | Alleles |
| --- | --- |
| B*07:174, 07:202, 07:222, B*81:02 | B*51:104, 51:118N, 51:147, B*58:08:01-58:08:02 |
| B*08:26:01-08:26:02, 08:50, 08:62, 08:85, 08:94, B*42:07 | B*51:148, B*53:14 |
| B*13:04, 13:10, 13:21, 13:35, 13:59, 13:71-13:72, B*44:135, 44:158, 44:184 | B*53:30, 53:39, B*57:45, 57:51, 57:69 |
| B*13:46, B*44:213 | B*55:01:07, 55:02:01:01-55:02:08, 55:07, 55:10, 55:12, 55:16, 55:19, 55:26, 55:30, 55:35, 55:37, 55:39, 55:41-55:43, 55:47-55:48, 55:50, 55:57, 55:61-55:63, 55:65, 55:67, 55:69-55:70, B*56:10 |
| B*14:08:01-14:08:02, 39:01:19, B* 39:25N, 39:30, 39:32-39:34, 39:43, 39:47, 39:50, 39:74, 39:82 | B*55:04, 55:08, 55:13, 55:23, 55:27, 55:32, 55:46, 55:49, B*56:01:05, 56:15, 56:18-56:19N, 56:22, 56:31-56:32 |
| B*18:29, 18:72, 18:92, 18:102, B*35:09:01-35:09:03, 35:18, 35:31-35:32:02, 35:37, 35:53N, 35:64:01-35:64:02, 35:68:01-35:68:02, 35:75, 35:88, 35:99, 35:118-35:119, 35:127, 35:151, 35:174, 35:205, 35:234-35:235, 35:273 | B*57:01:01-57:01:04, 57:01:06-57:03:02, 57:05-57:06, 57:08, 57:10, 57:15-57:20, 57:22-57:23, 57:25-57:30, 57:32-57:44, 57:46, 57:48-57:50, 57:52-57:58, 57:60-57:68, 57:70, 57:72-57:73, B*58:36 |
| B*41:09, 41:29, 41:32, B*45:02-45:03 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

DR low resolution Interpretation Table

1HLA-A, HLA-B, HLA-DRB and HLA-DQB1 alleles listed on the IMGT/HLA web page 2015-January-19, release 3.19.0, www.ebi.ac.uk/imgt/hla.

The DRB4*02:01N and DRB4*03:01N null alleles will not be amplified by the DR low resolution primer set.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all DRB alleles is not known. The grouping of not serologically defined alleles is taken from Tissue Antigens 73, 95-170, 2009.

4The HLA-DRB1, -DRB3, -DRB4 and –DRB5  alleles will be grouped into their corresponding serological specificities, except that following alleles give rise to identical amplification patterns:

Alleles

DRB1*03:76, DRB1*13:176

DRB1*08:20, DRB1*13:18, 13:47, 13:55, 13:158, 13:164

DRB1*08:31, 08:41, DRB1*11:67

DRB1*03:11:01, DRB1*13:02:02

DRB1*13:13, 13:119, 13:154, 13:156, DRB1*14:84, 14:116, 14:144

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

DQ low resolution Interpretation Table

1HLA-A, HLA-B, HLA-DRB and HLA-DQB1 alleles listed on the IMGT/HLA web page 2015-January-19, release 3.19.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological split of the DQB1*05:05-05:89, DQB1*06:06 to 06:07 alleles, the DQB1*06:10, 06:13, 06:15-06:24 and 06:27 to 06:167, the DQB1*02:04-02:53Q, the DQB1*03:07-03:09 and 03:11-03:167 alleles and the DQB1:04:0301-04:27 alleles is not known. In this table we have used the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170, and also inferred the serological grouping from the naming of the sequence-defined allele.

Change in revision R01 compared to R00:

- Primer mix 69 does not amplify the DRB1*03:11:01, *03:97, *11:53 and *15:100 alleles. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- Due to sharing of sequence motifs in codon 38 and 47, DRB3*01:23 will be amplified in primer mix 69, in addition to primer mix 85.

- The DRB1*13:02:02 allele is amplified in primer mix 70.

The Specificity and Interpretation Tables have been changed.

Change in revision R03 compared to R02:

- Primer mix 20 may generate a false positive band of about 560 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings. A footnote has been added in the Specificity Table.

Change in revision R04 compared to R03:

- Primer mix 70 does not amplify the DRB1*14:137N and 14:152N alleles. Primer mix 79 does not amplify the DRB1*14:137N allele. This has been corrected in the Specificity and Interpretation Tables.