PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DPA1

Prod. No:	101.331-24/24u

101.331-06/06u

Lot No:    	73G

Expiry Date:    	2012-January-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for DPA1 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DPA13 alleles |
| --- | --- | --- | --- |
| 14,5 | 85 bp | 515 bp | *010301, 010302, 010304-0105, 0107-0110, 0401 |
| 2 | 255 bp | 515 bp | *010301-0104, 010601-0110 |
| 3 | 205 bp | 430 bp | *010301-010304, 010601-0107, 0109, 0110, 0301, 0302 |
| 44,6 | 115 bp | 430 bp | *0104, 0108, 0303 |
| 54 | 105 bp | 430 bp | *0105, 020101-0204, 0401 |
| 68 | 155 bp, 195 bp | 515 bp | *010601, 010602, 0110, 0204 |
| 74 | 100 bp | 430 bp | *010601, 010602, 020101-020106 |
| 84 | 100 bp | 430 bp | *020201-020203, 0204 |
| 9 | 205 bp | 430 bp | *020201-020203, 0204, 0302 |
| 104 | 85 bp | 515 bp | *010303, 0203, 0301-0303 |
| 114,7 | 90 bp | 515 bp | *0301, 0303 |
| 12 | 205 bp | 430 bp | *0401 |
| 13 | 135 bp | 430 bp | *0107 |
| 14 | 140 bp | 430 bp | *0108 |
| 15 | 245 bp | 430 bp | *010303, 0301-0303 |
| 16 | 220 bp | 430 bp | *0109 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DPA1 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DPA1 typing.

In addition, wells number 2, 6, 10 and 11 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For some DPA1 alleles only partial 2nd exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences are conserved within allelic groups.

4Specific PCR fragments shorter than 125 bp are less intense and less sharp compared to longer specific PCR fragments.

5Primer mix 1 may give rise to primer dimer formation.

6Primer mix 4 may faintly amplify the DPA1*0401 allele.

7Primer mix 11 may give rise to non-specific amplification.

8Primer mix 6: Specific PCR fragment of 155 bp in the DPA1*0110 and 0204 alleles. Specific PCR fragment of 195 bp in the DPA1*010601 and DPA1*010602 alleles.

1The sequence of the DPA1*0101 allele has been shown to be identical to DPA1*0103.

The sequence of the DPA1*0102 allele has been shown to be identical to DPA1*0103.

2Primer mix 6: Specific PCR fragment of 155 bp in the DPA1*0110 and 0204 alleles. Specific PCR fragment of 195 bp in the DPA1*010601 and DPA1*010602 alleles.