PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DPB1-DQ-DR Combi Tray

Prod. No:	101.709-24/24u

101.709-06/06u

Lot No:    	12N

Expiry Date:    	2014-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 48 primer mixes used for DPB1 SSP typing

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DPB1 alleles |
| --- | --- | --- | --- |
| 1 | 245 bp | 515 bp | *01:01:01-01:01:03, 03:01:01-03:01:02, 13:01, 25:01, 26:01:01-26:01:02, 29:01, 37:01, 44:01, 50:01, 52:01, 56:01, 61:01N, 78:01, 79:01, 88:01, 90:01, 92:01, 104:01, 107:01, 110:01, 111:01, 114:01, 118:01, 122:01, 124:01, 127:01, 132:01, 133:01 |
| 2 | 220 bp | 515 bp | *01:01:01-01:01:03, 03:01:01-03:01:02, 18:01, 20:01:01-20:01:02, 25:01, 26:01:01-26:01:02, 27:01, 36:01, 50:01, 52:01, 56:01, 61:01N, 78:01, 79:01, 85:01, 87:01, 89:01, 90:01, 92:01, 104:01, 111:01, 114:01, 118:01, 122:01, 124:01, 127:01, 130:01, 132:01 |
| 33,4,5 | 100 bp | 430 bp | *01:01:01-01:01:03, 04:01:01:01-04:01:02, 11:01:01-11:01:02, 13:01, 15:01, 26:01:01-26:01:02, 27:01, 31:01, 33:01, 39:01, 40:01, 56:01, 65:01, 66:01, 72:01, 74:01, 85:01, 89:01, 90:01, 96:01, 99:01, 102:01, 107:01, 110:01, 117:01, 118:01, 120:01N, 121:01, 125:01, 126:01, 127:01, 128:01, 133:01, 134:01 |
| 4 | 140 bp | 430 bp | *01:01:01-01:01:03, 04:01:01:01-04:01:02, 24:01, 26:01:01-26:01:02, 27:01, 28:01, 39:01, 40:01, 49:01, 51:01, 53:01, 56:01, 65:01, 66:01, 72:01, 76:01, 85:01, 89:01, 90:01, 92:01, 96:01, 99:01, 102:01, 118:01, 120:01N, 121:01, 122:01, 125:01, 126:01, 127:01, 128:01, 134:01 |
| 5 | 235 bp | 515 bp | *01:01:01-01:01:03, 08:01, 19:01, 50:01, 57:01, 65:01, 68:01, 73:01, 75:01, 84:01, 90:01, 106:01, 118:01, 122:01, 127:01 |
| 6 | 165 bp | 430 bp | *01:01:01-01:01:03, 13:01, 26:01:01-26:01:02, 65:01, 107:01, 118:01, 122:01, 127:01, 133:01 |
| 7 | 160 bp | 430 bp | *01:01:01-01:01:03, 13:01, 26:01:01-26:01:02, 56:01, 65:01, 76:01, 90:01, 92:01, 107:01, 110:01, 118:01, 122:01, 125:01, 127:01, 133:01 |
| 8 | 210 bp | 515 bp | *02:01:02-02:02, 04:01:01:01-04:02:01:02, 05:01:01-05:01:02, 08:01, 16:01, 19:01, 22:01, 23:01, 24:01, 32:01, 33:01, 38:01, 39:01, 40:01, 46:01, 47:01, 48:01, 49:01, 51:01, 53:01, 60:01, 62:01, 63:01, 65:01, 68:01, 71:01, 75:01, 77:01, 80:01, 81:01, 82:01, 84:01, 94:01, 96:01, 97:01, 99:01, 100:01, 101:01, 103:01, 105:01, 106:01, 112:01, 115:01, 116:01, 120:01N, 121:01, 123:01, 126:01, 128:01, 134:01 |
| 9 | 165 bp | 430 bp | *02:01:02-02:01:05, 02:01:06w, 02:01:07-02:02, 04:02:01:01-04:02:01:02, 05:01:01-05:01:02, 06:01, 16:01, 17:01, 18:01, 19:01, 20:01:01-20:01:02, 21:01, 22:01, 23:01, 30:01, 32:01, 34:01, 36:01, 38:01, 41:01:01-41:01:02, 46:01, 47:01, 48:01, 55:01, 58:01, 59:01, 60:01, 62:01, 63:01, 64:01N, 69:01, 71:01, 77:01, 80:01, 82:01, 83:01, 86:01, 87:01, 91:01, 93:01, 94:01, 95:01, 97:01, 98:01, 100:01, 101:01, 103:01, 105:01, 106:01, 108:01, 109:01, 112:01, 113:01, 115:01, 116:01, 123:01, 129:01, 130:01, 131:01 |
| 10 | 220 bp | 515 bp | *02:01:02-02:02, 08:01, 16:01, 19:01, 22:01, 32:01, 33:01, 41:01:01-41:01:02, 46:01, 47:01, 48:01, 71:01, 81:01, 95:01, 101:01, 103:01, 106:01, 115:01, 123:01 |
| 11 | 185 bp | 515 bp | *02:01:02-02:01:07, 04:01:01:01-04:02:01:02, 23:01, 24:01, 32:01, 33:01, 41:01:01-41:01:02, 46:01, 47:01, 51:01, 59:01, 60:01, 66:01, 71:01, 72:01, 73:01, 75:01, 77:01, 80:01, 81:01, 82:01, 83:01, 86:01, 94:01, 99:01, 105:01, 108:01, 109:01, 113:01, 115:01, 119:01, 120:01N, 121:01, 123:01, 126:01, 128:01, 129:01, 134:01 |
| 123 | 80 bp | 515 bp | *02:01:02, 02:01:03w, 02:01:04-02:01:07, 06:01, 08:01, 09:01, 10:01, 16:01, 17:01, 29:01, 32:01, 37:01, 41:01:01-41:01:02, 44:01, 46:01, 48:01, 64:01N, 81:01, 86:01, 88:01, 93:01, 109:01, 113:01, 115:01, 123:01, 131:01 |
| 13 | 180 bp | 430 bp | *02:02, 05:01:01-05:01:02, 19:01, 22:01, 24:01, 38:01, 47:01, 97:01, 100:01, 101:01, 106:01 |
| 143 | 105 bp, 175 bp | 430 bp | *02:02, 05:01:01-05:01:02, 21:01, 22:01, 30:01, 36:01, 38:01, 54:01, 97:01, 100:01, 101:01, 114:01 |
| 15 | 200 bp | 430 bp | *02:02, 48:01, 95:01, 99:01, 100:01, 101:01, 112:01, 116:01 |
| 163 | 80 bp, 165 bp, 210 bp | 430 bp | *03:01:01-03:01:02, 04:02:01:01-04:02:01:02, 14:01, 18:01, 20:01:01-20:01:02, 25:01, 28:01, 35:01:01-35:01:02, 45:01, 49:01, 50:01, 51:01, 53:01, 57:01, 59:01, 60:01, 61:01N, 68:01, 70:01, 73:01, 75:01, 76:01, 77:01, 78:01, 79:01, 80:01, 82:01, 83:01, 91:01, 92:01, 98:01, 99:01, 104:01, 105:01, 108:01, 111:01, 116:01, 119:01, 122:01, 124:01, 127:01, 129:01, 130:01, 132:01 |
| 17 | 260 bp | 515 bp | *03:01:01-03:01:02, 06:01, 09:01, 10:01, 11:01:01-11:01:02, 13:01, 14:01, 17:01, 20:01:01-20:01:02, 21:01, 25:01, 26:01:01-26:01:02, 27:01, 29:01, 30:01, 35:01:01-35:01:02, 36:01, 37:01, 44:01, 45:01, 52:01, 54:01, 55:01, 56:01, 58:01, 61:01N, 64:01N, 67:01, 69:01, 70:01, 76:01, 78:01, 79:01, 85:01, 87:01, 88:01, 91:01, 92:01, 93:01, 98:01, 104:01, 107:01, 110:01, 111:01, 114:01, 119:01, 124:01, 125:01, 130:01, 131:01, 132:01, 133:01 |
| 18 | 180 bp | 430 bp | *03:01:01-03:01:02, 06:01, 09:01, 14:01, 17:01, 20:01:01-20:01:02, 29:01, 35:01:01-35:01:02, 44:01, 61:01N, 64:01N, 69:01, 70:01, 76:01, 78:01, 86:01, 88:01, 91:01, 92:01, 98:01, 104:01, 111:01, 119:01, 124:01, 130:01, 131:01, 132:01 |
| 19 | 200 bp | 430 bp | *03:01:01-03:01:02, 06:01, 11:01:01-11:01:02, 14:01, 20:01:01-20:01:02, 25:01, 29:01, 44:01, 45:01, 52:01, 56:01, 61:01N, 64:01N, 67:01, 69:01, 70:01, 74:01, 76:01, 78:01, 87:01, 91:01, 92:01, 104:01, 111:01, 114:01, 119:01, 124:01, 130:01 |
| 20 | 235 bp | 515 bp | *03:01:01-03:01:02, 09:01, 10:01, 13:01, 14:01, 25:01, 26:01:01-26:01:02, 29:01, 35:01:01-35:01:02, 37:01, 44:01, 45:01, 52:01, 54:01, 56:01, 61:01N, 67:01, 70:01, 76:01, 78:01, 79:01, 88:01, 92:01, 104:01, 107:01, 110:01, 111:01, 114:01, 119:01, 124:01, 125:01, 132:01, 133:01 |
| 213 | 100 bp | 430 bp | *03:01:01-03:01:02, 09:01, 14:01, 29:01, 35:01:01-35:01:02, 44:01, 50:01, 57:01, 61:01N, 70:01, 76:01, 78:01, 88:01, 92:01, 104:01, 111:01, 119:01, 124:01, 132:01 |
| 22 | 185 bp | 430 bp | *04:01:01:01-04:01:02, 24:01, 33:01, 39:01, 49:01, 51:01, 66:01, 72:01, 81:01, 96:01, 99:01, 117:01, 120:01N, 121:01, 126:01, 128:01, 134:01 |
| 23 | 145 bp, 210 bp | 430 bp | *05:01:01-05:01:02, 34:01, 36:01, 38:01, 62:01, 63:01, 97:01, 100:01, 112:01, 114:01, 116:01, 127:01 |
| 24 | 195 bp, 265 bp | 515 bp | *05:01:01-05:01:02, 08:01, 16:01, 19:01, 22:01, 31:01, 38:01, 57:01, 63:01, 65:01, 68:01, 84:01, 97:01, 101:01, 102:01, 103:01, 106:01 |
| 25 | 190 bp, 265 bp | 515 bp | *05:01:01-05:01:02, 21:01, 22:01, 36:01, 38:01, 44:01, 58:01, 63:01, 70:01, 97:01, 114:01 |
| 263,5 | 75 bp, 205 bp | 515 bp | *06:01, 09:01, 17:01, 29:01, 44:01, 46:01, 64:01N, 86:01, 88:01, 101:01, 123:01, 131:01 |
| 273 | 95 bp | 430 bp | *06:01, 17:01, 20:01:01-20:01:02, 46:01, 64:01N, 69:01, 80:01, 86:01, 91:01, 98:01, 108:01, 130:01, 131:01 |
| 28 | 200 bp, 260 bp | 430 bp | *09:01, 10:01, 13:01, 17:01, 21:01, 26:01:01-26:01:02, 27:01, 30:01, 35:01:01-35:01:02, 36:01, 37:01, 54:01, 55:01, 58:01, 66:01, 79:01, 85:01, 86:01, 88:01, 93:01, 98:01, 107:01, 109:01, 110:01, 113:01, 117:01, 119:01, 125:01, 129:01, 131:01, 132:01, 133:01 |
| 29 | 220 bp | 515 bp | *09:01, 10:01, 17:01, 30:01, 54:01, 55:01, 58:01, 86:01, 113:01, 131:01 |
| 30 | 185 bp | 430 bp | *09:01, 14:01, 17:01, 35:01:01, 76:01, 86:01, 91:01, 98:01, 119:01, 131:01 |
| 313,6 | 120 bp | 430 bp | *01:01:01-01:01:03, 05:01:01-05:01:02, 08:01, 10:01, 11:01:01-11:01:02, 13:01, 16:01, 19:01, 21:01, 22:01, 25:01, 26:01:01-26:01:02, 27:01, 30:01, 31:01, 36:01, 37:01, 38:01, 45:01, 52:01, 54:01, 55:01, 56:01, 58:01, 63:01, 65:01, 67:01, 68:01, 79:01, 84:01, 85:01, 87:01, 89:01, 90:01, 93:01, 97:01, 102:01, 103:01, 106:01, 107:01, 110:01, 114:01, 118:01, 122:01, 125:01, 127:01, 133:01 |
| 323,6 | 75 bp, 200 bp | 430 bp | *11:01:01-11:01:02, 78:01, 130:01 |
| 33 | 140 bp, 205 bp | 430 bp | *13:01, 33:01, 81:01, 107:01, 110:01, 117:01, 123:01, 133:01 |
| 34 | 210 bp | 430 bp | *14:01, 45:01, 67:01, 76:01, 91:01, 119:01 |
| 35 | 195 bp | 430 bp | *15:01, 74:01 |
| 36 | 260 bp | 515 bp | *15:01, 18:01, 28:01, 34:01, 40:01, 53:01, 62:01 |
| 377 | 185 bp, 255 bp | 430 bp | *04:02:01:01-04:02:01:02, 08:01?, 15:01, 21:01?, 22:01?, 24:01?, 25:01?, 27:01?, 28:01, 29:01?, 31:01?, 32:01?, 34:01?, 35:01:01?, 36:01?, 37:01?, 38:01?, 40:01, 41:01:01?-41:01:02?, 44:01?, 46:01?, 47:01?, 48:01?, 49:01?, 51:01?, 52:01?, 53:01, 54:01?, 56:01?, 57:01?, 58:01?, 59:01, 60:01?, 61:01N?, 62:01?, 63:01?, 64:01N?, 65:01?, 66:01?, 67:01?, 68:01?, 69:01?, 70:01?, 71:01?, 72:01?, 73:01?, 74:01, 75:01?, 76:01?, 77:01?, 78:01?, 79:01?, 80:01?, 81:01?, 82:01?, 83:01?, 84:01?, 86:01?, 87:01?, 89:01?, 90:01?, 91:01?, 92:01?, 93:01?, 94:01?, 95:01?, 96:01?, 97:01?, 98:01?, 99:01?, 100:01?, 101:01?, 102:01?, 103:01?, 108:01?, 109:01?, 110:01?, 111:01?, 112:01?, 113:01?, 114:01?, 115:01?, 116:01?, 117:01?, 118:01?, 119:01?, 120:01N?, 122:01?, 125:01?, 126:01, 127:01?, 128:01?, 129:01?, 130:01? |
| 383 | 100 bp | 515 bp | *23:01, 34:01, 52:01, 55:01, 58:01, 62:01, 63:01, 67:01, 71:01, 87:01, 95:01, 103:01, 112:01, 121:01 |
| 39 | 210 bp | 515 bp | *28:01, 31:01, 34:01, 57:01, 59:01, 72:01, 73:01, 95:01, 102:01, 108:01 |
| 403,4,7 | 95 bp, 180 bp | 430 bp | *06:01?, 08:01?, 10:01?, 11:01:01?-11:01:02?, 16:01?, 17:01?, 18:01?, 20:01:01?-20:01:02?, 21:01?, 22:01?, 23:01?, 24:01?, 25:01?, 26:01:01?-26:01:02?, 27:01?, 28:01?, 29:01?, 30:01, 31:01?, 32:01?, 33:01?, 34:01?, 35:01:01?-35:01:02?, 36:01?, 37:01?, 38:01?, 39:01?, 40:01?, 41:01:01?-41:01:02?, 44:01?, 45:01?, 46:01?, 47:01?, 48:01?, 49:01?, 50:01?, 51:01?, 52:01?, 53:01?, 54:01, 55:01?, 56:01?, 57:01?, 58:01?, 59:01?, 60:01?, 61:01N?, 62:01?, 63:01?, 64:01N?, 65:01?, 66:01?, 67:01?, 68:01?, 69:01?, 70:01?, 71:01?, 72:01?, 73:01?, 74:01?, 75:01?, 76:01?, 77:01?, 78:01?, 79:01?, 80:01?, 81:01?, 82:01?, 83:01?, 84:01?, 85:01?, 86:01?, 87:01?, 88:01?, 89:01?, 90:01?, 91:01?, 92:01?, 93:01?, 94:01?, 95:01?, 96:01?, 97:01?, 98:01?, 99:01?, 100:01?, 101:01?, 102:01?, 103:01?, 107:01, 108:01?, 109:01?, 110:01?, 111:01?, 112:01?, 113:01?, 114:01?, 115:01?, 116:01?, 117:01?, 118:01?, 119:01?, 120:01N?, 121:01?, 122:01?, 123:01?, 124:01?, 125:01?, 126:01?, 127:01?, 128:01?, 129:01?, 130:01?, 131:01?, 132:01?, 133:01?, 134:01? |
| 413 | 75 bp, 225 bp, 250 bp | 515 bp | *31:01, 34:01, 77:01, 94:01, 134:01 |
| 42 | 140 bp, 185 bp, 245 bp, 285 bp | 430 bp | *32:01, 82:01, 97:01, 128:01 |
| 433,4,7 | 65 bp, 190 bp | 430 bp | *05:01:01-05:01:02, 06:01?, 08:01?, 10:01?, 11:01:01?-11:01:02?, 16:01?, 17:01?, 18:01?, 19:01, 20:01:01?-20:01:02?, 21:01?, 22:01?, 23:01?, 24:01?, 25:01?, 26:01:01?-26:01:02?, 27:01?, 28:01?, 29:01?, 30:01?, 31:01?, 32:01?, 33:01?, 34:01, 35:01:01?-35:01:02?, 36:01?, 37:01?, 38:01?, 39:01?, 40:01?, 41:01:01?-41:01:02?, 44:01?, 45:01?, 46:01?, 47:01?, 48:01?, 49:01?, 50:01?, 51:01?, 52:01?, 53:01?, 54:01?, 55:01?, 56:01?, 57:01?, 58:01?, 59:01?, 60:01?, 61:01N?, 62:01, 63:01?, 64:01N?, 65:01?, 66:01?, 67:01?, 68:01?, 69:01?, 70:01?, 71:01?, 72:01?, 73:01?, 74:01?, 75:01?, 76:01?, 77:01?, 78:01?, 79:01?, 80:01?, 81:01?, 82:01?, 83:01?, 84:01?, 85:01?, 86:01?, 87:01?, 88:01?, 89:01?, 90:01?, 91:01?, 92:01?, 93:01?, 94:01?, 95:01?, 96:01?, 97:01?, 98:01?, 99:01?, 100:01?, 101:01?, 102:01?, 103:01?, 104:01, 108:01?, 109:01?, 110:01?, 111:01?, 112:01?, 113:01?, 114:01?, 115:01?, 116:01?, 117:01?, 118:01?, 119:01?, 120:01N?, 121:01?, 122:01?, 123:01?, 124:01?, 125:01?, 126:01?, 127:01?, 128:01?, 129:01?, 130:01?, 131:01?, 132:01?, 133:01?, 134:01? |
| 443 | 110 bp, 185 bp, 245 bp, 270 bp | 430 bp | *38:01, 64:01N, 111:01, 118:01, 120:01N |
| 453 | 120 bp, 190 bp, 245 bp | 430 bp | *39:01, 46:01, 49:01, 80:01, 86:01, 96:01, 108:01, 115:01, 117:01 |
| 463 | 70 bp, 205 bp | 430 bp | *41:01:01-41:01:02, 60:01, 69:01, 83:01, 96:01 |
| 474 | 165 bp, 255 bp | 430 bp | *02:01:02-02:02, 04:01:01:01-04:01:02, 08:01?, 17:01, 21:01?, 22:01?, 23:01, 24:01?, 25:01?, 26:01:01?, 27:01?, 28:01, 29:01?, 30:01, 31:01?, 32:01?, 33:01, 34:01?, 35:01:01?, 36:01?, 37:01?, 38:01?, 39:01, 40:01?, 41:01:01?-41:01:02?, 44:01, 46:01?, 47:01?, 48:01?, 49:01?, 51:01?, 52:01?, 53:01?, 54:01?, 55:01, 56:01?, 57:01?, 58:01?, 60:01?, 61:01N?, 62:01?, 63:01?, 64:01N?, 65:01?, 66:01?, 67:01?, 68:01?, 69:01?, 70:01?, 71:01?, 72:01?, 73:01?, 74:01?, 75:01?, 76:01?, 77:01?, 78:01?, 79:01?, 80:01?, 81:01?, 82:01?, 83:01?, 84:01?, 86:01?, 87:01?, 89:01?, 90:01?, 91:01?, 92:01?, 93:01?, 94:01?, 95:01?, 96:01?, 97:01?, 98:01?, 99:01?, 100:01?, 101:01?, 102:01?, 103:01?, 105:01, 106:01, 108:01?, 109:01?, 110:01?, 111:01?, 112:01?, 113:01?, 114:01, 115:01?, 116:01?, 117:01?, 118:01?, 119:01?, 120:01N?, 121:01, 122:01?, 123:01, 124:01, 125:01?, 127:01?, 128:01?, 129:01?, 130:01?, 133:01, 134:01 |
| 483 | 115 bp, 210 bp, 280 bp | 430 bp | *56:01, 61:01N, 66:01, 76:01, 85:01, 92:01, 110:01, 125:01 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DPB1 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DPB1 typing.

In addition, wells number 2, 5, 8, 10 to 12, 17, 20, 24 to 26, 29, 36, 38, 39 and 41 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Specific PCR fragments shorter than 125 base pairs are less intense and not as sharp as longer specific bands.

4Primer mixes 3, 40, 43 and 47 may yield less specific PCR products than the other DPB1 primer mixes.

5Primer mixes 3 and 26 may give rise to primer dimer formation.

6Primer mixes 31 and 32 may give rise to non-specific amplifications.

7The 1st, 3rd and 4th exon of many DPB1 alleles is not sequenced and thus the primer pairs in primer mixes 37, 40, 43 and 47 may amplify other DPB1 alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the DPB1*02011 allele has been shown to be identical to DPB1*02:01:02.

DPB1*0701 has never been assigned.

DPB1*1201 has never been assigned.

The sequence of the DPB1*4201 allele has been shown to be identical to DPB1*31:01.

The sequence of the DPB1*4301 allele has been shown to be identical to DPB1*28:01.

2The 1st, 3rd and 4th exon of many DPB1 alleles is not sequenced and thus the primer pairs in primer mixes 37, 40 and 43 may amplify other DPB1 alleles. We assume that unknown sequences in these exons and in the introns are conserved within loci and within allelic groups.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Specificities and sizes of the PCR products of the 24 primer mixes of the DQ low/DQB1*03 medium to high resolution primer set

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DQ serology3 | Amplified DQB1 alleles4 |
| --- | --- | --- | --- | --- |
| 49 | 225 bp | 515 bp | 5 | *05:01:01:01-05:12 |
| 50 | 220 bp, 270 bp | 430 bp | 1, 5, 6 | *06:01:01-06:44 |
| 51 | 210 bp | 430 bp | 2 | *02:01:01-02:06 |
| 52 | 220 bp | 515 bp | 3, 7 | *03:01:01:01-03:01:06, 03:04, 03:09-03:10, 03:13-03:14, 03:16, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36 |
| 535 | 130 bp | 515 bp | 6, 8 | *03:02:01-03:02:05, 03:05:01-03:05:04, 03:07-03:08, 03:11, 03:18, 03:32, 03:37, 06:29 |
| 547 | 135 bp | 515 bp | 2, 3, 9 | *02:03, 03:03:02:01-03:03:04, 03:06, 03:12, 03:15, 03:20, 03:25-03:26, 03:30-03:31, 03:33-03:34, 03:38, 04:03:01-04:03:02 |
| 55 | 145 bp, 185 bp | 515 bp | 3, 7, 8, 9 | *03:01:01:01-03:38 |
| 56 | 210 bp, 245 bp | 430 bp | 4 | *04:01:01-04:08 |
| 57 | 165 bp | 515 bp | 7 | *03:01:01:01-03:01:06, 03:04, 03:09, 03:13, 03:16, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36 |
| 58 | 135 bp | 430 bp | 3, 7 | *03:04, 03:14 |
| 595 | 95 bp, 130 bp | 430 bp | 8, 9 | *03:05:01-03:05:04, 03:17, 03:20 |
| 605 | 115 bp | 430 bp | 3 | *03:06, 03:25 |
| 615 | 110 bp, 145 bp | 430 bp | 7, 8,9 | *03:07, 03:15-03:16 |
| 62 | 135 bp | 430 bp | 6, 8 | *03:08, 06:02:02, 06:03:02 |
| 63 | 135 bp, 260 bp | 430 bp | 7, 8 | *03:09, 03:11, 03:26 |
| 646 | 135 bp | 430 bp | 3, 6, 9 | *03:10, 03:12, 03:14, 06:01:01-06:01:06, 06:43 |
| 65 | 260 bp | 515 bp | 3, 7, 9 | *03:01:01:01-03:01:06, 03:04, 03:09-03:10, 03:12-03:14, 03:16, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36 |
| 66 | 160 bp | 430 bp | 7 | *03:13, 03:21-03:22 |
| 675 | 130 bp | 515 bp | 3, 7 | *03:01:01:01-03:01:06, 03:04, 03:09-03:10, 03:13-03:14, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36 |
| 685 | 130 bp | 430 bp | 2, 3, 7, 9 | *02:03, 03:01:01:01-03:01:02, 03:01:04-03:01:06, 03:03:02:01-03:03:04, 03:06, 03:09-03:10, 03:12-03:13, 03:15-03:17, 03:19-03:24, 03:26-03:31, 03:33-03:36, 03:38 |
| 695,7 | 130 bp, 220 bp | 430 bp | 8 | *03:18, 03:24 |
| 708 | 175 bp | 515 bp | 3, 4, 7, 8, 9 | *03:02:01-03:03:04, 03:05:01-03:05:04, 03:06?-03:08?, 03:11?-03:18?, 03:19, 03:20?, 03:23?, 03:25, 03:26?, 03:30-03:32, 03:33?-03:34?, 03:37?, 03:38, 04:01:01-04:08 |
| 716,8 | 175 bp | 515 bp | 3, 7, 8, 9 | *03:01:01:01-03:01:06, 03:04, 03:06?-03:08?, 03:09-03:10, 03:11?-03:18?, 03:20?, 03:21-03:22, 03:23?, 03:24, 03:26?, 03:27-03:29, 03:33?-03:34?, 03:35-03:36, 03:37? |
| 72 | 160 bp | 515 bp | 1, 6 | *03:23, 06:03:01-06:03:02, 06:04:02, 06:07:01, 06:08:01, 06:11:01-06:11:02, 06:26N, 06:28, 06:30-06:32, 06:40-06:41, 06:44 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQ low resolution SSP subtypings. When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 49 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DQ low resolution typing. In addition, wells number 52 to 55, 57, 65, 67 and 70 to 72 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DQB alleles is not known. In this table we use the expert-assigned grouping in Tissue Antigens (2009) 73:95-170 and have also inferred the serological grouping from the naming of the sequence-defined allele.

4For several DQB1 alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 5’- and 3’-ends of the second exon of the DQB1 gene are conserved within allelic groups.

5Short PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mix 64 and 71 may give rise to nonspecific amplifications.

7Primer mix 54 and 69 may give rise to primer oligomer formation.

8For several DQB1 alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 5’- and 3’-ends of the second exon of the DQB1 gene are conserved within allelic groups.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the DQB1*03031 allele has been shown to be identical to DQB1*03:03:02.

2The serological reactivity of all DQB alleles is not known. In this table we use the information the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and have also inferred the serological grouping from the naming of the sequence-defined allele.

Specificities and sizes of the PCR products of the 23+1 primer mixes of the DR low resolution primer set

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DR serology3 | Amplified HLA-DRB4 alleles |
| --- | --- | --- | --- | --- |
| 736,7 | 205 bp, 255 bp | 515 bp | 1 | *01:01:01-01:02:05, 01:04-01:37 |
| 74 | 200 bp | 430 bp | 103/1 | *01:03 |
| 75 | 200 bp, 215 bp | 430 bp | 2, 15 | *15:01:01:01-15:58 |
| 76 | 210 bp | 430 bp | 16 | *16:01:01-16:05:02, 16:07-16:18 |
| 775, 6,10 | 120 bp, 220 bp | 430 bp | 3, 11, 17, 18 | *03:01:01:01-03:65, 11:07, 11:53, 11:103, 11:105, 11:107, 15:25 |
| 785,6,10 | 80 bp, 210 bp | 430 bp | 3, 6, 11, 13, 14, 17 | *03:01:01:01-03:01:14, 03:04:01-03:06, 03:08-03:16, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:31, 03:33-03:34, 03:36-03:37, 03:43-03:48, 03:50-03:52, 03:54-03:65, 08:40, 11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-11:41, 11:48, 11:59, 11:63, 11:65:01-11:65:02, 11:68, 11:70, 11:73, 11:76, 11:79-11:80, 11:83, 11:85-11:87, 11:93, 13:01:01-13:04, 13:08, 13:10, 13:15-13:17, 13:19-13:20, 13:22-13:24, 13:27-13:29, 13:31-13:41, 13:43, 13:45, 13:48, 13:51-13:54, 13:57, 13:59, 13:61:01-13:61:02, 13:63-13:66:02, 13:68-13:76, 13:78-13:81, 13:83-13:85, 13:87-13:99, 13:101-13:102, 13:104-13:107, 13:109, 13:111-13:117, 14:16, 14:19, 14:21, 14:82, 14:95, 14:109 |
| 795, 6 | 85 bp, 210 bp | 430 bp | 3, 6, 18, 11, 13, 14, 1403 | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 11:13:01w-11:13:02w, 11:26, 11:34, 13:15, 13:19, 13:26, 13:44, 13:53, 13:57, 13:85-13:86, 13:104, 14:02-14:03:02, 14:06:01-14:06:02, 14:09, 14:12:01-14:13, 14:17-14:21, 14:24, 14:27, 14:29-14:30, 14:32:01w-14:32:02w, 14:33, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:65w, 14:67, 14:77-14:78, 14:80-14:81, 14:83, 14:85, 14:89, 14:94, 14:98, 14:102, 14:106, 14:108-14:109, 14:115 |
| 805, 6 | 100 bp, 175 bp | 430 bp | 4, 3/4 | *04:01:01-04:102 |
| 81 | 210 bp, 235 bp | 430 bp | 7, 13, 14 | *07:01:01:01-07:01:04, 07:03-07:21, 12:22, 13:17, 13:116, 14:50 |
| 826 | 170 bp, 215 bp, 250 bp | 515 bp | 8, 11, 12, 14 | *08:01:01-08:19, 08:21-08:48, 11:67, 12:04, 12:16, 12:22, 14:11, 14:15, 14:68, 14:93 |
| 835,6 | 90 bp, 135 bp, 180 bp | 430 bp | 3, 9, 11 | *03:08, 03:65, 09:01:02-09:16, 11:07, 11:53, 11:103, 11:105, 11:107 |
| 847 | 205 bp | 430 bp | 10 | *10:01:01-10:03 |
| 855,6 | 100 bp, 170 bp | 430 bp | 3, 8, 11, 14 | *03:08, 03:65, 08:31, 08:41, 11:01:01-11:70, 11:72-11:113 |
| 865,6 | 85 bp, 105 bp | 430 bp | 12 | *08:32, 12:01:01-12:34 |
| 87 | 215 bp | 430 bp | 6, 8, 11, 13, 14, 1403 | *08:20-08:21, 11:01:01-11:04:08, 11:06:01-11:06:02, 11:08:01-11:12:02, 11:14:01-11:16, 11:18-11:21, 11:23-11:25, 11:27:01-11:33, 11:35-11:51, 11:54:01-11:54:02, 11:56-11:66, 11:68, 11:70, 11:72-11:81, 11:83-11:88, 11:90-11:97, 11:99-11:102:01, 11:106, 11:108-11:113, 13:01:01-13:08, 13:10-13:16, 13:18-13:43, 13:45-13:85, 13:87-13:115, 13:117-13:119, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:16, 14:19, 14:21-14:22, 14:25, 14:27, 14:40, 14:53, 14:63, 14:67, 14:69, 14:74, 14:77-14:78, 14:84-14:85, 14:98, 14:102, 14:105, 14:109, 14:115-14:116, DRB3*02:27 |
| 886 | 195 bp, 210 bp | 430 bp | 6, 8, 11, 12, 13, 14 | *08:01:01-08:02:04, 08:04:01-08:09, 08:11, 08:16-08:17, 08:20-08:22, 08:24, 08:26, 08:28, 08:31, 08:39, 08:41-08:44, 11:01:01-11:06:02, 11:09-11:12:02, 11:14:01-11:16, 11:20-11:21, 11:23-11:25, 11:27:01-11:30, 11:32-11:33, 11:35-11:41, 11:43-11:44, 11:46:01-11:51, 11:54:01-11:56, 11:58:01-11:63, 11:65:01-11:70, 11:72, 11:74:01-11:78, 11:80-11:88, 11:90-11:97, 11:99-11:102:01, 11:106, 11:108-11:113, 12:02:01-12:02:05, 12:13, 12:15-12:16, 12:18-12:21, 12:23, 12:26-12:27, 12:31N-12:33, 13:01:01-13:02:01, 13:02:03-13:02:05, 13:04-13:05:02, 13:07:01-13:09, 13:11:01-13:11:02, 13:14:01-13:24, 13:26-13:29, 13:31-13:32, 13:34-13:36, 13:38-13:43, 13:45-13:55, 13:57, 13:59, 13:61:01-13:65, 13:67-13:76, 13:78-13:80, 13:83-13:84, 13:87, 13:91-13:93, 13:96:01-13:100, 13:102-13:109, 13:111-13:114, 13:116-13:117, 14:15-14:16, 14:22, 14:24-14:25, 14:27, 14:37, 14:53, 14:73, 14:105 |
| 8910 | 175 bp | 430 bp | 3, 6, 11, 13, 14, 1403, 17, 18 | *03:01:01:01-03:07, 03:09, 03:11:01-03:41, 03:43-03:45, 03:47-03:63, 08:20, 13:01:01-13:16, 13:18-13:42, 13:44, 13:46-13:66:02, 13:68-13:102, 13:104-13:115, 13:117-13:119, 14:02-14:03:02, 14:05:01-14:06:02, 14:09, 14:12:01-14:14, 14:17-14:21, 14:23:01, 14:23:03-14:24, 14:27, 14:29-14:30, 14:33, 14:36-14:37, 14:40-14:45, 14:47-14:48, 14:51, 14:56, 14:59, 14:63-14:65, 14:67, 14:77-14:78, 14:80-14:81, 14:83-14:85, 14:89, 14:91, 14:94-14:96, 14:98, 14:100, 14:102-14:103, 14:106, 14:108-14:109, 14:115-14:116 |
| 905,9 | 100 bp, 140 bp, 155 bp | 430 bp | 4, 6, 8,13, 14, 1404 | *04:62, 04:69, 04:73, 08:08, 11:69, 11:82, 13:45, 14:01:01-14:01:02, 14:04, 14:07:01-14:07:02, 14:10, 14:16, 14:22, 14:25-14:26, 14:28, 14:31-14:32:02, 14:35, 14:37-14:39, 14:49-14:50, 14:53-14:55, 14:57-14:58, 14:60-14:62, 14:68-14:71, 14:73-14:76, 14:79, 14:82, 14:86-14:88, 14:90, 14:93, 14:99, 14:101, 14:104-14:105, 14:107, 14:110-14:114, 14:117, DRB4*01:03:01:02N |
| 915,6,8 | 110 bp, 135 bp, 170 bp | 430 bp | 3, 6, 9, 11, 13, 14, 1404 | *03:10, 09:01:02-09:01:05, 09:01:07-09:02:02, 09:04-09:16, 11:13:01-11:13:02, 11:17, 11:52, 13:43, 14:01:01-14:02, 14:04-14:11, 14:13-14:14, 14:16-14:18, 14:19w, 14:20, 14:21w, 14:22-14:23:03, 14:26, 14:28-14:36, 14:38-14:39, 14:41, 14:43-14:52, 14:54-14:57, 14:59-14:62, 14:64-14:65, 14:68, 14:70-14:76, 14:79-14:83, 14:86-14:88, 14:90-14:97, 14:99-14:101, 14:103-14:108, 14:109w, 14:110-14:114, 14:117, 15:27, 15:34 |
| 925,6,7 | 110 bp, 175 bp, 225 bp | 430 bp | 2, 3, 6, 8, 11, 13, 14,1403, 1404,16 | *03:10, 08:09, 08:20-08:21, 08:32, 08:35, 11:13:01-11:13:02, 11:17, 11:23, 11:25, 11:31, 11:45, 11:52, 11:55, 11:64, 11:89, 11:96, 13:13, 13:18, 13:43, 13:45, 13:47, 13:55, 13:119, 14:01:01-14:01:03, 14:03:01-14:05:03, 14:07:01-14:08, 14:10-14:12:02, 14:14-14:16, 14:18, 14:22-14:23:03, 14:25-14:28, 14:31-14:32:02, 14:34-14:36, 14:38-14:40, 14:42-14:45, 14:49-14:50, 14:53-14:65, 14:67-14:79, 14:81-14:82, 14:84-14:93, 14:95-14:97, 14:99-14:105, 14:107, 14:110-14:117, 15:21w, 16:04w, 16:18w |
| 936,10 | 160 bp, 240 bp | 430 bp | 52 | DRB3*01:01:02:01-01:15, DRB3*02:01-02:28, DRB3*03:01:01-03:03 |
| 947 | 215 bp | 430 bp | 53 | DRB4*01:01:01:01-01:08 |
| 95 | 175 bp | 430 bp | 51 | DRB5*01:01:01-01:14, DRB5*02:02-02:05 |
| 9611 | - | - |  | Negative Control |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 75, 80, 90, 91 and 92.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 73 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DR low resolution typing.

In addition, well number 82 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DRB alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170, and have also inferred the serological grouping from the naming of the sequence-defined allele.

4For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP set are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

The DRB1*08:09 and DRB1*14:15 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:20 and the DRB1*13:18, 13:47 and13:55 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31 and 08:41 and the DRB1*11:67 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

The DRB1*13:13 and 13:119 and the DRB1*14:84 and 14:116 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 73, 77 to 80, 82 to 83, 85, 86, 88 and 91 to 93.

7Primer mix 73, 84, 92 and 94 may give rise a primer oligomer formation.

8Primer mix 91 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

9The DRB4*01:03:01:02N allele is amplified by the primer pair in well No. 90, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer pair.

10Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 in addition to primer mix 93.

11Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the DRB1*0702 allele has been shown to be identical to DRB1*07:01:01:01.

The sequence of the DRB1*080301 allele has been shown to be identical to DRB1*08:03:02.

The sequence of the DRB1*090101 allele has been shown to be identical to DRB1*09:01:02.

The sequence of the DRB1*1171 allele has been shown to be identical to DRB1*11:02:01.

The sequence of the DRB1*120301 allele has been shown to be identical to DRB1*12:01:01.

The DRB1*1466 allele has been renamed DRB1*14:32:02.

The sequence of the DRB1*1606 allele has been shown to be identical to DRB1*16:05:01.

The sequence of the DRB3*010101 allele has been shown to be identical to DRB3*01:01:02:01.

The DRB4*0101102N allele has been renamed DRB4*01:03:10:02N.

The sequence of the DRB5*0201 allele has been shown to be identical to DRB5*02:02.

Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 77, 78 and 89 in addition to primer mix 93.

The DRB1*08:09 and DRB1*14:15 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:20 and the DRB1*13:18, 13:47 and13:55 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31 and 08:41 and the DRB1*11:67 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

The DRB1*13:13 and 13:119 and the DRB1*14:84 and 14:116 alleles yield identical amplification patterns in the DPB1-DQ-DR kit. These alleles can be separated by the respective high resolution primer sets.

2The serological reactivity of all DRB alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and have also inferred the serological grouping from the naming of the sequence-defined allele.

3Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘ser’, serological HLA specificity. ‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- The DPB1*31:01 allele is not amplified by primer mix 4. This has been corrected in the Specificity and Interpretation Tables.