PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DQ low resolution

Prod. No:	101.201-48/48u

101.201-12/12u

Lot No:    	91F

Expiry Date:    	2011-May-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 8 primer mixes used for DQ low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DQ serology3 | Amplified DQB1 alleles4 |
| --- | --- | --- | --- | --- |
| 16 | 225 bp | 515 bp | 5 | 050101-0505 |
| 2 | 220, 270 bp | 430 bp | 6 | 060101-0634 |
| 3 | 210 bp | 430 bp | 2 | 020101-0205 |
| 45 | 135, 150 bp | 515 bp | 2, 3, 6, 8 | 020101-0202, 0204, 0205, 030201-030204, 030501w, 030503w, 0307, 0308, 0311, 0318, 0629 |
| 55 | 220 bp | 515 bp | 3, 7, 8 | 030101-030104, 0304, 0309, 0310, 0313, 0314, 0316, 0319, 0321, 0322, 0324 |
| 65 | 220 bp | 515 bp | 2, 3, 4, 8, 9 | 020101-0205, 030201-030204, 030302-030303, 0306-0308, 0311, 0312, 0315, 0318, 0320, 0323, 0403 |
| 75,6 | 135, 155 bp | 515 bp | 3, 7, 8, 9 | 030101-030204, 030302-0324 |
| 85 | 210 bp | 430 bp | 4 | 0401-0403 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQ low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DQ low resolution typing.

In addition, wells number 4, 5, 6 and 7 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of the DQB1*0505, DQB1*0606 to 0608 alleles, the DQB1*0610 to 0634, the DQB1*030202-030204, DQB1*030303, DQB1*030502 to 0324 and the DQB1*0403 alleles is not known. In this table we have inferred the serological grouping from the naming of the sequence-defined allele.

4For several DQB1 alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 5’- and 3’-ends of the second exon of the DQB1 gene are conserved within allelic groups.

5These primer mixes may yield somewhat less intense specific PCR fragments than the other DQ low resolution primer mixes.

6Thess primer mixes may give rise to nonspecific amplifications.

1The sequence of the DQB1*030301 allele has been shown to be identical to DQB1*030302.

2The serological reactivity of the DQB1*0505, DQB1*0606 to 0608 alleles, the DQB1*0610 to 0634, the DQB1*030202-030204, DQB1*030303, DQB1*030502 to 0324 and the DQB1*0403 alleles is not known. In this table we have inferred the serological grouping from the naming of the sequence-defined allele.

Changes in revision R02 compared to R01:

- The DQB1*030501 and DQB1*030503 are weakly amplified by primer mix 4.