Gel Documentation Form and Worksheet

DQ-DR Combi Tray	Lot No: 19Y	Expiry Date: 2017-September-01

(101.704-48/12, -48u/12u)

Name:________________

Sample ID:_______________

DNA Conc.(ng/ul):________

Test  Date: ______________

Tested By: ______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	                        Failed lanes: ___________                                Comments:_________________

Gel Picture

PHOTO DOCUMENT

DQ low resolution

DR low resolution

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 4 to 6, 8 and 12 may give rise to a lower yield of HLA-specific PCR product than the other DQ low primer mixes.

Primer mix 6, 7, 14, 16, 21, 28, 29, 33, 35 and 43 may have a tendency to giving rise to primer oligomer formation.

Primer mix 13, 16, 17 and 22 may have a tendency of unspecific amplification.

Due to sharing of sequence motifs in codon 38 and 47, DRB3*01:14 will also be amplified in primer mixes 18, 19 and 30 and DRB3*01:23 in primer mix 18 in addition to primer mix 42.

Primer mix 32 has a tendency of giving rise to primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

The DRB4*01:03:01:02N allele is amplified by primer mix 43, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer pair.

Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 2, 8, 14, 18 to 21, 23, 24, 26, 29, 31 to 33, 35, 40 and 42.

Primer mix 45 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

DQ low resolution Interpretation Table

1DQB1 alleles listed on the IMGT/HLA web page 2014-October-10, release 3.18.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all DQ alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele. The DQB1*03:10 allele has been assigned type DQ7 by NMDP.

‘w’, might be weakly amplified.

DR low resolution Interpretation Table

1DRB alleles listed on the IMGT/HLA web page 2014-October-10, release 3.18.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological split of all DRB1 alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4The DRB1*03:76 and the DRB1*13:176 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*03:11:01 and DRB1*13:02:02 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- Primer mix 18 does not amplify the DRB1*03:11:01, *03:97, *11:53 and *15:100 alleles. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R02 compared to R01:

- The DR low Interpretation Table has been corrected.

Change in revision R03 compared to R02:

- Primer mix 35 does not amplify the DRB1*08:14 allele. This has been corrected in the Specificity and Interpretation Tables.

Change in revision R04 compared to R03:

- Due to sharing of sequence motifs in codon 38 and 47, DRB3*01:14 will also be amplified in primer mixes 18, 19 and 30 and DRB3*01:23 in mix 18, in addition to primer mix 42.

- The DRB1*13:02:02 allele is amplified in primer mix 19.

The Specificity and Interpretation Tables have been changed.