PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM DQA1

Prod. No:	101.231-24/04 – including Taq pol.

101.231-24u/04u – without Taq pol.

Lot No:    	Y46

Expiry Date:    	2009-October-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for DQA1 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DQA1 alleles3 |
| --- | --- | --- | --- |
| 14 | 145 bp | 515 bp | 010101-010102, 010401-010402, 0105, 0107 |
| 2 | 170 bp | 515 bp | 010101-010102, 010201-010204, 010401-010402, 0105, 0106, 0107 |
| 34 | 145 bp | 430 bp | 010201-010204, 0103, 0106 |
| 4 | 170 bp | 430 bp | 0103 |
| 5 | 220 bp | 430 bp | 010401-010402, 0105, 0106?, 0107? |
| 64 | 100 bp | 430 bp | 010401-010402, 0106?, 0107? |
| 74 | 95 bp | 430 bp | 0106 |
| 84 | 65 bp | 430 bp | 010101-010203, 0103, 0106?, 0107?, 0201-0602 |
| 9 | 175 bp | 430 bp | 0201 |
| 10 | 185 bp | 430 bp | 030101, 0302, 0303 |
| 115 | 215 bp | 430 bp | 0302 |
| 125 | 225 bp | 515 bp | 0302, 0303 |
| 13 | 225 bp | 515 bp | 010101-030101, 040101, 0402-0602 |
| 144 | 125 bp | 430 bp | 040101-040102, 0402, 0404, 050101-0509 |
| 15 | 165 bp | 430 bp | 050101-0509 |
| 164 | 90 bp | 430 bp | 0502 |
| 175 | 200 bp | 430 bp | 050101-050102, 0502?, 0504?, 0505, 0508, 0509 |
| 185 | 200 bp | 430 bp | 0502?, 0503, 0504?, 0506, 0507 |
| 19 | 205 bp | 430 bp | 050101-0503, 0505-0509 |
| 204 | 130 bp | 430 bp | 0504 |
| 214,5,6 | 110 bp | 430 bp | 0502?, 0504?, 0505, 0508, 0509 |
| 224 | 110 bp | 515 bp | 060101-060102, 0602 |
| 234,5 | 85 bp | 430 bp | 040101-0404, 060101-0602 |
| 245 | 240 bp | 430 bp | 010101-0107, 0201weakly, 030101, 0302-050102, 0502?, 0503, 0504?, 0506, 0507, 060101-0602 |
| 257 | 175, 275 bp | 430 bp | 0107, 0509weakly |
| 26 | 150 bp | 430 bp | 0402 |
| 274,5 | 90 bp | 430 bp | 0403N |
| 284 | 105 bp | 430 bp | 0404 |
| 29 | 160 bp | 430 bp | 0602 |
| 30 | 215 bp | 430 bp | 0506 |
| 314,5 | 100 bp | 430 bp | 0507 |
| 324 | 135 bp | 430 bp | 0508 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQA1 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 15 bp or more. Size differences shorter than 15 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in tubes 11, 12, 17, 18, 21, 23 and 27.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most tubes, or a band of 515 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DQA1 typing.

In addition, tubes number 2, 12, 13 and 22 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DQA1 alleles only partial 1st, 2nd, 3rd and 4th exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences are conserved within allelic groups.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mixes 11, 12, 17, 18, 21, 23, 24, 27 and 31 may give rise to nonspecific amplifications.

6Primer mix 21 may yield somewhat less specific PCR product than the other DQA1 primer mixes.

7Primer mix 25: Specific PCR product of 175 bp in the DQA1*0107 allele. Specific PCR product of 275 bp in the DQA1*0509 allele.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the DQA1*030102 allele has been shown to be identical to DQA1*0302.

DQA1*050103 has been renamed to DQA1*0505.

2Primer mix 25: Specific PCR product of 175 bp in the DQA1*0107 allele. Specific PCR product of 275 bp in the DQA1*0509 allele.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- The DQA1*0509 allele is weakly amplified by primer mix 25.