PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DQA1

Prod. No:	101.231-24/24u

101.231-04/04u

Lot No:    	82R

Expiry Date:    	2015-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for DQA1 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DQA1 alleles3 |
| --- | --- | --- | --- |
| 1 | 145 bp | 515 bp | *01:01:01-01:01:02, 01:04:01:01-01:05, 01:07 |
| 2 | 170 bp | 515 bp | *01:01:01-01:02:04, 01:04:01:01-01:09, 01:11 |
| 3 | 145 bp | 430 bp | *01:02:01:01-01:03:01:02, 01:06, 01:08-01:11 |
| 4 | 170 bp | 430 bp | *01:03:01:01-01:03:01:02, 01:10 |
| 512 | 205 bp | 430 bp | *01:05 |
| 64 | 100 bp | 430 bp | *01:04:01:01-01:04:02, 01:06?, 01:07 |
| 74,5 | 95 bp | 430 bp | *01:06 |
| 84 | 65 bp | 430 bp | *01:01:01-01:02:03, 01:03:01:01-01:03:01:02, 01:06?, 01:08-01:11, 02:01, 03:01:01-03:03:02, 04:01:01-04:04, 05:01:01:01-05:11, 06:01:01-06:02 |
| 9 | 175 bp | 430 bp | *02:01 |
| 10 | 185 bp | 430 bp | *03:01:01-03:03:02 |
| 11 | 215 bp | 430 bp | *03:02 |
| 12 | 225 bp | 515 bp | *03:02-03:03:02 |
| 13 | 225 bp | 515 bp | *01:01:01-01:11, 02:01, 03:01:01, 04:01:01, 04:02-04:04, 05:01:01:01-05:11, 06:01:01-06:02 |
| 144,5 | 125 bp | 430 bp | *04:01:01-04:02, 04:04, 05:01:01:01-05:11 |
| 15 | 165 bp | 430 bp | *05:01:01:01-05:09, 05:11 |
| 164 | 90 bp | 430 bp | *05:02 |
| 17 | 200 bp | 430 bp | *05:01:01:01-05:01:02, 05:02?, 05:04?, 05:05:01:01-05:05:01:03, 05:08-05:11 |
| 18 | 200 bp | 430 bp | *05:02?, 05:03, 05:04?, 05:06-05:07 |
| 19 | 205 bp | 430 bp | *05:01:01:01-05:03, 05:05:01:01-05:09, 05:11 |
| 20 | 135 bp | 430 bp | *05:04 |
| 214,7 | 100 bp, 210 bp | 430 bp | *01:09, 05:02?, 05:04?, 05:05:01:01-05:05:01:03, 05:08-05:09, 05:10?, 05:11 |
| 224,6,8 | 120 bp, 215 bp | 515 bp | *01:10, 06:01:01-06:02 |
| 234 | 85 bp | 430 bp | *04:01:01-04:04, 06:01:01-06:02 |
| 24 | 220 bp | 430 bp | *01:01:01-01:09, 02:01w, 03:01:01-03:03:02, 04:01:01-04:04, 05:01:01:01-05:01:02, 05:02?, 05:03, 05:04?, 05:06-05:07, 05:10?, 06:01:01-06:02 |
| 254,9 | 80 bp, 175 bp | 430 bp | *01:07, 05:09 |
| 26 | 160 bp | 430 bp | *04:02, 05:10 |
| 274,6,10 | 90 bp, 135 bp | 430 bp | *01:11, 04:03N |
| 284 | 105 bp | 430 bp | *04:04 |
| 29 | 150 bp | 430 bp | *01:08, 06:02 |
| 304,11 | 115 bp, 215 bp | 430 bp | *05:06, 05:11 |
| 314 | 100 bp | 430 bp | *05:07 |
| 32 | 135 bp | 430 bp | *05:08 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQA1 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 11, 12, 17, 18, 21, 23 and 27.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DQA1 typing.

In addition, wells number 2, 12, 13 and 22 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DQA1 alleles 1st and/or 3rd exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

5Primer mixes 7 and 14 may have a tendency to giving rise to primer oligomer formation.

6Primer mixes 22 and 27 may have tendencies of unspecific amplifications.

7Primer mix 21: Specific PCR product of 100 bp in the DQA1*05:02?, 05:04?, 05:05:01:01-05:05:01:03, 05:08-05:09, 05:10? and 05:11 alleles. Specific PCR product of 210 bp in the DQA1*01:09 allele.

8Primer mix 22: Specific PCR product of 120 bp in the DQA1*06:01:01-06:02 alleles. Specific PCR product of 215 bp in the DQA1*01:10 allele.

9Primer mix 25: Specific PCR product of 80 bp in the DQA1*05:09 allele. Specific PCR product of 175 bp in the DQA1*01:07 allele.

10Primer mix 27: Specific PCR product of 90 bp in the DQA1*04:03N allele. Specific PCR product of 135 bp in the DQA1*01:11 allele.

11Primer mix 30: Specific PCR product of 115 bp in the DQA1*05:11 allele. Specific PCR product of 215 bp in the DQA1*05:06 allele.

12Primer mix 5 may faintly amplify the *01:01:01-01:02:04 and 01:04:01:01-01:04:02 alleles.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1The sequence of the DQA1*03012 allele has been shown to be identical to DQA1*03:02.

DQA1*05013 has been renamed to DQA1*05:05.

2Primer mix 21: Specific PCR product of 100 bp in the DQA1*05:02?, 05:04?, 05:05:01:01-05:05:01:03, 05:08-05:09, 05:10? and 05:11 alleles. Specific PCR product of 210 bp in the DQA1*01:09 allele.

Primer mix 22: Specific PCR product of 120 bp in the DQA1*06:01:01-06:02 alleles. Specific PCR product of 215 bp in the DQA1*01:10 allele.

Primer mix 25: Specific PCR product of 80 bp in the DQA1*05:09 allele. Specific PCR product of 175 bp in the DQA1*01:07 allele.

Primer mix 27: Specific PCR product of 90 bp in the DQA1*04:03N allele. Specific PCR product of 135 bp in the DQA1*01:11 allele.

Primer mix 30: Specific PCR product of 115 bp in the DQA1*05:11 allele. Specific PCR product of 215 bp in the DQA1*05:06 allele.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- Primer mix 5 may faintly amplify the DQA1*01:01:01-01:02:04 and 01:04:01:01-01:04:02 alleles. A footnote has been added in the Specificity Table.