Olerup SSP® ## PHOTO DOCUMENT

Olerup SSP® DQA1*02,05;DQB1*02,03:02

Prod. No:	101.903-24/24u

Lot No:    	59N

Expiry Date:    	2014-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 15+1 primer mixes used for DQA1*02,05;DQB1*02,03:02 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DQA1 alleles3 | Amplified DQB1 alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 175 bp | 430 bp | *02:01 |  |
| 2 | 165 bp | 515 bp | *05:01:01:01-05:09, 05:11 |  |
| 36 | 210 bp | 515 bp |  | *02:01:01-02:06 |
| 44 | 80 bp | 515 bp |  | *03:01:01:01-03:01:06, 03:03:02:01-03:03:04, 03:09-03:10, 03:12-03:13, 03:15-03:17, 03:19-03:24, 03:26-03:31, 03:33-03:36, 03:38-03:39, *06:01:01w-06:01:06w, 06:07:01w-06:07:02w, 06:15w, 06:32w, 06:35w, 06:37w |
| 55,6 | 135 bp | 430 bp |  | *02:01:01-02:02, 02:04-02:06, 1*03:02:01-03:02:05, 03:07-03:08, 03:11, 03:18, 03:32, 03:37, *06:29 |
| 64 | 110 bp | 515 bp |  | *03:02:01-03:03:04, 03:06, 03:08, 03:11-03:12, 03:15, 03:18, 03:20, 03:25-03:26, 03:30-03:34, 03:37-03:39, *04:03:01-04:03:02, *06:19 |
| 74,6 | 115 bp | 515 bp |  | *03:06, 03:25 |
| 86,7 | 150 bp, 190 bp, 220 bp | 515 bp |  | *03:07, 03:18, 03:37 |
| 98,12 | 135 bp, 175 bp | 515 bp |  | *03:08, 03:32, *06:02:02, 06:03:02 |
| 109 | 135 bp, 260 bp | 515 bp |  | *03:09, 03:11, 03:26 |
| 116,10 | 145 bp, 185 bp | 515 bp |  | *03:01:01:01-03:39 |
| 124 | 90 bp | 515 bp | *05:02 |  |
| 13 | 200 bp | 430 bp | *05:01:01:01-05:01:02, 05:02?, 05:04?, 05:05:01:01-05:05:01:03, 05:08-05:11 |  |
| 14 | 200 bp | 515 bp | *05:02?, 05:03, 05:04?, 05:06-05:07 |  |
| 15 | 205 bp | 515 bp | *05:01:01:01-05:03, 05:05:01:01-05:09, 05:11 |  |
| 1611 | - | - |  | Negative control |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQA1 and DQB1 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 515 base pairs, for most wells, or a band of 430 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 430 bp, internal positive control band in order to help in the correct orientation of the DQA1*02,05;DQB1*02,03:02 typing.

In addition, wells number 5 and 13 contain the primer pair giving rise to the shorter, 430 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DQA1 and DQB1 alleles only partial second exon nucleotide sequences are available or nucleotide sequence information is not available for the 1st and 3rd exons. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences of DQA1 and DQB1 alleles are conserved within allelic groups.

4Specific PCR fragments shorter than 125 base pairs are less intense and not as sharp as longer specific bands.

5Primer mix 5 may yield less specific PCR products than the other DQB1 primer mixes.

6Primer mixes 3, 5, 7, 8 and 11 have a tendency of primer oligomer formation.

7Primer mix 8: Specific PCR fragment of 150 bp in the DQB1*03:07 allele. Specific PCR fragment of 190 bp in the DQB1*03:37 allele. Specific PCR fragment of 220 bp in the DQB1*03:18 allele.

8Primer mix 9: Specific PCR fragment of 135 bp in the DQB1*03:08 and the DQB1*06:02:02 and 06:03:02 alleles. Specific PCR fragment of 175 bp in the DQB1*03:32 alleles.

9Primer mix 10: Specific PCR fragment of 135 bp in the DQB1*03:09 allele. Specific PCR fragment of 260 bp in the DQB1*03:11 and 03:26 alleles.

10Primer mix 11: Specific PCR fragment of 145 bp in the DQB1*03:01:02, 03:02:03, 03:03:03, 03:05:01-03:05:02, 03:20, 03:23 and 03:37 alleles. Specific PCR fragment of 185 bp in the DQB1*03:25 allele. Specific PCR fragment of 145 and 185 bp in the DQB1*03:01:01:01-03:01:01:03, 03:01:03-03:02:02, 03:02:04-03:03:02:03, 03:04, 03:05:03-03:19, 03:21-03:22, 03:24, 03:26-03:36 and 03:38-03:39 alleles. Both bands may not always be visible.

11Primer mix 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

12Primer mix 9 may have tendencies of unspecific amplifications.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

1DQA1*05013 has been renamed to DQA1*05:05.

The sequence of the DQB1*03031 allele has been shown to be identical to DQB1*03:03:02.

2Primer mix 9: Specific PCR fragment of 135 bp in the DQB1*03:08 and the DQB1*06:02:02 and 06:03:02 alleles. Specific PCR fragment of 175 bp in the DQB1*03:32 alleles.

Primer mix 10: Specific PCR fragment of 135 bp in the DQB1*03:09 allele. Specific PCR fragment of 260 bp in the DQB1*03:11 and 03:26 alleles.

Primer mix 11: Specific PCR fragment of 145 bp in the DQB1*03:01:02, 03:02:03, 03:03:03, 03:05:01-03:05:02, 03:20, 03:23 and 03:37 alleles. Specific PCR fragment of 185 bp in the DQB1*03:25 allele. Specific PCR fragment of 145 and 185 bp in the DQB1*03:01:01:01-03:01:01:03, 03:01:03-03:02:02, 03:02:04-03:03:02:03, 03:04, 03:05:03-03:19, 03:21-03:22, 03:24, 03:26-03:36 and 03:38-03:39 alleles. Both bands may not always be visible.

Primer mix 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by a control primer pair. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘w’, might be weakly amplified.‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- The control primer pair for the Negative Control, well 16, has been corrected in the Specificity table.

Change in revision R02 compared to R01:

- Primer mix 9 may have tendencies of unspecific amplifications. A footnote has been added in the Specificity Table.