Gel Documentation Form and Worksheet

DQB1*03 (101.214-24/06, -24u/06u) 	Lot No: 94X	Expiry Date: 2017-August-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:________________

Reviewed By:_________________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 3, 5, 23, 25, 34 and 46 may have a tendency to giving rise to primer oligomer formation.

Primer mixes 6, 11, 14, 16 and 24 may have tendencies of unspecific amplifications.

Primer mixes 18, 19 and 27 may give rise to a lower yield of HLA-specific PCR product than the other DQB1*03 primer mixes.

Primer mix 15 may faintly amplify the DQB1*03:16 allele.

Primer mix 48 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-DQB1 alleles listed on the IMGT/HLA web page 2013-July-25, release 3.13.1, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 6: Specific PCR fragment of 95 bp in the DQB1*03:20 allele. Specific PCR fragment of 130 bp in the DQB1*03:05:01-03:05:04, 03:17:01-03:17:02, 03:61 and 03:72 and the DQB1*04:09 alleles.

Primer mix 7: Specific PCR fragment of 115 bp in the DQB1*03:03:06, 03:06, 03:25, 03:42 and 03:88 and the DQB1*04:01:03 alleles. Specific PCR fragment of 150 bp in the DQB1*03:84N alleles.

Primer mix 8: Specific PCR fragment of 110 bp in the DQB1*03:15 allele. Specific PCR fragment of 140 bp in the DQB1*03:07 and 03:16 alleles.

Primer mix 9: Specific PCR fragment of 65 bp in the DQB1*03:40 allele. Specific PCR fragment of 135 bp in the DQB1*03:08 and the DQB1*06:02:02 and 06:03:02 alleles.

Primer mix 10: Specific PCR fragment of 135 bp in the DQB1*03:09 allele. Specific PCR fragment of 260 bp in the DQB1*03:11 and 03:26 alleles.

Primer mix 11: Specific PCR fragment of 135 bp in the DQB1*03:10:01, 03:12, 03:14:01-03:14:02 and 03:70 and the DQB1*06:01:01-06:01:06, 06:01:08-06:01:10, 06:43, 06:54N-06:58, 06:98 and 06:100-06:105 alleles. Specific PCR fragment of 255 bp in the DQB1*03:47 allele.

Primer mix 13: Specific PCR fragment of 165 bp in the DQB1*03:13 allele. Specific PCR fragment of 205 bp in the DQB1*03:48 allele.

Primer mix 17: Specific PCR fragment of 165 bp in the DQB1*03:43 and 03:87 alleles. Specific PCR fragment of 220 bp in the DQB1*03:18 and 03:85 alleles.

Primer mix 23: Specific PCR fragment of 95 bp in the DQB1*03:41 allele. Specific PCR fragment of 125 bp in the DQB1*03:24, 03:79 and 03:89 alleles.

Specific PCR fragment of 150 bp in the DQB1*03:84N. Specific PCR fragment of 185 bp in the DQB1*03:45 allele.

Primer mix 24: Specific PCR fragment of 120 bp in the DQB1*03:39 allele. Specific PCR fragment of 210 bp in the DQB1*03:27, 03:48 and 03:56 alleles.

Primer mix 25: Specific PCR fragment of 130 bp in the DQB1*03:24 and 03:79 alleles. Specific PCR fragment of 165 bp in the DQB1*03:29 and 03:80 alleles. Specific PCR fragment of 240 bp in the DQB1*03:34 and the DQB1*02:14:01 alleles.

Primer mix 27: Specific PCR fragment of 170 bp in the DQB1*03:32 and 03:87 alleles. Specific PCR fragment of 225 bp in the DQB1*03:85 allele. Specific PCR fragment of 250 bp in the DQB1*03:31 allele.

Primer mix 29: Specific PCR fragment of 110 bp in the DQB1*03:33 allele. Specific PCR fragment of 165 bp in the DQB1*03:35 allele.

Primer mix 31: Specific PCR fragment of 200 bp in the DQB1*03:28 and 03:56 alleles. Specific PCR fragment of 225 bp in the DQB1*03:83 alleles.

Primer mix 34: Specific PCR fragment of 175 bp in the DQB1*03:50 and 03:58 alleles. Specific PCR fragment of 200 bp in the DQB1 *03:82 allele.

Primer mix 37: Specific PCR fragment of 180 bp in the DQB1*03:60 and 03:77 allele. Specific PCR fragment of 230 bp in the DQB1*03:53 and 03:62 alleles.

Primer mix 38: Specific PCR fragment of 175 bp in the DQB1*03:58 allele. Specific PCR fragment of 230 bp in the DQB1*03:57 and 03:83 alleles.

Primer mix 39: Specific PCR fragment of 145 bp in the DQB1*03:66N allele. Specific PCR fragment of 190 bp in the DQB1*03:68 allele. Specific PCR fragment of 235 bp in the DQB1*03:86 allele.

Primer mix 45: Specific PCR fragment of 90 bp in the DQB1*03:65 allele. Specific PCR fragment of 230 bp in the DQB1*03:81 allele.

4This lot of the DQB1*03 subtyping kit cannot distinguish the DQB1*03:01:01:01-03:01:02, 03:01:04-03:01:07, 03:01:09-03:01:12, 03:01:14 and 03:01:16-03:01:20 alleles and the DQB1*03:22 alleles.

The DQB1*03 subtyping kit cannot distinguish the following silent mutations: the DQB1*03:01:01:01-03:01:02, 03:01:04-03:01:07, 03:01:09-03:01:12, 03:01:14, 03:01:16-03:01:20, the DQB1*03:01:13 and 03:01:15, the DQB1*03:02:01-03:02:09, the DQB1*03:03:02:01-03:03:05 and DQB1*03:03:07, the DQB1*03:05:01-03:05:04, the DQB1*03:14:01-03:14:02 and the DQB1*03:17:01-03:17:02 alleles.

‘?’, nucleotide sequence information not available for the primer matching region.

Change in revision R01 compared to R00:

- Primer mix 21 does not amplify the DQB1*03:22 allele. This has been corrected in the Specificity and Interpretation Tables.