Gel Documentation Form and Worksheet

DQB1*03 (101.214-24/06, -24u/06u) 	Lot No: 4D9	Expiry Date: 2018-08-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:________________

Reviewed By:_________________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 3, 5, 23, 25, 34, 43 and 46 may have a tendency to giving rise to primer oligomer formation.

Primer mixes 6, 11, 14, 16, 24 and 50 may have tendencies of unspecific amplifications.

Primer mixes 1, 2, 18, 19 and 27 may give rise to a lower yield of HLA-specific PCR product than the other DQB1*03 primer mixes.

Primer mixes 55 and 58 may give rise to a long unspecific amplification product. This should be disregarded when interpreting the DQB1*03 typings.

Primer mix 60 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-DQB1 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.22.0, October 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 6: Specific PCR fragment of 95 bp in the DQB1*03:20 allele. Specific PCR fragment of 130 bp in the DQB1*03:05:01-03:05:04, 03:17:01-03:17:02, 03:61, 03:72, 03:100, 03:132 and 03:181 and the DQB1*04:09 alleles.

Primer mix 7: Specific PCR fragment of 115 bp in the DQB1*03:03:06, 03:06, 03:25, 03:42 and 03:88 and the DQB1*04:01:03 alleles. Specific PCR fragment of 150 bp in the DQB1*03:84N alleles.

Primer mix 8: Specific PCR fragment of 110 bp in the DQB1*03:15 allele. Specific PCR fragment of 145 bp in the DQB1*03:07 and 03:16 alleles.

Primer mix 9: Specific PCR fragment of 65 bp in the DQB1*03:40 allele. Specific PCR fragment of 145 bp in the DQB1*03:08, 03:119, 03:137 and 03:194 and the DQB1*06:02:02 and 06:03:02 alleles.

Primer mix 10: Specific PCR fragment of 135 bp in the DQB1*03:09 allele. Specific PCR fragment of 260 bp in the DQB1*03:11 and 03:26 alleles.

Primer mix 11: Specific PCR fragment of 135 bp in the DQB1*03:10:01-03:10:02, 03:12, 03:14:01-03:14:02, 03:70, 03:179, 03:183 and 03:195 and the DQB1*06:01:01-06:01:06, 06:01:08-06:01:15, 06:06? 06:43, 06:54N-06:58, 06:98, 06:99:02-06:105, 06:108, 06:120, 06:132, 06:140, 06:153, 06:157, 06:167-06:168, 06:177, 06:181 and 06:194 alleles. Specific PCR fragment of 255 bp in the DQB1*03:47 allele.

Primer mix 13: Specific PCR fragment of 165 bp in the DQB1*03:13 allele. Specific PCR fragment of 205 bp in the DQB1*03:48 allele.

Primer mix 17: Specific PCR fragment of 165 bp in the DQB1*03:43 and 03:87 alleles. Specific PCR fragment of 225 bp in the DQB1*03:18 and 03:85 alleles.

Primer mix 23: Specific PCR fragment of 95 bp in the DQB1*03:41 allele. Specific PCR fragment of 125 bp in the DQB1*03:24, 03:79 and 03:89 alleles.

Specific PCR fragment of 150 bp in the DQB1*03:84N allele. Specific PCR fragment of 185 bp in the DQB1*03:45 allele.

Primer mix 24: Specific PCR fragment of 120 bp in the DQB1*03:39 allele. Specific PCR fragment of 210 bp in the DQB1*03:27, 03:48 and 03:56 alleles.

Primer mix 25: Specific PCR fragment of 105 bp in the DQB1*03:206 allele. Specific PCR fragment of 130 bp in the DQB1*03:24 and 03:79 alleles. Specific PCR fragment of 160 bp in the DQB1*03:29 and 03:80 alleles. Specific PCR fragment of 240 bp in the DQB1*03:34 and the DQB1*02:14:01 alleles.

Primer mix 27: Specific PCR fragment of 170 bp in the DQB1*03:32 and 03:87 alleles. Specific PCR fragment of 225 bp in the DQB1*03:85 allele. Specific PCR fragment of 250 bp in the DQB1*03:31 allele.

Primer mix 29: Specific PCR fragment of 110 bp in the DQB1*03:33 allele. Specific PCR fragment of 165 bp in the DQB1*03:35 allele.

Primer mix 31: Specific PCR fragment of 200 bp in the DQB1*03:28 and 03:56 alleles. Specific PCR fragment of 225 bp in the DQB1*03:83 allele.

Primer mix 35: Specific PCR fragment of 125 bp in the DQB1*03:190 and the DQB1*05:102 alleles. Specific PCR fragment of 190 bp in the DQB1*03:51 and 03:54 alleles. Specific PCR fragment of 235 bp in the DQB1*03:117 allele.

Primer mix 37: Specific PCR fragment of 180 bp in the DQB1*03:60 and 03:77 allele. Specific PCR fragment of 230 bp in the DQB1*03:53 and 03:62 alleles.

Primer mix 38: Specific PCR fragment of 175 bp in the DQB1*03:58 allele. Specific PCR fragment of 230 bp in the DQB1*03:57 and 03:83 alleles.

Primer mix 39: Specific PCR fragment of 145 bp in the DQB1*03:66N and the DQB1*06:193N alleles. Specific PCR fragment of 190 bp in the DQB1*03:68 allele. Specific PCR fragment of 235 bp in the DQB1*03:86 allele.

Primer mix 42: Specific PCR fragment of 80 bp in the DQB1*03:188 allele. Specific PCR fragment of 165 bp in the DQB1*03:64 and 03:76-03:77 alleles.

Primer mix 43: Specific PCR fragment of 160 bp in the DQB1*03:49 and the DQB1*02:10 alleles. Specific PCR fragment of 185 bp in the DQB1*03:198 allele. Specific PCR fragment of 250 bp in the DQB1*03:97 and the DQB1*06:86 and 06:104 alleles.

Primer mix 44: Specific PCR fragment of 160 bp in the DQB1*03:211 allele. Specific PCR fragment of 220 bp in the DQB1*03:78 allele. Specific PCR fragment of 270 bp in the DQB1*03:71 allele.

Primer mix 45: Specific PCR fragment of 90 bp in the DQB1*03:65 allele. Specific PCR fragment of 190 bp in the DQB1*03:204 allele. Specific PCR fragment of 230 bp in the DQB1*03:81 and 03:159 alleles.

Primer mix 46: Specific PCR fragment of 115 bp in the DQB1*03:88-03:89 alleles. Specific PCR fragment of 175 bp in the DQB1*03:191 allele.

Primer mix 49: Specific PCR fragment of 150 bp in the DQB1*03:94. 03:164 and 03:182 alleles. Specific PCR fragment of 175 bp in the DQB1*03:191 allele.

Primer mix 50: Specific PCR fragment of 105 bp in the DQB1*03:206 allele. Specific PCR fragment of 150 bp in the DQB1*03:98, 03:165 and 03:182 alleles.

Primer mix 51: Specific PCR fragment of 105 bp in the DQB1*03:90N allele. Specific PCR fragment of 160 bp in the DQB1*03:115 and 03:213N and the DQB1*06:109 alleles.

Primer mix 52: Specific PCR fragment of 80 bp in the DQB1*03:188 allele. Specific PCR fragment of 190 bp in the DQB1*03 99Q, 03:118N, 03:162 and 03:183 alleles.

Primer mix 53: Specific PCR fragment of 125 bp in the DQB1*03:130 allele. Specific PCR fragment of 155 bp in the DQB1*03:139 allele. Specific PCR fragment of 190 bp in the DQB1*03:162, 03:183 and 03:204 alleles.

Primer mix 54: Specific PCR fragment of 95 bp in the DQB1*03:91Q allele. Specific PCR fragment of 170 bp in the DQB1*03:158 and 03:170 alleles. Specific PCR fragment of 220 bp in the DQB1*03:181 allele.

Primer mix 55: Specific PCR fragment of 155 bp in the DQB1*03:102 allele. Specific PCR fragment of 195 bp in the DQB1*03:114 allele.

Primer mix 56: Specific PCR fragment of 120 bp in the DQB1*03:95N, 03:116 and 03:210 alleles. Specific PCR fragment of 160 bp in the DQB1*03:170 allele.

Primer mix 57: Specific PCR fragment of 100 bp in the DQB1*03:196 and the DQB1*05:45 alleles. Specific PCR fragment of 215 bp in the DQB1*03:169 allele.

Primer mix 58: Specific PCR fragment of 155 bp in the DQB1*03:102 allele. Specific PCR fragment of 180 bp in the DQB1*03:197Q allele.

4The following DQB1*03 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| DQB1*03:40, 03:137 | 9 |
| DQB1*03:169, 03:196 | 57 |

The DQB1*03 subtyping kit cannot distinguish the following silent mutations: the DQB1*03:01:01:01-03:01:02, 03:01:04-03:01:07, 03:01:09-03:01:12, 03:01:14, 03:01:16-03:01:24, 03:01:26-03:01:31, the DQB1*03:01:13, 03:01:15 and 03:01:25, the DQB1*03:02:01-03:02:19, the DQB1*03:03:02:01-03:03:05 and DQB1*03:03:07-03:03:13, the DQB1*03:04:01-03:04:02, the DQB1*03:05:01-03:05:04, the DQB1*03:14:01-03:14:02 and the DQB1*03:17:01-03:17:02 alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching region.

Change in revision R01 compared to R00:

- Primer mix 49 amplifies the DQB1*03:164 allele. This has been corrected in the Specificity and Interpretation Tables.