DQB1\*03 (101.214-24/06, -24u/06u) Lot No: 4N9 Expiry Date: 2025-10-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

*Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |







Abbreviations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 3, 5, 23, 25, 30, 43, 45, 46, 48, 51, 52 and 61 may have a tendency to giving rise to primer oligomer formation.

Primer mixes 6, 11, 14, 16, 19, 24 and 50 may have tendencies of unspecific amplifications.

Primer mixes 1, 2, 18, 27 and 33 may give rise to a lower yield of HLA-specific PCR product than the other DQB1\*03 primer mixes.

Primer mixes 55 and 58 may give rise to a long unspecific amplification product. This should be disregarded when interpreting the DQB1\*03 typings.

Primer mix 64 contains a negative control, which will amplify the majority of the HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.





















**1**HLA-DQB1 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.29.0, August 2017.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following DQB1\*03 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified DQB1\*03alleles | Other amplifiedDQB1 alleles |
|  **6**  | 95 bp 130 bp | \*03:20\*03:05:01-03:05:05, 03:17:01-03:17:02, 03:61, 03:72, 03:100, 03:132, 03:181, 03:226, 03:244, 03:250, 03:262, 03:272, 03:346 | \*04:09, 04:63, 04:75, 05:131, 05:276 |
|  **7** | 115 bp150 bp | \*03:03:06, 03:06, 03:25:01-03:25:02, 03:42, 03:88\*03:84N | \*04:01:03, 04:02:16, 04:03:03 |
|  **8** | 110 bp145 bp | \*03:15\*03:07, 03:16, 03:438 |  |
|  **9** | 65 bp 145 bp | \*03:40, 03:353\*03:08, 03:119, 03:137, 03:194 | \*06:02:02, 06:03:02, 06:03:34 |
| **10** | 135 bp 260 bp | \*03:09\*03:11, 03:26, 03:244 |  |
| **11** | 140 bp 255 bp | \*03:10:01-03:10:03, 03:12, 03:14:01-03:14:02, 03:70, 03:179, 03:183, 03:195\*03:47 | \*06:01:01:01-06:01:06, 06:01:08-06:01:30, 06:06?, 06:43, 06:54N-06:58, 06:98, 06:99:02-06:105, 06:108, 06:120, 06:132, 06:140, 06:153:01-06:153:02, 06:157, 06:167-06:168, 06:177, 06:181, 06:194, 06:205, 06:209, 06:214, 06:229, 06:239, 06:243, 06:245-06:247, 06:251, 06:257-06:258, 06:260, 06:263, 06:268, 06:274, 06:277, 06:285, 06:302, 06:305, 06:307, 06:309-06:310, 06:312, 06:321, 06:323, 06:330N, 06:342, 06:359, 06:377, 06:382 |
| **13** | 115 bp170 bp 205 bp | \*03:255\*03:13, 03:232\*03:48 |  |
| **17** | 165 bp 225 bp | \*03:43, 03:87\*03:18, 03:85 |  |
| **20** | 150 bp240 bp | \*03:21, 03:452\*03:246 | \*02:64, 02:79, 05:35, 05:182, 05:259, 06:243, 06:255\*06:310 |
| **23**  | 95 bp 125 bp150 bp 185 bp220 bp | \*03:41, 03:224\*03:24-03:25:02, 03:79, 03:89, 03:234, 03:239, 03:447\*03:84N\*03:45:01-03:45:02\*03:258 | \*04:03:01-04:03:03, 04:81 |
| **24**  | 120 bp 210 bp | \*03:39\*03:27, 03:48, 03:56 |  |
| **25**  | 105 bp130 bp155 bp180 bp 240 bp | \*03:206\*03:24, 03:79, 03:447\*03:80, 03:257\*03:29, 03:118N, 03:237N\*03:34 | \*04:81\*02:14:01:01-02:14:01:02 |
| **27** | 170 bp 225 bp250 bp | \*03:32, 03:87\*03:85\*03:31 |  |
| **28** | 200 bp230 bp | \*03:37-03:38:02, 03:46, 03:252, 03:277\*03:225 |  |
| **29** | 110 bp 165 bp255 bp | \*03:33, 03:245, 03:328\*03:35, 03:315\*03:243 | \*05:281, 06:228, 06:304N\*02:126 |
| **30** | 135 bp175 bp | \*03:36, 03:149\*03:150 |  |
| **31** | 115 bp195 bp225 bp | \*03:255\*03:28\*03:56, 03:83 |  |
| **33** | 85 bp160 bp | \*03:59\*03:204 | \*06:37, 06:374 |
| **35** | 135 bp190 bp235 bp | \*03:190, 03:205, 03:214\*03:51, 03:54, 03:287\*03:117 | \*02:157, 04:15, 04:42, 05:102, 06:223\*04:35 |
| **36** | 145 bp175 bp200 bp | \*03:205, 03:214\*03:55, 03:60, 03:250\*03:269N | \*04:15 |
| **37** | 180 bp 230 bp | \*03:60, 03:77\*03:53, 03:62, 03:225 |  |
| **38** | 170 bp 225 bp | \*03:58, 03:194, 03:232\*03:57, 03:83 | \*06:82, 06:233, 06:302 |
| **39** | 105 bp140 bp 190 bp235 bp | \*03:240\*03:25:01-03:25:02, 03:66N, 03:234, 03:239\*03:68\*03:86 | \*04:03:01-04:03:03, 06:193N \*02:151 |
| **40** | 180 bp230 bp | \*03:67-03:68, 03:228\*03:104, 03:212 | \*06:17, 06:24, 06:30, 06:42, 06:318\*02:47 |
| **41** | 135 bp175 bp210 bp | \*03:282N\*03:63, 03:75, 03:168\*03:201, 03:269N |  |
| **42** | 80 bp165 bp220 bp | \*03:188\*03:64, 03:76-03:77\*03:201 |  |
| **43** | 175 bp250 bp | \*03:49:01:01-03:49:01:02, 03:198:01, 03:222, 03:247, 03:285-03:286, 03:402\*03:97, 03:329, 03:349, 03:381 | \*02:10, 02:81, 05:152, 06:359, 06:381\*06:86, 06:104, 06:366 |
| **44** | 160 bp220 bp270 bp | \*03:199, 03:211\*03:78\*03:71 |  |
| **45** | 90 bp135 bp170 bp230 bp | \*03:65\*03:282N\*03:133, 03:200, 03:462\*03:81, 03:159 | \*02:84, 04:14\*02:116 |
| **46**  | 115 bp175 bp | \*03:88-03:89\*03:191 |  |
| **49** | 150 bp175 bp | \*03:94, 03:164, 03:182\*03:191 | \*04:44 |
| **50**  | 105 bp150 bp | \*03:206\*03:98, 03:165, 03:182 |  |
| **51** | 105 bp140 bp165 bp220 bp | \*03:90N, 03:240\*03:239\*03:115, 03:213N, 03:228, 03:368\*03:258  | \*02:106, 04:43, 05:246, 06:17, 06:24, 06:30, 06:42, 06:109 |
| **52** | 80 bp185 bp | \*03:188\*03:99Q, 03:118N, 03:133, 03:162, 03:183, 03:200, 03:223, 03:233, 03:237N, 03:351, 03:462 | \*02:84, 04:14 |
| **53** | 125 bp155 bp190 bp | \*03:130, 03:271\*03:139, 03:199\*03:162, 03:183, 03:204, 03:223, 03:233, 03:351 |  |
| **54** | 95 bp170 bp220 bp | \*03:91Q\*03:146, 03:158, 03:170, 03:207\*03:181 |  |
| **55** | 155 bp195 bp | \*03:102\*03:114 |  |
| **56** | 155 bp200 bp | \*03:95N, 03:210\*03:150, 03:170 |  |
| **57** | 100 bp215 bp240 bp | \*03:196, 03:442\*03:169\*03:246 | \*05:45, 05:201, 06:323\*06:327\*06:310 |
| **58** | 155 bp180 bp | \*03:102\*03:197Q |  |
| **61** | 115 bp155 bp255 bp | \*03:116\*03:221, 03:450\*03:243 | \*06:51:01-06:51:02, 06:377\*02:126 |

**4**The following DQB1\*03 alleles can be distinguished by the different sizes of the HLA-specific PCR product:



**5**The following alleles cannot be separated by the HLA-DQB1\*03 primer set. These alleles can be distinguished by the DQ low resolution kit and/or the HLA-DQB1\*04 high resolution kit.

|  |
| --- |
| Alleles |
| DQB1\*03:25:01, DQB1\*04:03:03 |
| DQB1\*03:72, 03:100, DQB1\*04:63 |
| DQB1\*03:132, DQB1\*04:75 |

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence information not available for the primer matching region.