PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DR low resolution

Prod. No:	101.101-48/12 – including Taq pol.

101.101-48u/12u – without Taq pol.

Lot No:    	40E

Expiry Date:    	2010-May-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 23 primer mixes used for DR low resolution SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | DR serology3 | Amplified HLA-DRB4 alleles |
| --- | --- | --- | --- | --- |
| 1 | 205 bp | 515 bp | 1 | 010101-010204, 0104-0119 |
| 2 | 195 bp | 430 bp | 103 | 0103 |
| 3 | 215, 260 bp | 430 bp | 15 | 150101-1527 |
| 4 | 210 bp | 430 bp | 16 | 160101-160502, 1607-1613N |
| 5 | 220 bp | 430 bp | 3, 17, 18, 11 | 030101-0339, 1107, 1153 |
| 65,6 | 80, 210 bp | 430 bp | 3, 17, 11, 13, 14 | 030101-030106, 0304-0306, 0308-0316, 0318-0320, 0322, 0323, 0325, 0326, 0328, 0330, 0331, 0333, 0334, 0336, 0337, 110201-1103, 111101-111102, 111401-111402, 1116, 1120, 1121, 1136, 1140, 1141, 1148, 1159, 1163, 116501-116502, 1168, 130101-1304, 1308, 1310, 1315-1317, 1319, 1320, 1322-1324, 1327-1329, 1331-1341, 1343, 1345, 1348, 1351-1354, 1357, 1359, 1361, 1363-1366, 1368-1376, 1378-1381, 1383, 1416, 1419, 1421 |
| 75,6 | 80, 210 bp | 430 bp | 3, 18, 11, 13, 14 | 030201-0303, 0327, 0329, 0338, 1126, 1134, 1315, 1319, 1326, 1344, 1353, 1357, 1402-140302, 1406, 1409, 1412, 1413, 1417-1421, 1424, 1427, 1429, 1430, 1433, 1440, 1441, 1447-1449, 1451, 1463, 1467, 1477 |
| 85,6 | 100, 175 bp | 430 bp | 4 | 040101-0473 |
| 9 | 205, 230 bp | 430 bp | 7, 13, 14 | 070101-070102, 0703-0715, 1317, 1450 |
| 10 | 165, 215, 250 bp | 515 bp | 8, 11, 12, 14 | 080101-080203, 080302-0819, 0821-0834, 1167, 1204, 1216, 1411, 1415, 1468 |
| 115 | 85, 130, 180 bp | 430 bp | 3, 9, 11 | 0308, 090102-0906, 1107, 1153 |
| 12 | 205 bp | 430 bp | 10 | 100101-1002 |
| 135,8 | 100, 170 bp | 430 bp | 11, 3, 8 | 0308, 0831, 110101-1168 |
| 145 | 85, 105 bp | 430 bp | 12, 8 | 0832, 120101-120202, 120302-1216 |
| 15 | 210 bp | 430 bp | 13, 8, 11, 14 | 0820, 110101-110404, 110601-110602, 110801-111202, 111401-1116, 1118-1121, 1123-1125, 112701-1133, 1135-1151, 115401-115402, 1156-1166, 1168, 130101-1308, 1310-1316, 1318-1343, 1345-1383, 140301-140302, 1412, 1416, 1419, 1421, 1422, 1425, 1427, 1440, 1453, 1463, 1467, 1469, 1474, 1477 |
| 166 | 195, 210 bp | 430 bp | 13, 8, 11, 12, 14 | 080101-080203, 080401-0809, 0811, 0816, 0817, 0820-0822, 0824, 0826, 0828, 0831, 110101-110602, 1109-111202, 111401-1116, 1120, 1121, 1123-1125, 112701-1130, 1132, 1133, 1135-1141, 1143, 1144, 1146-1151, 115401-1156, 1158-1163, 116501-1168, 120201-120202, 1213, 1215, 1216, 130101-130201, 130203, 1304-130502, 130701-1309, 1311, 131401-1324, 1326-1329, 1331, 1332, 1334-1336, 1338-1343, 1345-1355, 1357, 1359, 1361-1365, 1367-1376, 1378-1380, 1383, 1415, 1416, 1422, 1424, 1425, 1427, 1437, 1453, 1473 |
| 17 | 175 bp | 430 bp | 3, 8, 13, 14 | 030101-0307, 0309, 0311-0339, 0820, 130101-1316, 1318-1342, 1344, 1346-1366, 1368-1383, 1402-140302, 140501-1406, 1409, 1412-1414, 1417-1421, 142301, 1424, 1427, 1429, 1430, 1433, 1436, 1437, 1440-1445, 1447, 1448, 1451, 1456, 1459, 1463-1467, 1477 |
| 185-7 | 100, 140 bp | 430 bp | 14, 4, 8,13 | 0462, 0469, 0473, 0808, 1345, 140101-140102, 1404, 140701-140702, 1410, 1416, 1422, 1425, 1426, 1428, 1431-143202, 1435, 1437-1439, 1449, 1450, 1453-1455, 1457, 1458, 1460-1462, 1468-1471, 1473-1476 |
| 195-9 | 110, 135, 170 bp | 515 bp | 14, 3, 9, 11, 12, 13, 15 | 0308, 0310, 090102-0902, 0904-0906, 110404, 1109, 1110, 111202, 1113, 1116, 1117, 1120, 1122, 1140, 1146, 1152, 1158, 1159, 1204, 1343, 140101-1402, 1404-1411, 1413, 1414, 1416-142302, 1426, 1428-1436, 1438, 1439, 1441, 1443-1452, 1454-1457, 1459-1462, 1464, 1465, 1468, 1470-1476, 1527 |
| 205-7 | 110, 175, 225 bp | 430 bp | 14, 3, 8, 11, 13, 15, 16 | 0310, 0809, 0820, 0821, 0832, 1113, 1117, 1123, 1125, 1131, 1145, 1152, 1155, 1164, 1313, 1318, 1343, 1345, 1347, 1355, 140101-140103, 140301-140503, 140701-1408, 1410-1412, 1414-1416, 1418, 1422-142302, 1425-1428, 1431-143202, 1434-1436, 1438-1440, 1442-1445, 1449, 1450, 1453-1477, 1521weakly, 1604weakly |
| 21 | 230 bp | 430 bp | 52 | DRB3*01010201-0111, 0201-0223, 030101-0303 |
| 2210 | 215 bp | 430 bp | 53 | DRB4*01010101-0107 |
| 23 | 260 bp | 430 bp | 51 | DRB5*010101-0113, 0202-0205 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 3, 18, 19 and 20.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DR low resolution typing.

In addition, wells number 10 and 19 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DRB alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

4For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP set are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

The DRB1*0809, DRB1*0821 and DRB1*1415 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1415 allele.

The DRB1*0820, DRB1*1318, DRB1*1347 and DRB1*1355 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1318, DRB1*1347 and DRB1*1355 alleles.

The DRB1*0831 and DRB1*1167 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1167 allele.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to two differently sized specific PCR fragments in these primer mixes.

7These primer mixes may give rise to nonspecific amplifications.

8Primer mix 13 may give rise a primer oligomer formation.

9Primer mix 19 has a tendency of primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.

10The DRB4*01030102N allele is amplified by the primer pair in well No. 22, whereas the DRB4*0201N and DRB4*0301N null alleles are not amplified by this primer pair.

1The sequence of the DRB1*0702 allele has been shown to be identical to DRB1*070101.

The sequence of the DRB1*080301 allele has been shown to be identical to DRB1*080302.

The sequence of the DRB1*090101 allele has been shown to be identical to DRB1*090102.

The sequence of the DRB1*120301 allele has been shown to be identical to DRB1*1201.

The sequence of the DRB1*1606 allele has been shown to be identical to DRB1*1605.

The DRB4*0101102N allele has been renamed DRB4*0103102N.

The sequence of the DRB5*0201 allele has been shown to be identical to DRB5*0202.

The DRB1*0809, DRB1*0821 and DRB1*1415 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1415 allele.

The DRB1*0820, DRB1*1318, DRB1*1347 and DRB1*1355 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1318, DRB1*1347 and DRB1*1355 alleles.

The DRB1*0831 and DRB1*1167 alleles yield identical amplification patterns except for the specific PCR product yielded by the DRB3 gene in linkage disequilibrium with the DRB1*1167 allele.

2The serological reactivity of all DRB alleles is not known. In this table we use the information in the HLA Dictionary 2004 on the www.ebi.ac.uk/imgt/hla web site and the information available at the www.worldmarrow.org web site and have also inferred the serological grouping from the naming of the sequence-defined allele.

‘ser’, serological HLA specificity.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- The length of the specific PCR product in primer mix 12 has been corrected to 205 bp in the Interpretation Table.