PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*01

Prod. No:	101.111-24/24u

101.111-06/06u

Lot No:    	50F

Expiry Date:    	2011-January-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for DRB1*01 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB1*013 alleles |
| --- | --- | --- | --- |
| 14 | 255 bp | 515 bp | 010101-010109, 0103, 0105, 0107-0119 |
| 2 | 235 bp | 430 bp | 010201-010205 |
| 3 | 200 bp | 515 bp | 010101-010205, 0104-0114, 0116-0120 |
| 4 | 210 bp | 430 bp | 0103, 0110 |
| 5 | 230 bp | 430 bp | 0104, 0111 |
| 6 | 130 bp | 430 bp | 0105 |
| 7 | 210 bp | 430 bp | 0103, 0115, 0117 |
| 8 | 210 bp | 430 bp | 0107 |
| 95,6 | 110, 255 bp | 430 bp | 0104, 0106, 0108, 0120 |
| 10 | 210 bp | 430 bp | 0106, 0109, 0115 |
| 114,7 | 140, 210 bp | 430 bp | 0110, 0118 |
| 12 | 215 bp | 430 bp | 0112 |
| 135 | 85 bp | 430 bp | 0113 |
| 14 | 170 bp | 430 bp | 0114 |
| 15 | 215 bp | 430 bp | 0116 |
| 16 | 225 bp | 430 bp | 0119 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB*01 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*01 subtyping.

In addition, well number 3 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

4Primer mixes 1 and 11 have a tendency of giving rise to primer oligomer formation.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Primer mix 9: Specific PCR fragment of 110 bp in the DRB1*0108 allele. Specific PCR fragment of 255 bp in the DRB1*0104, 0106 and 0120 alleles.

7Primer mix 11: Specific PCR fragment of 140 bp in the DRB1*0118 allele. Specific PCR fragment of 210 bp in the DRB1*0110 allele.

1Primer mix 9: Specific PCR fragment of 110 bp in the DRB1*0108 allele. Specific PCR fragment of 255 bp in the DRB1*0104, 0106 and 0120 alleles.

Primer mix 11: Specific PCR fragment of 140 bp in the DRB1*0118 allele. Specific PCR fragment of 210 bp in the DRB1*0110 allele.