PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*01

Prod. No:	101.111-24/24u

101.111-06/06u

Lot No:    	90G

Expiry Date:    	2012-March-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 19 primer mixes used for DRB1*01 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB1*013 alleles |
| --- | --- | --- | --- |
| 15 | 255 bp | 515 bp | *010101-010117, 0103, 0105, 0107-0119, 0121, 0122w, 0124, 0125, 0127-0131 |
| 2 | 235 bp | 430 bp | *010201-010205, 0123, 0126 |
| 3 | 200 bp | 515 bp | *010101-010205, 0104-0114, 0116-0121, 0122w, 0125-0131 |
| 4 | 210 bp | 430 bp | *0103, 0110 |
| 5 | 220 bp | 430 bp | *0104, 0111, 0129 |
| 66,7 | 135 bp, 210 bp | 430 bp | *0105, 0129, 0131 |
| 7 | 210 bp | 430 bp | *0103, 0115, 0117 |
| 88 | 175 bp, 210 bp | 430 bp | *0107, 0127 |
| 94,9 | 110 bp, 255 bp | 430 bp | *0104, 0106, 0108, 0120 |
| 10 | 210 bp | 430 bp | *0106, 0109, 0115 |
| 115,7,10 | 140 bp, 210 bp | 430 bp | *0110, 0118 |
| 1211 | 140 bp, 180 bp, 215 bp | 430 bp | *0112, 0126, 0131 |
| 134,12 | 85 bp, 150 bp, 215 bp | 430 bp | *0113, 0123, 0128 |
| 1413 | 170 bp, 210 bp | 430 bp | *0114, 0123, 0124 |
| 1514 | 150 bp, 220 bp | 430 bp | *0116, 0121, 0128 |
| 1615 | 200 bp, 230 bp | 430 bp | *0119, 0121, 0124 |
| 1716 | 230 bp, 260 bp | 430 bp | *0122, 0127 |
| 18 | 230 bp | 430 bp | *0125 |
| 194 | 125 bp | 430 bp | *0130 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB*01 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*01 subtyping.

In addition, well number 3 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

4Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

5Primer mixes 1 and 11 have a tendency of giving rise to primer oligomer formation.

6Primer mix 6 has a tendency of giving rise to non-specific amplification.

7Primer mix 6: Specific PCR fragment of 135 bp in the DRB1* 0105 and DRB1*0131 alleles. Specific PCR fragment of 210 bp in the DRB1* 0129 allele.

8Primer mix 8: Specific PCR fragment of 175 bp in the DRB1*0127 allele. Specific PCR fragment of 210 bp in the DRB1*0107 allele.

9Primer mix 9: Specific PCR fragment of 110 bp in the DRB1*0108 allele. Specific PCR fragment of 255 bp in the DRB1*0104, 0106 and 0120 alleles.

10Primer mix 11: Specific PCR fragment of 140 bp in the DRB1*0118 allele. Specific PCR fragment of 210 bp in the DRB1*0110 allele.

11Primer mix 12: Specific PCR fragment of 140 bp in the DRB1*0131 allele. Specific PCR fragment of 180 bp in the DRB1*0126 allele. Specific PCR fragment of 215 bp in the DRB1*0112 allele.

12Primer mix 13: Specific PCR fragment of 85 bp in the DRB1*0113 allele. Specific PCR fragment of 150 bp in the DRB1*0128 allele. Specific PCR fragment of 215 bp in the DRB1*0123 allele.

13Primer mix 14: Specific PCR fragment of 170 bp in the DRB1*0114 allele. Specific PCR fragment of 210 bp in the DRB1* 0123 and DRB1*0124 alleles.

14Primer mix 15: Specific PCR fragment of 150 bp in the DRB1*0128 allele. Specific PCR fragment of 220 bp in the DRB1*0116 and DRB1*0121 alleles.

15Primer mix 16: Specific PCR fragment of 200 bp in the DRB1*0124 allele. Specific PCR fragment of 230 bp in the DRB1* 0119 and DRB1*0121 alleles.

16Primer mix 17: Specific PCR fragment of 230 bp in the DRB1*0127 allele. Specific PCR fragment of 260 bp in the DRB1*0122 allele.

’w’, might be weakly amplified.

1Primer mix 6: Specific PCR fragment of 135 bp in the DRB1* 0105 and DRB1*0131 alleles. Specific PCR fragment of 210 bp in the DRB1* 0129 allele.

Primer mix 8: Specific PCR fragment of 175 bp in the DRB1*0127 allele. Specific PCR fragment of 210 bp in the DRB1*0107 allele.

Primer mix 9: Specific PCR fragment of 110 bp in the DRB1*0108 allele. Specific PCR fragment of 255 bp in the DRB1*0104, 0106 and 0120 alleles.

Primer mix 11: Specific PCR fragment of 140 bp in the DRB1*0118 allele. Specific PCR fragment of 210 bp in the DRB1*0110 allele.

Primer mix 12: Specific PCR fragment of 140 bp in the DRB1*0131 allele. Specific PCR fragment of 180 bp in the DRB1*0126 allele. Specific PCR fragment of 215 bp in the DRB1*0112 allele.

Primer mix 13: Specific PCR fragment of 85 bp in the DRB1*0113 allele. Specific PCR fragment of 150 bp in the DRB1*0128 allele. Specific PCR fragment of 215 bp in the DRB1*0123 allele.

Primer mix 14: Specific PCR fragment of 170 bp in the DRB1*0114 allele. Specific PCR fragment of 210 bp in the DRB1* 0123 and DRB1*0124 alleles.

Primer mix 15: Specific PCR fragment of 150 bp in the DRB1*0128 allele. Specific PCR fragment of 220 bp in the DRB1*0116 and DRB1*0121 alleles.

Primer mix 16: Specific PCR fragment of 200 bp in the DRB1*0124 allele. Specific PCR fragment of 230 bp in the DRB1* 0119 and DRB1*0121 alleles.

Primer mix 17: Specific PCR fragment of 230 bp in the DRB1*0127 allele. Specific PCR fragment of 260 bp in the DRB1*0122 allele.

’w’, might be weakly amplified.