Gel Documentation Form and Worksheet

DRB1*01 (101.111-24/06 -24u/06u) 	Lot No: 85V	Expiry Date: 2016-November-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Interpretation:_______________

Failed lanes:    _________

Comments:    ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 1 has a tendency of giving rise to primer oligomer formation.

Primer mix 15 may give rise to a lower yield of HLA-specific PCR product than the other DRB1*01 resolution primer mixes in the DRB1*01:21 allele.

Primer mix 18 may have tendency of unspecific amplification and may give rise to a lower yield of HLA-specific PCR product than the other DRB1*01 primer mixes.

Primer mix 24 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1DRB1*01 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.15.0, January 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 6: Specific PCR fragment of 135 bp in the DRB1*01:05 and DRB1*01:31 alleles. Specific PCR fragment of 210 bp in the DRB1*01:29:01-01:29:02 alleles.

Primer mix 8: Specific PCR fragment of 115 bp in the DRB1*01:50. Specific PCR fragment of 175 bp in the DRB1*01:27 and the DRB1*14:112 alleles. Specific PCR fragment of 210 bp in the DRB1*01:07 allele.

Primer mix 9: Specific PCR fragment of 110 bp in the DRB1*01:08 allele. Specific PCR fragment of 255 bp in the DRB1*01:04, 01:06 and 01:20 alleles.

Primer mix 10: Specific PCR fragment of 210 bp in the DRB1*01:06, 01:09 and 01:15 alleles. Specific PCR fragment of 250 bp in the DRB1*01:52N allele.

Primer mix 11: Specific PCR fragment of 140 bp in the DRB1*01:18 and 01:44 alleles. Specific PCR fragment of 210 bp in the DRB1*01:10 allele.

Primer mix 12: Specific PCR fragment of 140 bp in the DRB1*01:31 allele. Specific PCR fragment of 180 bp in the DRB1*01:26 and 01:37 alleles. Specific PCR fragment of 215 bp in the DRB1*01:12 allele.

Primer mix 13: Specific PCR fragment of 85 bp in the DRB1*01:13 allele. Specific PCR fragment of 150 bp in the DRB1*01:28 allele. Specific PCR fragment of 215 bp in the DRB1*01:23 and *01:51 alleles.

Primer mix 14: Specific PCR fragment of 170 bp in the DRB1*01:14 allele. Specific PCR fragment of 210 bp in the DRB1*01:23, 01:24:01-01:24:02 and *01:51 alleles.

Primer mix 15: Specific PCR fragment of 150 bp in the DRB1*01:28 allele. Specific PCR fragment of 225 bp in the DRB1*01:16, 01:21 and 01:55 alleles.

Primer mix 16: Specific PCR fragment of 200 bp in the DRB1*01:24:01-01:24:02 and 01:54 alleles. Specific PCR fragment of 230 bp in the DRB1* 01:19 and 01:21 alleles.

Primer mix 17: Specific PCR fragment of 125 bp in the DRB1*01:30 and the DRB1*13:165 alleles. Specific PCR fragment of 230 bp in the DRB1*01:27 allele. Specific PCR fragment of 260 bp in the DRB1*01:22 allele.

Primer mix 18: Specific PCR fragment of 115 bp in the DRB1*01:50 allele. Specific PCR fragment of 230 bp in the DRB1*01:25 allele.

Primer mix 21: Specific PCR fragment of 85 bp in the DRB1*01:34 allele. Specific PCR fragment of 250 bp in the DRB1*01:32 allele.

Primer mix 22: Specific PCR fragment of 205 bp in the DRB1*01:36 allele. Specific PCR fragment of 250 bp in the DRB1*01:52N allele.

’w’, might be weakly amplified.

The DRB1*01 primer set cannot distinguish the silent mutations in DRB1*01:01:01-01:01:26 alleles, the DRB1*01:02:01-01:02:05 and 01:02:07-01:02:09 alleles, the DRB1*01:11:01-01:11:02 or the DRB1*01:24:01-01:24:02 alleles.