PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*03

Prod. No:	101.113-24/04 – including Taq pol.

101.113-24u/04u – without Taq pol.

Lot No:    	36E

Expiry Date:    	2010-March-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 29 primer mixes used for DRB1*03 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB1*03 alleles3 | Other amplified DRB alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 220 bp | 515 bp | 030101-030106, 0304, 0306, 0308, 0310-0313, 0315, 0318-0320, 0322, 0323, 0325, 0326, 0328, 0330-0334, 0336, 0337, 0339 | 0906, 1159, 1327, 1371 |
| 25 | 85 bp | 430 bp | 030201-0303, 0327, 0329, 0338 | 1315, 1319, 1326, 1353, 1357, 1402-140302, 1406, 1412, 1413, 1418-1420, 1424, 1427, 1429, 1440, 1441, 1447-1449, 1451, 1463, 1467 |
| 36 | 190 bp | 515 bp | 030201-030202, 030501-030502, 0309, 0314, 0327, 0329, 0338 | 1109, 1120, 130201-130203, 130501-130502, 1326, 1329, 1331, 1334, 1336, 1339, 1341, 135002, 1356, 1363, 1365, 1367, 1373, 1374, 1402-140302, 1409, 1413, 1419, 1424, 1427, 1430, 1446-1448, 1451, 1463, 1467, 1608, DRB3*0108, DRB3*0206, DRB3*0220 |
| 4 | 190 bp | 430 bp | 030101-030106, 0303, 0306-0308, 0310-0313, 0315, 0318-0320, 0322-0324, 0326, 0328, 0332-0334, 0336, 0337, 0339 | 1116, 1140, 1158, 1159, 130101-130103, 1306, 1309, 1310, 1315, 1318, 1320, 1327, 1328, 1332, 1335, 1340, 1342, 1343, 1351, 1353, 1359, 1361, 1369, 1371, 1377-1380, 1406, 1412, 1417, 1418, 1421, 1429, 1433, 1452, 1459, 1464 |
| 55 | 75 bp | 430 bp | 0304 |  |
| 65 | 105 bp | 430 bp | 030101-030106, 0304-030502, 0308-0315, 0316weakly, 0318-0320, 0322, 0323, 0325, 0328, 0330, 0332-0334, 0336, 0337, 0339 | 1159, 1327, 1341, 1371 |
| 75 | 135 bp | 430 bp | 0308 | 1159 |
| 87 | 185 bp | 430 bp | 030201-0303, 0307, 0321, 0329, 0338 | 0116, 0422, 1107, 1525, DRB3*0109, DRB3*0204, DRB3*0219, DRB3*0222, DRB3*0303 |
| 95 | 135 bp | 430 bp | 030101-030106, 0304-0306, 0309, 0311, 0313weakly, 0314-0316, 0318-0320, 0322, 0323, 0325, 0326, 0328, 0330-0334, 0336, 0337, 0339 | 1327, 1341, 1371, DRB3*0107 |
| 105 | 145 bp | 430 bp | 030201-0303, 0306, 0317, 0326, 0327, 0331, 0335, 0338 | 0820, 1117, 1137, 1141, 1148, 1150, 1152, 1156, 130301-130302, 130701-1308, 1312, 1313, 1319, 1326, 1332, 1333, 1336-1338, 1340, 1347-1349, 1353, 1355, 1358, 1360, 1365, 1370, 1372, 1376, 1381, 140101-140301, 140501-1409, 1412-1414, 1416, 1418-1420, 1422-1427, 1429, 143201-143202, 1434, 1436-1441, 1443-1445, 1447-1449, 1451, 1454-1456, 1458-1460, 1462, 1463, 1467, 1469, 1470, 1474, 1475 |
| 115,8 | 105, 180 bp | 515 bp | 0309, 0313, 0323, 0338 | 1339, 1354, 1377-1379, 1448, 1464 |
| 125 | 135 bp | 430 bp | 0310 |  |
| 13 | 225 bp | 430 bp | 0311, 0317, 0324, 0327, 0335 | 1438 |
| 145,9 | 105, 135 bp | 515 bp | 0312, 0316 |  |
| 15 | 155 bp | 430 bp | 0317, 0337 | 1139, 1143, 1150, 1459 |
| 16 | 195 bp | 430 bp | 0314, 0315 | 0905, 1159, 1327, 1341, 1371 |
| 17 | 255 bp | 515 bp | 0320 | 110601-110602, 1121, 1147, 1157, 1358, 1381, 1429 |
| 185 | 120 bp | 430 bp | 030101-030106, 0304-030502, 0307-0310, 0312-0316, 0318-0323, 0325, 0328-0330, 0332-0334, 0336, 0337, 0339 | 1107, 1525 |
| 196 | 175 bp | 515 bp | 0322 |  |
| 205,10 | 75, 120 bp | 430 bp | 0304weakly, 0325, 0328, 0331 |  |
| 215 | 70 bp | 430 bp | 0321, 0330 | 1110, 111201-111202weakly, 1113, 1205, 1214, 1215, 1357, 1364, 1435, 1465, 1472 |
| 225 | 95 bp | 430 bp | 0318 |  |
| 235,11 | 65, 165 bp | 430 bp | 0309, 0319, 0323 |  |
| 245,7 | 65 bp | 430 bp | 030201-0303, 0327, 0329, 0338 | 1315, 1326, 1353, 1402-140302, 1406, 1412, 1413, 1418, 1419, 1424, 1427, 1429, 1446-1448, 1451, 1452, 1463, 1467, DRB3*0220 |
| 25 | 180 bp | 430 bp | 0332 |  |
| 265 | 115 bp | 515 bp | 0326, 0333 | 1144, 1149, 1150 |
| 27 | 230 bp | 515 bp | 0334 |  |
| 2812 | 160 bp | 430 bp | 0336 |  |
| 29 | 180 bp | 430 bp | 0339 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*03SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*03 subtyping. In addition, wells number 3, 11, 14, 17, 19, 26 and 27 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, several DRB1*11, several DRB1*12 as well as many DRB1*13 and DRB1*14 alleles are amplified by some of the DRB1*03 primer mixes. In addition, the DRB1*0116 allele, the DRB1*0422 allele, the DRB1*0820 allele, the DRB1*0905 and 0906 alleles, the DRB1*1525 allele and DRB1*1608 allele are amplified by a few of the DRB1*03 primer mixes. Finally, the DRB3*0107 allele is amplified by primer mix 9, the DRB3*0108 and DRB3*0206 alleles are amplified by primer mix 3, the DRB3*0109, DRB3*0204, DRB3*0219, DRB3*0222 and DRB3*0303 alleles by primer mix 8 and the DRB1*0220 allele by primer mixes 3 and 24.

5Specific PCR fragments shorter than 150 base pairs have a lower intensity than longer PCR bands.

6Primer mixes 3 and 19 have a tendency of giving rise to primer oligomer formation. Most pronounced for primer mix 19.

7The specific primer pairs in primer mixes 8 and 24 yield less product than than the other DRB1*03 primer mixes.

8Primer mix 11: Specific PCR fragment of 105 bp in the DRB1*0309 and 0323 alleles. Specific PCR fragment of 180 bp in the DRB1*0313 and 0338 and the 1339, 1354, 1377-1379, 1448 and 1464 alleles.

9Primer mix 14: Specific PCR fragment of 105 bp in the DRB1*0316 allele. Specific PCR fragment of 135 bp in the DRB1*0312 allele.

10Primer mix 20: Specific PCR fragment of 75 bp in the DRB1*0304weakly, 0325 and 0331 alleles. Specific PCR fragment of 120 bp in the DRB1*0328 allele.

11Primer mix 23: Specific PCR fragment of 65 bp in the DRB1*0309 and 0323 alleles. Specific PCR fragment of 165 bp in the DRB1*0319 allele.

12Primer mix 28 may give to nonspecific amplifications.

1Primer mix 11: Specific PCR fragment of 105 bp in the DRB1*0309 and 0323 alleles. Specific PCR fragment of 180 bp in the DRB1*0313 and 0338 and the 1339, 1354, 1377-1379, 1448 and 1464 alleles.

Primer mix 14: Specific PCR fragment of 105 bp in the DRB1*0316 allele. Specific PCR fragment of 135 bp in the DRB1*0312 allele.

10Primer mix 20: Specific PCR fragment of 75 bp in the DRB1*0304weakly, 0325 and 0331 alleles. Specific PCR fragment of 120 bp in the DRB1*0328 allele.

Primer mix 23: Specific PCR fragment of 65 bp in the DRB1*0309 and 0323 alleles. Specific PCR fragment of 165 bp in the DRB1*0319 allele.

“w”, may be weakly amplified.

Change in revision R01 compared to R00:

- Footnote has been to the Specificity Table for primer mix 28.