PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*03

Prod. No:	101.113-24/24u

101.113-04/04u

Lot No:    	15S

Expiry Date:    	2015-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for DRB1*03 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB1*03 alleles3 | Other amplified DRB alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 220 bp | 515 bp | *03:01:01:01-03:01:21, 03:04:01-03:04:02, 03:06, 03:08, 03:10-03:13:02, 03:15, 03:16?, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:34, 03:36-03:37, 03:39, 03:42-03:48, 03:50-03:52, 03:55-03:56, 03:58-03:62, 03:64-03:73, 03:75-03:84, 03:86-03:87, 03:89 | *09:06, 11:59, 11:83, 11:135, 13:27, 13:71, 13:129, 13:144, 14:82, 14:95, 14:132 |
| 25,8 | 85 bp | 430 bp | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 03:74, 03:88 | *13:15, 13:19, 13:26:01-13:26:02, 13:53, 13:57, 13:85-13:86, 13:104, 14:02-14:03:02, 14:06:01-14:06:02, 14:12:01-14:13, 14:18-14:20, 14:24, 14:27, 14:29, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:67, 14:77-14:78, 14:81, 14:83, 14:85, 14:89, 14:94, 14:98, 14:102, 14:106, 14:108-14:109, 14:115, 14:135 |
| 3 | 190 bp | 515 bp | *03:02:01-03:02:02, 03:05:01-03:05:03, 03:09, 03:14, 03:27, 03:29, 03:38, 03:53, 03:74, 03:81, 03:88 | *11:09, 11:20, 11:28:02, 11:87, 11:122, 13:02:01-13:02:03, 13:02:05-13:02:08, 13:05:01-13:05:02, 13:26:01-13:26:02, 13:29, 13:31, 13:34, 13:36, 13:39, 13:41, 13:50:02, 13:56, 13:63, 13:65, 13:67, 13:73-13:74, 13:85, 13:96:01-13:96:02, 13:99, 13:103-13:104, 13:119-13:120, 13:123-13:124, 13:126, 13:128, 13:136, 13:139-13:140, 13:142N-13:143, 13:145, 13:147, 13:155, 13:157-13:158, 14:02-14:03:02, 14:09, 14:13, 14:19, 14:24, 14:27, 14:30, 14:46-14:48, 14:51, 14:63, 14:67, 14:85, 14:89, 14:98, 14:109, 14:115, 14:134-14:135, 16:08, DRB3*01:08, DRB3*02:06, DRB3*02:20 |
| 4 | 190 bp | 430 bp | *03:01:01:01-03:01:21, 03:03, 03:06-03:08, 03:10-03:13:02, 03:15, 03:16?, 03:18-03:20, 03:22-03:24, 03:26, 03:28, 03:32-03:34, 03:36-03:37, 03:39, 03:42, 03:45-03:48, 03:50-03:51, 03:54-03:56, 03:58-03:62, 03:64-03:73, 03:76-03:84, 03:86-03:87, 03:89 | *11:16, 11:40, 11:58:01, 11:59, 11:83, 11:113, 11:115, 13:01:01-13:01:13, 13:06, 13:09-13:10, 13:15, 13:18, 13:20, 13:27-13:28, 13:32, 13:35, 13:40, 13:42-13:43, 13:51, 13:53, 13:59, 13:61:01-13:61:02, 13:69, 13:71, 13:77-13:80, 13:87, 13:91-13:93, 13:98, 13:102, 13:105, 13:110-13:114, 13:117, 13:121, 13:125, 13:127, 13:129-13:131, 13:137N-13:138, 13:141, 13:144, 13:146, 13:148, 13:153, 13:156, 14:06:01-14:06:02, 14:12:01-14:12:02, 14:17-14:18, 14:21, 14:29, 14:33, 14:52, 14:59, 14:64, 14:78, 14:80-14:81, 14:83, 14:95, 14:106, 14:108, 14:121, 14:132-14:133, 15:32 |
| 55 | 75 bp | 430 bp | *03:04:01-03:04:02 |  |
| 65 | 105 bp | 430 bp | *03:01:01:01-03:01:21, 03:04:01-03:05:03, 03:08-03:15, 03:16w, 03:18-03:20, 03:22-03:23, 03:25, 03:28, 03:30, 03:32-03:34, 03:36-03:37, 03:39, 03:42-03:48, 03:50-03:52, 03:55-03:59, 03:61-03:73, 03:76-03:79, 03:81, 03:83-03:84, 03:86-03:87, 03:89 | *11:59, 11:80, 11:87, 11:135, 11:142, 13:27, 13:41, 13:71, 13:129, 13:144, 14:95, 14:132 |
| 7 | 130 bp | 430 bp | *03:08, 03:65, 03:72 | *11:59, 11:80, 11:83, 11:87, 11:135, 11:142 |
| 8 | 185 bp | 430 bp | *03:02:01-03:03, 03:07, 03:21, 03:29, 03:38, 03:40-03:41, 03:49, 03:53, 03:74, 03:88 | *01:16, 01:55, 04:22, 11:07, 11:103, 11:107, 11:125, 14:76, 14:79, 15:25, DRB3*01:09, DRB3*02:04, DRB3*02:19, DRB3*02:22, DRB3*03:03 |
| 9 | 135 bp | 430 bp | *03:01:01:01-03:01:21, 03:04:01-03:06, 03:09, 03:11:01, 03:13:01w-03:13:02w, 03:14-03:16, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:34, 03:36-03:37, 03:39, 03:42-03:45, 03:47-03:48, 03:50-03:52, 03:55-03:63, 03:66, 03:67Nw, 03:68N-03:73, 03:75-03:78, 03:80-03:84, 03:86-03:87, 03:89 | *09:01:08, 09:02:01-09:02:02, 13:27, 13:41, 13:71, 13:129, 13:144, 14:95, DRB3*01:07, DRB3*01:15 |
| 10 | 145 bp | 430 bp | *03:02:01-03:03, 03:06, 03:17, 03:26-03:27, 03:31, 03:35, 03:38, 03:41, 03:60, 03:74-03:75, 03:80, 03:88 | *08:20, 11:17, 11:37:01-11:37:02, 11:41, 11:48, 11:50, 11:52, 11:56, 11:83, 11:107, 13:03:01-13:03:06, 13:07:01-13:08, 13:12:01-13:13, 13:19, 13:26:01-13:26:02, 13:32-13:33:03, 13:36-13:38, 13:40, 13:47-13:49, 13:53, 13:55, 13:58, 13:60, 13:65, 13:70, 13:72, 13:76, 13:81, 13:84-13:85, 13:88-13:90, 13:94:01-13:95, 13:101, 13:108, 13:115, 13:118, 13:120, 13:122, 13:133-13:135, 13:149, 13:151-13:152, 14:01:01-14:03:01, 14:05:01-14:09, 14:12:01-14:14, 14:16, 14:18-14:20, 14:22-14:27, 14:29, 14:32:01-14:32:02, 14:34, 14:36-14:41, 14:43-14:45, 14:47-14:49, 14:51, 14:54:01-14:56, 14:58-14:60, 14:62-14:63, 14:67, 14:69-14:70, 14:74-14:75, 14:77-14:78, 14:80-14:92N, 14:94, 14:96-14:106, 14:108-14:119, 14:121-14:125, 14:127, 14:129-14:131, 14:133, 14:135-14:137N, 14:139-14:140, DRB3*01:14 |
| 115,9 | 105 bp, 180 bp | 515 bp | *03:09, 03:13:01-03:13:02, 03:23, 03:38, 03:79, 03:88 | *11:116, 11:120, 13:39, 13:54, 13:77-13:79, 13:100, 14:48, 14:64, 14:109, 14:115, DRB3*01:14 |
| 1210 | 135 bp, 190 bp | 430 bp | *03:10, 03:51, 03:76 | *11:59w, 11:80w, 11:83w, 11:87w, 11:135w, 11:142w, 13:27w, 13:41w, 13:71w, 13:129w, 14:82, 14:132 |
| 1311 | 165 bp, 225 bp | 430 bp | *03:11:01, 03:17, 03:24, 03:27, 03:35, 03:47, 03:62, 03:81 | *11:136, 14:38, 14:127 |
| 145,6,12 | 105 bp, 135 bp, 245 bp | 515 bp | *03:12, 03:16, 03:68N | *09:08 |
| 15 | 155 bp | 430 bp | *03:17, 03:37, 03:62, 03:80 | *11:39, 11:43, 11:50, 11:115, 13:118, 13:131, 13:135, 13:150, 14:59, 14:86, 14:96 |
| 16 | 195 bp | 430 bp | *03:14-03:15 | *09:05, 11:59, 11:80, 11:83, 11:87, 11:135, 11:142, 13:27, 13:41, 13:71, 13:129, 13:144 |
| 17 | 255 bp | 515 bp | *03:20 | *11:06:01-11:06:03, 11:21, 11:47, 11:57, 11:129, 13:58, 13:81, 14:29, 14:99 |
| 185 | 120 bp | 430 bp | *03:01:01:01-03:01:21, 03:04:01-03:05:03, 03:07-03:10, 03:12-03:16, 03:18-03:23, 03:25, 03:28-03:30, 03:32-03:34, 03:36-03:37, 03:39-03:40, 03:42-03:52, 03:54-03:59, 03:61-03:73, 03:78-03:79, 03:82-03:83, 03:86-03:87 | *11:07, 11:103, 11:105, 11:125, 15:25 |
| 19 | 175 bp | 515 bp | *03:22 |  |
| 205,13 | 75 bp, 120 bp, 245 bp | 430 bp | *03:25, 03:28, 03:31, 03:68N, 03:75 | *11:80, 11:135, 11:142 |
| 215,14 | 70 bp, 150 bp | 430 bp | *03:21, 03:30, 03:42-03:43, 03:87 | *11:10:01-11:10:02, 11:12:01-11:13:02, 11:30, 11:114, 12:05, 12:14-12:15, 13:57, 13:64, 13:67, 13:83, 13:107, 14:35, 14:65, 14:72, 14:138 |
| 225 | 95 bp | 430 bp | *03:18 |  |
| 235,15 | 65 bp, 165 bp | 430 bp | *03:09, 03:19, 03:23 |  |
| 245 | 65 bp | 430 bp | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 03:74, 03:88 | *13:15, 13:26:01-13:26:02, 13:53, 13:85, 13:104, 14:02-14:03:02, 14:06:01-14:06:02, 14:12:01-14:13, 14:18-14:19, 14:24, 14:27, 14:29, 14:46-14:48, 14:51-14:52, 14:63, 14:67, 14:78, 14:81, 14:85, 14:89, 14:98, 14:106, 14:108-14:109, 14:115, 14:121, 14:135, DRB3*02:20 |
| 255,16 | 75 bp, 180 bp, 230 bp | 430 bp | *03:32, 03:67N, 03:69 |  |
| 265,7 | 115 bp | 515 bp | *03:26, 03:33, 03:43-03:44, 03:71 | *11:44, 11:49:01-11:50, 11:76, 11:85, 11:114, 11:127, 13:34, 13:64, 13:106, 13:136, 14:41, 14:77, 14:97, 14:110, DRB3*01:14 |
| 275,17 | 115 bp, 150 bp, 175 bp, 230 bp | 515 bp | *03:34, 03:43, 03:46, 03:53, 03:71 | *11:76, 11:114, 13:34, 13:64, 13:136, 14:41, 14:77, 14:110, DRB3*01:14 |
| 285,18 | 75 bp, 160 bp, 195 bp | 430 bp | *03:36, 03:45, 03:67N |  |
| 2919 | 155 bp, 180 bp | 430 bp | *03:39-03:41, 03:52, 03:75 | *11:07, 11:103, 11:105, 11:107, 11:125 |
| 30 | 175 bp | 430 bp | *03:01:01:01-03:07, 03:09, 03:11:01-03:41, 03:43-03:45, 03:47-03:63, 03:66-03:68N, 03:70-03:84, 03:86, 03:88-03:89 | *08:20, 13:01:01-13:16, 13:18-13:42, 13:44, 13:46-13:66:02, 13:68-13:102, 13:104-13:115, 13:117-13:121, 13:123-13:158, 14:02-14:03:02, 14:05:01-14:06:02, 14:09, 14:12:01-14:14, 14:17-14:21, 14:23:01, 14:23:03-14:24, 14:27, 14:29-14:30, 14:33, 14:36-14:37, 14:40-14:45, 14:47-14:48, 14:51, 14:56, 14:59, 14:63-14:65, 14:67, 14:77-14:78, 14:80-14:81, 14:83-14:85, 14:89, 14:91, 14:94-14:96, 14:98, 14:100, 14:102-14:103, 14:106, 14:108-14:109, 14:115-14:116, 14:121, 14:123, 14:127, 14:134-14:136 |
| 3120 | 195 bp, 245 bp | 515 bp | *03:45, 03:50 |  |
| 32 | 215 bp | 430 bp | *03:48 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*03 SSP subtypings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*03 subtyping.

In addition, wells number 3, 11, 14, 17, 19, 26, 27 and 31 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB1 alleles 1st and/or 3rd exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*01:01 consensus sequence.

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, non-DRB1*03 alleles are amplified by primer mixes 1 to 4, 6 to 18, 20, 21, 24, 26, 27, 29 and 30. In addition, primer mix 3 will amplify the DRB3*01:08, DRB3*02:06 and DRB3*02:20 alleles, primer mix 8 will amplify the DRB3*01:09, DRB3*02:04, DRB3*02:19, DRB3*02:22 and DRB3*03:03 alleles, primer mix 9 will amplify the DRB3*01:07 and  DRB3*01:15 allele, primer mix 24 will amplify the DRB3*02:20 allele and the DRB3*01:14 allele will be amplified by primer mixes 10, 11 26 and 27.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mix 14 may have a tendency of giving rise to primer oligomer formation.

7Primer mix 26 may have a tendency of unspecific amplification.

8Primer mix 2 may give rise to a lower yield of specific PCR product than the other DRB1*03 primer mixes.

9Primer mix 11: Specific PCR fragment of 105 bp in the DRB1*03:09 and 03:23 alleles. Specific PCR fragment of 180 bp in the DRB1*03:13:01-03:13:02, 03:38, 03:79 and 03:88 and in the DRB1*11:116, 11:120, 13:39, 13:54, 13:77-13:79, 13:100, 14:48, 14:64, 14:109 and 14:115 and in the DRB3*01:14 alleles.

10Primer mix 12: Specific PCR fragment of 135 bp in the DRB1*03:10 and the DRB1*14:82 and 14:132 alleles. Specific PCR fragment of 190 bp in the DRB1*03:51 and 03:76 and the DRB1*11:59w, 11:80w, 11:83w, 11:87w, 11:135w, 11:142w, 13:27w, 13:41w, 13:71w and 13:129w alleles.

11Primer mix 13: Specific PCR fragment of 165 bp in the DRB1*03:47 and 03:62 alleles. Specific PCR fragment of 225 bp in the DRB1 *03:11, 03:17, 03:24, 03:27, 03:35 and 03:81 and in the DRB1*11:136, 14:38 and 14:127 alleles.

12Primer mix 14: Specific PCR fragment of 105 bp in the DRB1*03:16 allele. Specific PCR fragment of 135 bp in the DRB1*03:12 and the DRB1*09:08 alleles. Specific PCR fragment of 245 bp in the DRB1*03:68N allele.

13Primer mix 20: Specific PCR fragment of 75 bp in the DRB1*03:25, 03:31 and 03:75 and the DRB1* 11:80, 11:135 and 11:142 alleles. Specific PCR fragment of 120 bp in the DRB1*03:28 allele. Specific PCR fragment of 245 bp in the DRB1*03:68N allele.

14Primer mix 21: Specific PCR fragment of 70 bp in the DRB1*03:21, 03:30 and 03:43 and the DRB1*11:10:01-11:10:02, 11:12:01-11:13:02, 11:114, 12:05, 12:14-12:15, 13:57, 13:64, 13:83, 13:107, 14:35, 14:65, 14:72 and 14:138 alleles. Specific PCR fragment of 150 bp in the DRB1*03:42 and 03:87 and the DRB1*11:30 and 13:67 alleles.

15Primer mix 23: Specific PCR fragment of 65 bp in the DRB1*03:09 and 03:23 alleles. Specific PCR fragment of 165 bp in the DRB1*03:19 allele.

16Primer mix 25: Specific PCR fragment of 75 bp in the DRB1*03:67N allele. Specific PCR fragment of 180 bp in the DRB1*03:32 allele. Specific PCR fragment of 230 bp in the DRB1*03:69 allele.

17Primer mix 27: Specific PCR fragment of 115 bp in the DRB1*03:43, 03:71 and the DRB1*11:76, 11:114, 13:34, 13:64, 13:136, 14:41, 14:77 and 14:110 and the DRB3*01:14 alleles. Specific PCR fragment of 150 bp in the DRB1*03:53 allele. Specific PCR fragment of 175 bp in the DRB1*03:46 allele. Specific PCR fragment of 230 bp in the DRB1*03:34 allele.

18Primer mix 28: Specific PCR fragment of 75 bp in the DRB1*03:67N allele. Specific PCR fragment of 160 bp in the DRB1*03:36 allele. Specific PCR fragment of 195 bp in the DRB1*03:45 allele.

19Primer mix 29: Specific PCR fragment of 155 bp in the DRB1*03:40-03:41, 03:52 and 03:75 and the DRB1*11:07, 11:103, 11:105, 11:107 and 11:125 alleles. Specific PCR fragment of 180 bp in the DRB1*03:39 allele.

20Primer mix 31: Specific PCR fragment of 195 bp in the DRB1*03:45 allele. Specific PCR fragment of 245 bp in the DRB1*03:50 allele.

‘?’, nucleotide sequence information not available for the primer matching sequence.

“w”, may be weakly amplified.

1The DRB1*03:11:02 allele has been renamed DRB1*03:81.

2DRB1*03 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.11.0, January 2013.

3The DRB1*03 primer set cannot separate the DRB1*03:35 and DRB1*14:127. These alleles can be distinguished by the DR low resolution kit and/or the DRB1*14 high resolution kit.

4The DRB1*03 primer set cannot separate the DRB1*03:77, *03:84 and *03:89 and the DRB1*14:95 alleles. These alleles can be distinguished by the DR low resolution kit and/or the DRB1*14 high resolution kit.

5Primer mix 11: Specific PCR fragment of 105 bp in the DRB1*03:09 and 03:23 alleles. Specific PCR fragment of 180 bp in the DRB1*03:13:01-03:13:02, 03:38, 03:79 and 03:88 and in the DRB1*11:116, 11:120, 13:39, 13:54, 13:77-13:79, 13:100, 14:48, 14:64, 14:109 and 14:115 and in the DRB3*01:14 alleles.

Primer mix 12: Specific PCR fragment of 135 bp in the DRB1*03:10 and the DRB1*14:82 and 14:132 alleles. Specific PCR fragment of 190 bp in the DRB1*03:51 and 03:76 and the DRB1*11:59w, 11:80w, 11:83w, 11:87w, 11:135w, 11:142w, 13:27w, 13:41w, 13:71w and 13:129w alleles.

Primer mix 13: Specific PCR fragment of 165 bp in the DRB1*03:47 and 03:62 alleles. Specific PCR fragment of 225 bp in the DRB1 *03:11, 03:17, 03:24, 03:27, 03:35 and 03:81 and in the DRB1*11:136, 14:38 and 14:127 alleles.

Primer mix 14: Specific PCR fragment of 105 bp in the DRB1*03:16 allele. Specific PCR fragment of 135 bp in the DRB1*03:12 and the DRB1*09:08 alleles. Specific PCR fragment of 245 bp in the DRB1*03:68N allele.

Primer mix 20: Specific PCR fragment of 75 bp in the DRB1*03:25, 03:31 and 03:75 and the DRB1* 11:80, 11:135 and 11:142 alleles. Specific PCR fragment of 120 bp in the DRB1*03:28 allele. Specific PCR fragment of 245 bp in the DRB1*03:68N allele.

Primer mix  21: Specific PCR fragment of 70 bp in the DRB1*03:21, 03:30 and 03:43 and the DRB1*11:10:01-11:10:02, 11:12:01-11:13:02, 11:114, 12:05, 12:14-12:15, 13:57, 13:64, 13:83, 13:107, 14:35, 14:65, 14:72 and 14:138 alleles. Specific PCR fragment of 150 bp in the DRB1*03:42 and 03:87 and the DRB1*11:30 and 13:67 alleles.

Primer mix 23: Specific PCR fragment of 65 bp in the DRB1*03:09 and 03:23 alleles. Specific PCR fragment of 165 bp in the DRB1*03:19 allele.

Primer mix 25: Specific PCR fragment of 75 bp in the DRB1*03:67N allele. Specific PCR fragment of 180 bp in the DRB1*03:32 allele. Specific PCR fragment of 230 bp in the DRB1*03:69 allele.

Primer mix 27: Specific PCR fragment of 115 bp in the DRB1*03:43, 03:71 and the DRB1*11:76, 11:114, 13:34, 13:64, 13:136, 14:41, 14:77 and 14:110 and the DRB3*01:14 alleles. Specific PCR fragment of 150 bp in the DRB1*03:53 allele. Specific PCR fragment of 175 bp in the DRB1*03:46 allele. Specific PCR fragment of 230 bp in the DRB1*03:34 allele.

Primer mix 28: Specific PCR fragment of 75 bp in the DRB1*03:67N allele. Specific PCR fragment of 160 bp in the DRB1*03:36 allele. Specific PCR fragment of 195 bp in the DRB1*03:45 allele.

Primer mix 29: Specific PCR fragment of 155 bp in the DRB1*03:40-03:41, 03:52 and 03:75 and the DRB1*11:07, 11:103, 11:105, 11:107 and 11:125 alleles. Specific PCR fragment of 180 bp in the DRB1*03:39 allele.

Primer mix 31: Specific PCR fragment of 195 bp in the DRB1*03:45 allele. Specific PCR fragment of 245 bp in the DRB1*03:50 allele.

‘?’, nucleotide sequence information not available for the primer matching sequence.

“w”, may be weakly amplified.