Gel Documentation Form and Worksheet

DRB1*04 (101.114-24/03, -24u/03u) 	Lot No: 18V	Expiry Date: 2016-July-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Interpretation:_______________

Failed lanes:    _________

Comments:    ___________________

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 21, 22 and 33 may have tendencies of unspecific amplifications.

Primer mixes 14, 31 and 38 have a tendency to giving rise to primer oligomer formation.

Primer mixes 24 and 32 may give rise to a lower yield of HLA-specific PCR product than the other DRB1*04 primer mixes.

Primer mix 48 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1DRB1*04 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.14.0, October 2013.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 11: Specific PCR fragment of 175 bp in the DRB1*04:15, 04:61, 04:138, 04:145 and 04:154 and the DRB1*11:22 alleles. Specific PCR fragment of 220 bp in the DRB1*04:22 and 04:98:02 alleles.

Primer mix 12: Specific PCR fragment of 180 bp in the DRB1*04:16 allele. Specific PCR fragment of 265 bp in the DRB1*04:23, 04:52 and 04:151 alleles.

Primer mix 13: Specific PCR fragment of 170 bp in the DRB1*04:46 allele. Specific PCR fragments of 225 bp in the DRB1*04:12, 04:18, 04:25, 04:31, 04:54-04:55, 04:58, 04:86, 04:103 and 04:126alleles.

Primer mix 19: Specific PCR fragment of 140 bp in the DRB1*04:39 allele. Specific PCR fragment of 210 bp in the DRB1*04:24, 04:69, 04:78, 04:101 and the DRB1*14:10 and 14:143 alleles.

Primer mix 20: Specific PCR fragment of 100 bp in the DRB1*04:63, 04:110 and 04:161 alleles. Specific PCR fragment of 175 bp in the DRB1*04:40 allele. Specific PCR fragment of 255 bp in the DRB1*04:27 allele.

Primer mix 21: Specific PCR fragment of 155 bp in the DRB1*04:29 allele. Specific PCR fragment of 205 bp in the DRB1*04:37, 04:58, 04:73, 04:105:01-04:105:02, 04:122 and 04:132 alleles.

Primer mix 25: Specific PCR fragment of 125 bp in the DRB1*04:26 allele. Specific PCR fragment of 230 bp in the DRB1*04:45 allele. Specific PCR fragment of 260 bp in the DRB1*04:64 allele.

Primer mix 27: Specific PCR fragment of 120 bp in the DRB1*04:10:01-04:12, 04:60, 04:67, 04:91, 04:126, 04:136 and 04:147 and the DRB1*03:12, 08:06, 08:10, 08:12, 08:22, 08:54, 13:04, 13:32, 13:48, 13:58, 13:75, 13:81, 13:89, 13:93-13:94:02, 13:108, 13:164, 14:65, 14:78 and 15:12 alleles. Specific PCR fragment of 265 bp in the DRB1*04:49 allele.

Primer mix 29: Specific PCR fragment of 75 bp in the DRB1*04:07:01-04:07:04, 04:17:01-04:17:02, 04:20, 04:69, 04:74, 04:92, 04:102, 04:129 and 04:160 and the DRB1*01:17, 09:01:02-09:01:05, 09:01:08-09:02:02, 09:04-09:05, 09:07-09:08, 09:10-09:20, 09:22, 14:07:01-14:07:02, 14:14, 14:36, 14:42, 14:44:01-14:44:03, 14:51, 14:68, 14:93, 14:134, 14:143 and in the DRB5*01:12 alleles. Specific PCR fragment of 260 bp in the DRB1*04:64 allele.

Primer mix 31: Specific PCR fragment of 85 bp in the DRB1*04:14, 04:106 and 04:132 and the DRB1*01:03, 01:39N, 08:40, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:20, 11:68, 11:73, 11:93, 11:122, 13:02:01-13:02:03, 13:02:05-13:03:06, 13:23:01-13:23:02, 13:29, 13:31, 13:33:01-13:34, 13:36-13:39, 13:41, 13:45, 13:63, 13:65-13:67, 13:73-13:74, 13:85, 13:88, 13:90, 13:95-13:97:02, 13:99, 13:101, 13:103-13:104, 13:107, 13:115, 13:120, 13:122-13:124, 13:126, 13:128, 13:135, 13:139, 13:143, 13:145, 13:147, 13:151-13:152, 13:155 and 13:159 alleles. Specific PCR fragment of 260 bp in the DRB1*04:64 allele.

Primer mix 35: Specific PCR fragment of 210 bp in the DRB1*04:81N allele. Specific PCR fragment of 245 bp in the DRB1*04:51 and 04:70 alleles.

Primer mix 36: Specific PCR fragment of 175 bp in the DRB1*04:96 allele. Specific PCR fragment of 240 bp in the DRB1*04:71 allele.

Primer mix 38: Specific PCR fragment of 140 bp in the DRB1*04:89 allele. Specific PCR fragment of 215 bp in the DRB1*04:73, 04:75, 04:105:01-04:105:02 and 04:122 alleles.

Primer mix 39: Specific PCR fragment of 100 bp in the DRB1*04:83 allele. Specific PCR fragment of 175 bp in the DRB1*04:15, 04:61-04:62, 04:69, 04:138, 04:145-04:146 and 04:154 and the DRB1*11:22, 14:10, 14:57 and 14:143 alleles.

Primer mix 40: Specific PCR fragment of 145 bp in the DRB1*04:30 and 04:85 alleles. Specific PCR fragment of 200 bp in the DRB1*04:38, 04:56, 04:88, 04:100, 04:107 and 04:125 alleles.

Primer mix 42: Specific PCR fragment of 110 bp in the DRB1*04:41, 04:77, 04:102, 04:113, 04:152 and 04:164 and the DRB1*14:10, 14:57 and 14:143 alleles. Specific PCR fragment of 140 bp in the DRB1*04:79 allele.

‘w’, might be weakly amplified.

4The following DRB1*04 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| DRB1*04:16, 04:151 | 12 |
| DRB1*04:27, 04:110 | 20 |
| DRB1*04:39, 04:101 | 19 |

The DRB1*04 subtyping kit cannot distinguish the silent mutation in the DRB1*04:01:01-04:01:13 alleles, the DRB1*04:02:01-04:02:03 alleles, the DRB1*04:03:01-04:03:07 alleles, the DRB1* 04:04:01-04:04:03 and 04:04:06-04:04:08 alleles, the DRB1*04:05:01, 04:05:03-04:05:07 and 04:05:09-04:05:11, 04:05:13-04:05:14 and 04:05:16 alleles, the 04:05:02 and 04:05:15 alleles, the DRB1*04:06:01-04:06:05 alleles, the DRB1*04:07:01-04:07:04 alleles, the DRB1*04:08:01-04:08:03 alleles, the DRB1*04:11:01-04:11:02 alleles, the DRB1*04:17:01-04:17:02 alleles and the DRB1*04:105:01-04:105:02 alleles.

Change in revision R01 compared to R00:

- Primer mix 33 may have tendencies of unspecific amplifications. A foot note has been added in the Specificity Tables.