Gel Documentation Form and Worksheet

DRB1*04 (101.114-24/03, -24u/03u) 	Lot No: 95X	Expiry Date: 2017-August-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 21, 22 and 33 may have tendencies of unspecific amplifications.

Primer mixes 14, 31 and 38 have a tendency to giving rise to primer oligomer formation.

Primer mixes 24 and 32 may give rise to a lower yield of HLA-specific PCR product than the other DRB1*04 primer mixes.

Primer mix 48 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1DRB1*04 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.18.0, October 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 11: Specific PCR fragment of 175 bp in the DRB1*04:15, 04:61, 04:138, 04:145, 04:154 and 04:170 and the DRB1*11:22 alleles. Specific PCR fragment of 220 bp in the DRB1*04:22 and 04:98:02 alleles.

Primer mix 12: Specific PCR fragment of 180 bp in the DRB1*04:16 and 04:180 alleles. Specific PCR fragment of 265 bp in the DRB1*04:23, 04:52 and 04:151 alleles.

Primer mix 13: Specific PCR fragment of 170 bp in the DRB1*04:46 allele. Specific PCR fragments of 225 bp in the DRB1*04:12, 04:18, 04:25, 04:31, 04:54-04:55, 04:58, 04:86, 04:103, 04:126 and 04:185 alleles.

Primer mix 15: Specific PCR fragment of 175 bp in the DRB1*04:123 and the DRB1*01:10 allele. Specific PCR fragment of 215 bp in the DRB1*04:34 allele.

Primer mix 19: Specific PCR fragment of 140 bp in the DRB1*04:39 allele. Specific PCR fragment of 210 bp in the DRB1*04:24, 04:69, 04:78, 04:101 and 04:168 and the DRB1*14:10 and 14:143 alleles.

Primer mix 20: Specific PCR fragment of 100 bp in the DRB1*04:63, 04:110 and 04:161 alleles. Specific PCR fragment of 175 bp in the DRB1*04:40 allele. Specific PCR fragment of 255 bp in the DRB1*04:27 allele.

Primer mix 21: Specific PCR fragment of 155 bp in the DRB1*04:29 allele. Specific PCR fragment of 210 bp in the DRB1*04:37, 04:58, 04:73, 04:105:01-04:105:02, 04:122 and 04:132 alleles.

Primer mix 25: Specific PCR fragment of 125 bp in the DRB1*04:26 allele. Specific PCR fragment of 230 bp in the DRB1*04:45 allele. Specific PCR fragment of 260 bp in the DRB1*04:64 allele.

Primer mix 27: Specific PCR fragment of 115 bp in the DRB1*04:10:01-04:12, 04:60, 04:67, 04:91, 04:126, 04:136 and 04:147 and in the DRB1*03:12, 08:06, 08:10, 08:12, 08:22, 08:54, 12:46, 13:04, 13:32, 13:48, 13:58, 13:75, 13:81, 13:89, 13:93-13:94:02, 13:108, 13:164, 13:169, 14:65, 14:78 and 15:12 alleles. Specific PCR fragment of 265 bp in the DRB1*04:49 allele.

Primer mix 29: Specific PCR fragment of 75 bp in the DRB1*04:07:01-04:07:05, 04:17:01-04:17:02, 04:20, 04:69, 04:74, 04:92, 04:102, 04:129, 04:160, 04:170 and 04:172 and the DRB1*01:17, 09:01:02-09:01:05, 09:01:08-09:02:02, 09:04-09:05, 09:07-09:08, 09:10-09:24, 14:07:01-14:07:02, 14:14, 14:36, 14:42, 14:44:01-14:44:03, 14:51, 14:68, 14:93, 14:134 and 14:143 and in the DRB5*01:12 alleles. Specific PCR fragment of 260 bp in the DRB1*04:64 allele.

Primer mix 31: Specific PCR fragment of 85 bp in the DRB1*04:14, 04:106 and 04:132 and the DRB1*01:03, 01:39N, 08:40, 11:11:01-11:11:03, 11:14:01-11:14:02, 11:20, 11:68, 11:73, *11:93, 11:122, 11:168, 13:02:01-13:02:03, 13:02:05-13:02:09, 13:03:01-13:03:06, 13:23:01-13:23:02, 13:29, 13:31, 13:33:01-13:34, 13:36-13:39, 13:41, 13:45, 13:63, 13:65-13:67, 13:73-13:74, 13:85, 13:88, 13:90, 13:95-13:97:02, 13:99, 13:101, 13:103-13:104, 13:107, 13:115, 13:120, 13:122-13:124, 13:126, 13:128, 13:135, 13:139, 13:143, 13:145, 13:147, 13:151-13:152, 13:155, 13:159, 13:165, 13:167-13:168, 13:170-13:172, 13:174, 13:179-13:180, 13:182, 13:188 alleles. Specific PCR fragment of 260 bp in the DRB1*04:64 allele.

Primer mix 35: Specific PCR fragment of 210 bp in the DRB1*04:81N allele. Specific PCR fragment of 245 bp in the DRB1*04:51 and 04:70 alleles.

Primer mix 36: Specific PCR fragment of 175 bp in the DRB1*04:96 allele. Specific PCR fragment of 240 bp in the DRB1*04:71 allele.

Primer mix 38: Specific PCR fragment of 140 bp in the DRB1*04:89 allele. Specific PCR fragment of 225 bp in the DRB1*04:73, 04:75, 04:105:01-04:105:02, 04:122 and 04:178N alleles.

Primer mix 39: Specific PCR fragment of 100 bp in the DRB1*04:83 allele. Specific PCR fragment of 175 bp in the DRB1*04:15, 04:61-04:62, 04:69, 04:138, 04:145-04:146, 04:154 and 04:170 and the DRB1*11:22, 14:10, 14:57 and 14:143 alleles.

Primer mix 40: Specific PCR fragment of 145 bp in the DRB1*04:30 and 04:85 alleles. Specific PCR fragment of 200 bp in the DRB1*04:38, 04:56:01-04:56:02, 04:88, 04:100, 04:107 and 04:125 alleles.

Primer mix 42: Specific PCR fragment of 110 bp in the DRB1*04:41, 04:77, 04:102, 04:113, 04:152 and 04:164 and the DRB1*14:10, 14:57 and 14:143 alleles. Specific PCR fragment of 140 bp in the DRB1*04:79 allele.

Primer mix 44: Specific PCR fragment of 165 bp in the DRB1*04:127 allele. Specific PCR fragment of 240 bp in the DRB1*04:87, 04:94N, 04:120N, 04:129, 04:157N and 04:186N alleles.

Primer mix 47: Specific PCR fragment of 140 bp in the DRB1*04:117 allele. Specific PCR fragment of 250 bp in the DRB1*04:119N and 04:142N alleles.

4The DRB1*04:34 and the DRB1*01:10 alleles will give rise to identical amplification patterns with the DRB1*04 subtyping kit. These alleles can be distinguished by the DR low resolution kit and/or the DRB1*01 subtyping kit.

5The following DRB1*04 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| DRB1*04:16, 04:151 | 12 | DRB1*04:39, 04:101 | 19 |
| DRB1*04:23, 04:180 | 12 | DRB1*04:89, 04:178N | 38 |
| DRB1*04:27, 04:110 | 20 | DRB1*04:117, 04:119N | 47 |

The DRB1*04 subtyping kit cannot distinguish the silent mutation in the DRB1*04:01:01-04:01:16 alleles, the DRB1*04:02:01-04:02:03 alleles, the DRB1*04:03:01-04:03:07 and 04:03:09-04:03:10 alleles, the DRB1*04:04:01-04:04:03 and 04:04:06-04:04:09 alleles, the DRB1*04:05:01, 04:05:03-04:05:07 and 04:05:09-04:05:11, 04:05:13-04:05:14 and 04:05:16 alleles, the 05:05:02 and 04:05:15 alleles, the DRB1*04:06:01-04:06:05 alleles, the DRB1*04:07:01-04:07:05 alleles, the DRB1*04:08:01-04:08:03 alleles, the DRB1*04:10:01 and 04:10:03 alleles, the DRB1*04:11:01-04:11:02 alleles, the DRB1*04:17:01-04:17:02 alleles,  the DRB1*04:56:01-04:56:02 alleles and the DRB1*04:105:01-04:105:02 alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

- Primer mix 33 may have tendencies of unspecific amplifications. A foot note has been added in the Specificity Tables.