Gel Documentation Form and Worksheet

DRB1\*08 (101.127-12/04, -12u/04u) Lot No: 0L6 Expiry Date: 2024-04-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

*Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes: \_\_\_\_\_\_\_\_\_\_\_ Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |



Abbreviations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 2 and 5 may have tendencies of unspecific amplifications.

Primer mixes 3 and 24 might have a tendency to giving rise to primer oligomer formation.

Primer mix 24 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 200 base pairs.







**1**DRB1\*08 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.31.0, January 2018.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following DRB1\*08 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified DRB1\*08 alleles | Other amplified DRB1 alleles |
| **9** | 125 bp  175 bp | \*08:26, 08:35-08:36:02  \*08:14 | \*03:12, 03:158, 04:202, 13:32, 13:65, 13:93, 13:120, 13:139, 14:13, 14:63, 14:65, 14:78, 14:85, 14:169-14:170 |
| **10** | 205 bp  250 bp | \*08:40, 08:92  \*08:12, 08:22 | \*11:229, 12:73, 13:17, 13:116, 13:175, 13:214  \*12:01:01:01-12:02:10, 12:04-12:15, 12:17-12:18, 12:20-12:21, 12:23-12:38, 12:40-12:41, 12:43-12:56, 12:58-12:82, 14:28 |
| **12** | 95 bp  145 bp  195 bp  220 bp | \*08:27  \*08:44  \*08:13, 08:48  \*08:43 | \*12:12, 12:62 |
| **13** | 135 bp  165 bp  260 bp | \*08:33  \*08:11  \*08:39 |  |
| **14** | 135 bp  250 bp | \*08:17, 08:28, 08:37, 08:45:01-08:45:02, 08:63  \*08:60N | \*03:92, 11:67, 11:193:01-11:193:02, 11:209, 11:229, 12:01:01:01-12:20, 12:22-12:37, 12:39-12:41, 12:43-12:52, 12:54-12:56, 12:58-12:75, 12:77-12:82, 13:17, 13:175, 13:214, 14:138, 14:175 |
| **17** | 75 bp  115 bp  265 bp | \*08:04:01, 08:04:04-08:04:07, 08:06, 08:10, 08:12, 08:20, 08:22, 08:28, 08:54, 08:59, 08:67, 08:70, 08:75, 08:80, 08:83, 08:94, 08:98-08:99  \*08:14 | \*04:12, 04:18, 04:25, 04:58, 11:25, 11:67, 11:119, 13:18, 13:144, 13:156, 13:164, 14:12:01-14:12:02, 14:15, 14:78, 14:84, 14:156, 14:173, 15:21  \*11:261  \*11:261 |
| **18** | 150 bp  225 bp | \*08:09:01-08:09:02, 08:21, \*08:32, 08:35, 08:87  \*08:20 | \*13:232, 14:15, 14:40, 14:55, 14:77, 14:84, 14:136, 14:164, 14:173  \*03:157, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:119, 13:13, 13:18, 13:47, 13:55, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158, 13:164, 13:232, 13:286, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:27:01-14:27:02, 14:40, 14:55, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89, 14:102, 14:115-14:116, 14:135-14:136, 14:144, 14:156, 14:164, 14:173-14:174, 14:181, 14:218 |
| **19** | 100 bp  170 bp | \*08:16, 08:38  \*08:07, 08:49 |  |
| **21** | 70 bp  120 bp | \*08:78N  \*08:17, 08:28, 08:37, 08:45:01-08:45:02, 08:63 | \*03:157, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:67, 11:119, 13:18, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158, 13:286, 15:21 |
| **22** | 130 bp  165 bp  215 bp | \*08:23  \*08:19  \*08:29 | \*12:66, 14:04:01:01-14:04:07, 14:11, 14:28, 14:68:01-14:68:02, 14:71, 14:73, 14:93, 14:120, 14:126:01-14:126:02, 14:138, 14:145, 14:148, 14:152N, 14:175, 14:180, 14:196, 14:201, 14:205-14:206, 14:210Q |
| **23** | 220 bp  250 bp | \*08:43  \*08:04:01, 08:04:04-08:04:07, 08:06, 08:10, 08:28, 08:31, 08:54, 08:59, 08:67, 08:70, 08:75, 08:83, 08:94, 08:98-08:99 | \*03:92, 11:67, 11:209, 11:229, 12:03:02-12:03:03, 12:19, 13:17, 13:116, 13:175, 13:214, 14:04:01:01-14:04:07, 14:11, 14:15, 14:31, 14:50, 14:52, 14:73, 14:76, 14:79, 14:107, 14:120, 14:126:01-14:126:02, 14:138, 14:145, 14:148, 14:152N, 14:175, 14:180, 14:196, 14:201, 14:205-14:206, 14:210Q |

**4**The following alleles cannot be separated by the HLA-DRB1\*08 primer set. These alleles can be distinguished by the DR low resolution kit and/or the DRB1\*14 high resolution kit.

|  |
| --- |
| Alleles |
| DRB1\*08:20, DRB1\*14:12:01-14:12:02, 14:84, 14:156, 14:173 |

Abbreviations

‘w’, may be weakly amplified.