PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM DRB1*09

Prod. No:	101.120-06 – including Taq polymerase

101.120-06u – without Taq polymerase

Lot No:    	Y45

Expiry Date:    	2009-September-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 7 primer mixes used for DRB1*09 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB1*09 alleles |
| --- | --- | --- | --- |
| 1 | 195 bp | 515 bp | 090102-090103, 0902-0904, 0906 |
| 2 | 130 bp | 430 bp | 090102-090103, 0903-0906 |
| 3 | 135 bp | 430 bp | 0902 |
| 4 | 215 bp | 430 bp | 0903 |
| 5 | 215 bp | 430 bp | 0904 |
| 6 | 250 bp | 430 bp | 0905 |
| 7 | 220 bp | 430 bp | 0906 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*09 SSP subtypings.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most tubes, or a band of 515 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*09 subtyping and in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

1The nucleotide sequence of the DRB1*090101 allele has been shown to contain errors and be identical to DRB1*090102.