PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*11

Prod. No:	101.115-24/24u

101.115-03/03u

Lot No:    	43F

Expiry Date:    	2010-December-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for DRB1*11 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB*11 alleles3 | Other amplified DRB alleles3,4 |
| --- | --- | --- | --- | --- |
| 1 | 175 bp | 515 bp | 110101-110404, 110601-1121, 1123-1129, 1131-1166, 1168-1170 | 0308 |
| 2 | 175 bp | 430 bp | 1105, 1167 | 0831, 1204, 1411 |
| 35 | 65 bp | 430 bp | 110101-110107, 1103-110602, 1109, 111001-111101-111202, 1115, 1122-1125, 112701-1130, 1132, 1133, 1135, 1137-1141, 1143, 1144, 1146, 1147, 1149-1151, 115401-1156, 1158-1163, 1166, 1167 | 0415, 0831, DRB5*0113 |
| 4 | 200 bp | 430 bp | 110201-110202, 111401-111402, 1116, 1118-1121, 1131, 1136 weakly, 1145, 1148, 116501-116502, 1168 weakly, 1170 | 130101-1304, 1306, 1308, 1310, 1312, 1313, 1315, 1316, 1319, 1320 weakly, 1322, 1323, 1327, 1328, 1329 weakly, 1330-1341, 1343 weakly, 1345, 1348, 1351-1353, 1357-1359, 1361, 1364-1366, 1368-1370, 1371 weakly, 1372, 1373, 1374 weakly, 1376, 1377, 1378 weakly, 1379-1385, 1387, 1416, 1463, 1467, 1478 |
| 55,6 | 85 bp | 515 bp | 110101-110107, 110401-110602, 110801-111001, 111201-1113, 1115, 1117-111902, 1122-1135, 1137-1139, 1142-1147, 1149, 1150, 1152, 115401-1158, 1160-1162, 1164, 1166, 1167 | 0415, 0461, 0831, 1204, 1411, DRB5*0113 |
| 65,6 | 80 bp | 515 bp | 110201-1103, 111101-111102, 111401-111402, 1116, 1120, 1121, 1136, 1140, 1141, 1148, 1155, 1159, 1163, 116501-116502, 1168, 1170 |  |
| 75 | 125 bp | 430 bp | 110101-110107, 1105, 110801-111001-111202, 111401- 1115, 111901-1120, 1122-1124, 1126-1133, 1137, 1139, 1145, 1149, 1151, 1153, 1161, 1162, 1164, 1166, 1168, 1169 | 0461, DRB3*0218, DRB5*0113 |
| 85 | 125 bp | 430 bp | 110201-110404, 110601-1107, 1113, 1116-1118, 1121, 1125, 1134-1136, 1138, 1140-1144, 1146-1148, 1150, 1152, 115401-1160, 1163, 116501-116502, 1167, 1170 | 0308, 0415, 0831, 1204, 1411 |
| 95 | 85 bp | 430 bp | 1107, 1123, 1125, 1132, 1145, 1153, 1155, 1164, 1167, 1170 | 0308 |
| 105,7 | 65, 85 bp | 430 bp | 1107, 1113, 1117, 1123, 1125, 1126, 1134, 1145, 1152, 1153weakly, 1155, 1164, 1167 | 0308 , 0461, 1411, DRB3*0218 |
| 118 | 200, 255 bp | 515 bp | 110601-110602, 110801-110802, 1121, 1136, 1142, 1147, 1157, 1164, 1168 | 0320, 1320, 1325, 1329, 1343, 1356, 1358, 1360, 1371, 1378, 1381, 140301-140302, 1412, 1429, 1440, 1469, 1474, 1477 |
| 125,9 | 95, 185 bp | 430 bp | 1130, 1161 | DRB3*0218 |
| 135,10 | 70, 105 bp | 430 bp | 1109, 1116, 1120, 1128, 1140, 1158-1160 | 0308 |
| 145,6 | 145 bp | 430 bp | 1117, 1137, 1141, 1148, 1150, 1152, 1156 | 030201-0303, 0306, 0317, 0326, 0327, 0331, 0335, 0338, 0820, 130301-130302, 130701-1308, 1312, 1313, 1319, 1326, 1332, 1333, 1336-1338, 1340, 1347-1349, 1353, 1355, 1358, 1360, 1365, 1370, 1372, 1376, 1381, 1384, 1385, 140101-140301, 140501-1409, 1412-1414, 1416, 1418-1420, 1422-1427, 1429, 143201-143202, 1434, 1436-1441, 1443-1445, 1447-1449, 1451, 1454-1456, 1458-1460, 1462, 1463, 1467, 1469, 1470, 1474, 1475, 1477, 1478, 1480-1483 |
| 155,11 | 60, 105 bp | 430 bp | 111001, 111201-1113, 1117, 1152, 1166 | 1411 |
| 165 | 110 bp | 430 bp | 1115, 1124, 1144, 1149, 1150 | 0326, 0333, 1362, 1368 |
| 175 | 90 bp | 430 bp | 1117, 1123, 1125, 1145, 1152-1155, 1164, 1167 | 1411 |
| 185,6 | 95 bp | 515 bp | 1107, 112701-112702 | 0308, DRB3*0218 |
| 1912 | 180, 225 bp | 430 bp | 1123, 1125, 1131, 1132, 1145, 1152, 1155, 1164, 1170 | 0310, 0820, 1313, 1318, 1343, 1345, 1347, 1355, 140101-140103, 140301-140302, 140701-1408, 1412, 1416, 1422, 1425-1427, 143201-143202, 1434, 1435, 1438-1440, 1449, 1453-1455, 1458, 1460, 1462, 1463, 1467, 1469, 1470, 1472, 1474, 1475, 1477, 1478, 1482 |
| 205,6 | 115 bp | 515 bp | 111001, 111201-1113, 1117, 1129, 1144, 1147, 1149, 1150, 1152, 1163 | 0304, 0321, 0326, 0330, 0333, 0335, 1308, 1319, 1357, 1360, 1372, 1376, 1383, 1384, 1387, 140101-140103, 140501-140503, 140701-1408, 1414, 1416, 1420, 1422-142302, 1426, 143201-143202, 1434-1440, 1443-1445, 1449, 1454-1456, 1460, 1462, 1465, 1470, 1472, 1474, 1475, 1482 |
| 2113 | 135, 175 bp | 430 bp | 1122, 1159 | 0308, 0415, 0461 |
| 225,14 | 115, 180, 215 bp | 430 bp | 1115, 1133, 1135, 116501-116502, 1169 | 1309, 1362, 1424, 1437 |
| 235 | 120 bp | 430 bp | 1138 |  |
| 2415 | 155, 255 bp | 430 bp | 110601-110602, 1121, 1139, 1143, 1147, 1150, 1157 | 0317, 0320, 0337, 1358, 1381, 1429, 1459 |
| 255 | 110 bp | 515 bp | 110404, 1109, 111001, 111202, 1113, 1116, 1117, 1120, 1122, 1140, 1146, 1152, 1158, 1159 | 0308, 1411 |
| 265 | 120 bp | 430 bp | 1109, 111001, 1113, 1116, 1117, 1120, 1140, 1146, 1152, 1159, 1162 | 0308, 1411 |
| 275 | 85 bp | 430 bp | 111101-111102, 111401-111402, 1120, 1168 | 0103, 0414, 130201-130302, 1323, 1329, 1331, 1333, 1334, 1336-1339, 1341, 1345, 1363, 1365, 1367, 1373, 1374 |
| 285 | 85 bp | 430 bp | 110201-1103, 1116, 1121, 1136, 1140, 1141, 1148, 1155, 1159, 1163, 116501-116502, 1170 | 0402, 0437, 130101-130103, 1304, 1308, 1310, 1315, 1317, 1319, 1320, 1322, 1324, 1327, 1328, 1332, 1335, 1340, 1343, 1348, 1351-1354, 1357, 1359, 1361, 1364, 1368-1372, 1375, 1376, 1378-1380, 1383, 1384, 1387,1416, 1457, 1510 |
| 295 | 100 bp | 515 bp | 111401-111402, 111901-1120, 1131, 1145 | 0103, 0115, 0414, 0438, 07010101-070103, 0703-0715, 080302, 0814, 0815, 0818, 0819, 0823, 0825, 0827, 0829, 0830, 0832-0836, 130201-130302, 1312, 1313, 1323, 1330, 1331, 1333, 1334, 1336-1339, 1341, 1345, 1365-1367, 1373, 1374, 1382, 1385, 1424, 1463, 1467, 150201-150206, 1511, 1514, 1519, 1526, 1529, 1530, 160501-160502, 1607, DRB3*0211, DRB5*0107, DRB5*0111, DRB5*0203 |
| 305 | 100 bp | 515 bp | 110101-110107, 1103-110602, 1109, 111001, 111101-111202, 1115, 1122-1125, 112701-1130, 1132, 1133, 1135, 1138-1140, 1143, 1144, 1146, 1147, 1149, 1151, 115401-1155, 1158-1163, 1166, 1167 | 0817, 0828, 120201-120203, 1213, 1215, 1216, 130501-130502, 131101-131102, 131401-131403, 1318, 1321, 1324, 1342, 1346, 135001-135002, 1354, 1362, 1363, 1375, 1453, 1609 |
| 315 | 85 bp | 515 bp | 110401-110404, 110601-110602, 1118, 1125, 1135, 1138, 1142-1144, 1146, 1147, 1150, 115401-115402, 1156-1158, 1160, 1167 | 0412, 0415, 0418, 0425, 0436, 0453, 0458, 080401, 080404, 0806, 0810, 0812, 0820, 0822, 0828, 0831, 120101-120203, 120302-1215, 1217, 1306, 131101-131102, 1318, 1342, 1358, 1377, 1412, 1415, 1473, 1474, 1478, 1521 |
| 325 | 75 bp | 430 bp | 1113, 1117, 1151, 1152 | 1411 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*11 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*11 subtyping.

In addition, wells number 5, 6, 11, 18, 20, 25, 29, 30 and 31 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

The nucleotide sequence of the DRB1*111002 allele is not yet retrievable. The DRB1*1171 allele is not included in the kit.

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, two DRB1*01 alleles, seventeen DRB1*03 alleles, twelve DRB1*04 alleles, most DRB1*07 alleles, most DRB1*08 alleles, all DRB1*12 alleles, many DRB1*13 and DRB1*14 alleles, nine DRB1*15 and four DRB1*16 alleles when present on the other haplotype. In addition, the DRB3*0211, DRB5*0107, DRB5*0111 and DRB5*0203 alleles are amplified by primer mix 29 and the DRB3*0218 allele is amplified by primer mixes 7, 10, 12 and 18 and the DRB5*0113 allele is amplified by primer mixes 3, 5, and 7.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity than longer PCR bands. In order to have as little co-amplification as possible of non-DRB1*11 alleles by the primer pairs of the DRB1*11 subtyping kit, the DRB1*11 characteristic sequence motif in codon 58, ‘Ag’, has been used in either the 5’-primer or the 3’-primer in many of the DRB1*11 primer mixes and consequently many of the specific PCR products are shorter than 100 base pairs.

6Primer mixes 5, 6, 14, 18 and 20 frequently gives rise to nonspecific amplification and/or a primer oligomer artifact.

7Primer mix 10: Specific PCR fragment of 65 bp in the DRB1*1113, 1117, 1126, 1134, 1152, and the 0461, 1411 and DRB3*0218 alleles. Specific PCR fragment of 85 bp in the DRB1*1107, 1123, 1125, 1145, 1153weakly, 1155, 1164, 1167 and the 0308 alleles.

8Primer mix 11: Specific PCR fragment of 200 bp in the DRB1*110801-110802, 1136, 1142, 1164, 1168 and the 1320, 1325, 1329, 1343, 1356, 1360, 1371, 1378, 140301-140302, 1412, 1440, 1469, 1474 and 1477 alleles. Specific PCR fragment of 255 bp in the DRB1*110601-110602, 1121, 1147, 1157 and the 0320, 1358, 1381 and 1429 alleles.

9Primer mix 12: Specific PCR fragment of 95 bp in the DRB1*1161 allele. Specific PCR fragment of 185 bp in the DRB1*1130 and the DRB3*0218 alleles.

10Primer mix 13: Specific PCR fragment of 70 bp in the DRB1*1160 allele. Specific PCR fragment of 105 bp in the DRB1*1109, 1116, 1120, 1128, 1140, 1158, 1159 and the 0308 alleles.

11Primer mix 15: Specific PCR fragment of 60 bp in the DRB1*1166 allele. Specific PCR fragment of 105 bp in the DRB1*1110, 111201-111202, 1113, 1117, 1152, 1166 and the 1411 alleles.

12Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*1131, 1152 and the 0310, 1343, 1345, 140101-140103, 140701-1408, 1416, 1422, 1425, 1426, 143201-143202, 1434, 1435, 1438, 1439, 1449, 1453-1455, 1463, 1469, 1470, 1472, 1474, 1475 and 1482 alleles. Specific PCR fragment of 225 bp in the DRB1*1123, 1125, 1132, 1145, 1155, 1164, 1170 and the 0820, 1313, 1318, 1347, 1355, 140301-140302, 1412, 1427, 1440, 1458, 1460, 1462, 1467, 1477 and 1478 alleles.

13Primer mix 21: Specific PCR fragment  of 135 bp in the DRB1*1159 and the 0308 alleles. Specific PCR fragment of 175 bp in the DRB1*1122 and the 0415 and 0461 alleles.

14Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*1115 and the 1362 alleles. Specific PCR fragment of 180 bp in the DRB1*1133, 1135 and 116501-116502 alleles. Specific PCR fragment of 215 bp in DRB1*1169 and 1309, 1424 and 1437 alleles.

15Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*1139, 1143, 1150 and the 0317, 0337 and 1459 alleles. Specific PCR fragment of 255 bp in the DRB1*110601-110602, 1121, 1147, 1157 and the 0320, 1358, 1381 and 1429 alleles.

1The nucleotide sequence of the DRB1*111002 allele is not yet retrievable. The DRB1*1171 allele is not included in the kit.

1Primer mix 10: Specific PCR fragment of 65 bp in the DRB1*1113, 1117, 1126, 1134, 1152, and the 0461, 1411 and DRB3*0218 alleles. Specific PCR fragment of 85 bp in the DRB1*1107, 1123, 1125, 1145, 1153weakly, 1155, 1164, 1167 and the 0308 alleles.

Primer mix 11: Specific PCR fragment of 200 bp in the DRB1*110801-110802, 1136, 1142, 1164, 1168 and the 1320, 1325, 1329, 1343, 1356, 1360, 1371, 1378, 140301-140302, 1412, 1440, 1469, 1474 and 1477 alleles. Specific PCR fragment of 255 bp in the DRB1*110601-110602, 1121, 1147, 1157 and the 0320, 1358, 1381 and 1429 alleles.

Primer mix 12: Specific PCR fragment of 95 bp in the DRB1*1161 allele. Specific PCR fragment of 185 bp in the DRB1*1130 and the DRB3*0218 alleles.

Primer mix 13: Specific PCR fragment of 70 bp in the DRB1*1160 allele. Specific PCR fragment of 105 bp in the DRB1*1109, 1116, 1120, 1128, 1140, 1158, 1159 and the 0308 alleles.

Primer mix 15: Specific PCR fragment of 60 bp in the DRB1*1166 allele. Specific PCR fragment of 105 bp in the DRB1*1110, 111201-111202, 1113, 1117, 1152, 1166 and the 1411 alleles.

Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*1131, 1152 and the 0310, 1343, 1345, 140101-140103, 140701-1408, 1416, 1422, 1425, 1426, 143201-143202, 1434, 1435, 1438, 1439, 1449, 1453-1455, 1463, 1469, 1470, 1472, 1474, 1475 and 1482 alleles. Specific PCR fragment of 225 bp in the DRB1*1123, 1125, 1132, 1145, 1155, 1164, 1170 and the 0820, 1313, 1318, 1347, 1355, 140301-140302, 1412, 1427, 1440, 1458, 1460, 1462, 1467, 1477 and 1478 alleles.

Primer mix 21: Specific PCR fragment  of 135 bp in the DRB1*1159 and the 0308 alleles. Specific PCR fragment of 175 bp in the DRB1*1122 and the 0415 and 0461 alleles.

Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*1115 and the 1362 alleles. Specific PCR fragment of 180 bp in the DRB1*1133, 1135 and 116501-116502 alleles. Specific PCR fragment of 215 bp in DRB1*1169 and 1309, 1424 and 1437 alleles.

Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*1139, 1143, 1150 and the 0317, 0337 and 1459 alleles. Specific PCR fragment of 255 bp in the DRB1*110601-110602, 1121, 1147, 1157 and the 0320, 1358, 1381 and 1429 alleles.

‘w’, may be weakly amplified.

Changes in revision R01 compared to R00:

- Length of the specific PCR products generated by primer mixes 12 and 13 have been corrected in the Interpretation Table.

Changes in revision R03 compared to R02:

- The DRB1*1136, 1168, 1320, 1329, 1343, 1371 and 1378 are weakly amplified by primer mix 4.