PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*11

Prod. No:	101.115-24/24u

101.115-03/03u

Lot No:    	55G

Expiry Date:    	2011-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for DRB1*11 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB*11 alleles3 | Other amplified DRB alleles3,4 |
| --- | --- | --- | --- | --- |
| 1 | 175 bp | 515 bp | *110101-110405, 110601-1121, 1123-1129, 1131-1166, 1168-1170, 1172-1184 | *0308 |
| 2 | 175 bp | 430 bp | *1105, 1167 | *0831, 1204, 1411 |
| 35 | 65 bp | 430 bp | *110101-110111, 1103-110602, 1109-111202, 1115, 1122-1125, 112701-1130, 1132, 1133, 1135, 1137-1141, 1143, 1144, 1146, 1147, 1149-1151, 115401-1156, 1158-1163, 1166, 1167, 1172, 1174-1178, 1180, 1181, 1183, 1184 | *0415, 0831, DRB5*0113 |
| 46 | 200 bp | 430 bp | *110201, 110202, 111401, 111402, 1116, 1118-1121, 1131, 1145, 1148, 116501, 116502, 1170, 1173 | *130101-1304, 1306, 1308, 1310, 1312, 1313, 1315, 1316, 1319, 1322, 1323, 1327, 1328, 1330-1341, 1345, 1348, 1351-1353, 1357-1359, 1361, 1364-1366, 1368-1370, 1372, 1373, 1376, 1377, 1379-1385, 1387, 1389-1394, 1416, 1463, 1467, 1478 |
| 55,7 | 85 bp | 515 bp | *110101-110111, 110401-110602, 110801-111002, 111201-111302, 1115, 1117-111902, 1122-1135, 1137-1139, 1142-1147, 1149, 1150, 1152, 115401-1158, 1160-1162, 1164, 1166, 1167, 1172, 1174, 1175, 1177, 1178, 1180, 1181, 1183, 1184 | *0415, 0461, 0831, 1204, 1411, DRB5*0113 |
| 65,7 | 80 bp | 515 bp | *110201-1103, 111101, 111102, 111401, 111402, 1116, 1120, 1121, 1136, 1140, 1141, 1148, 1155, 1159, 1163, 116501, 116502, 1168, 1170, 1176 |  |
| 75 | 125 bp | 430 bp | *110101-110111, 1105, 110801-111202, 111401-1115, 111901-1120, 1122-1124, 1126-1133, 1137, 1139, 1145, 1149, 1151, 1153, 1161, 1162, 1164, 1166, 1168, 1169, 1172-1175, 1179-1182 | *0461, DRB3*0218, DRB5*0113 |
| 85 | 125 bp | 430 bp | *110201-110405, 110601-1107, 111301, 111302, 1116-1118, 1121, 1125, 1134-1136, 1138, 1140-1144, 1146-1148, 1150, 1152, 115401-1160, 1163, 116501, 116502, 1167, 1170, 1176-1178, 1183, 1184 | *0308, 0415, 0831, 1204, 1411 |
| 95,7,10 | 85 bp, 115 bp | 430 bp | *1107, 1123, 1125, 1132, 1145, 1155, 1164, 1167, 1170, 1177 | *0308 |
| 105,11 | 65 bp, 85 bp, 120 bp | 430 bp | *1107, 111301, 111302, 1117, 1123, 1125, 1126, 1134, 1145, 1152, 1155, 1164, 1167, 1174, 1179 | *0308, 0461, 1411, DRB3*0218 |
| 117,12 | 155 bp, 200 bp, 255 bp | 515 bp | *110601, 110602, 110801, 110802, 1121, 1136, 1142, 1147, 1157, 1164, 1168, 1175 | *0320, 1320, 1325, 1329, 1343, 1356, 1358, 1360, 1371, 1378, 1381, 1388, 140301, 140302, 1412, 1429, 1440, 1469, 1474, 1477, 1484, 1485 |
| 125,13 | 95 bp, 185 bp | 430 bp | *1130, 1161 | DRB3*0218 |
| 135,7,14 | 70 bp, 105 bp | 430 bp | *1109, 1116, 1120, 112801, 112802, 1140, 1158-1160, 1183 | *0308 |
| 145,7 | 145 bp | 430 bp | *1117, 1137, 1141, 1148, 1150, 1152, 1156, 1183 | *030201-0303, 0306, 0317, 0326, 0327, 0331, 0335, 0338, 0341, 0820, 130301-130303, 130701-1308, 1312, 1313, 1319, 1326, 1332-133303, 1336-1338, 1340, 1347-1349, 1353, 1355, 1358, 1360, 1365, 1370, 1372, 1376, 1381, 1384, 1385, 1388-1390, 1394, 140101-140301, 140501-1409, 1412-1414, 1416, 1418-1420, 1422-1427, 1429, 143201, 143202, 1434, 1436-1441, 1443-1445, 1447-1449, 1451, 1454-1456, 1458-1460, 1462, 1463, 1467, 1469, 1470, 1474, 1475, 1477, 1478, 1480-1492N, DRB3*0114 |
| 155,15 | 60 bp, 105 bp | 430 bp | *111001, 111002, 111201-111302, 1117, 1152, 1166 | *1411 |
| 165 | 110 bp | 430 bp | *1115, 1124, 1144, 1149, 1150, 1176 | *0326, 0333, 0343, 0344, 1334, 1362, 1364, 1368, 1441, 1477, DRB3*0114 |
| 175,8 | 90 bp | 430 bp | *1117, 1123, 1125, 1145, 1152-1155, 1164, 1167 | *1411 |
| 185,7 | 95 bp | 515 bp | *1107, 112701, 112702, 1184 | *0308, DRB3*0218 |
| 197,16 | 180 bp, 225 bp | 430 bp | *1123, 1125, 1131, 1132, 1145, 1152, 1155, 1164, 1170 | *0310, 0820, 1313, 1318, 1343, 1345, 1347, 1355, 140101-140103, 140301, 140302, 140701-1408, 1412, 1416, 1422, 1425-1427, 143201, 143202, 1434, 1435, 1438-1440, 1449, 1453-1455, 1458, 1460, 1462, 1463, 1467, 1469, 1470, 1472, 1474, 1475, 1477, 1478, 1482, 1484-1490, 1492N |
| 205,7 | 115 bp | 515 bp | *111001, 111002, 111201-111302, 1117, 1129, 1144, 1147, 1149, 1150, 1152, 1163 | *0304, 0321, 0326, 0330, 0333, 0335, 0344, 1308, 1319, 1357, 1360, 1372, 1376, 1383, 1384, 1387, 140101-140103, 140501-140503, 140701-1408, 1414, 1416, 1420, 1422-142303, 1426, 143201, 143202, 1434-1440, 1443-1445, 1449, 1454-1456, 1460, 1462, 1465, 1470, 1472, 1474, 1475, 1482, 1484, 1486-1488, 1490-1492N |
| 2117 | 135 bp, 175 bp | 430 bp | *1122, 1159, 1172, 1180, 1183 | *0308, 0415, 0461 |
| 225,17 | 115 bp, 180 bp, 215 bp, 245 bp | 430 bp | *1115, 1133, 1135, 116501, 116502, 1169, 1181, 1182 | *1309, 1362, 1424, 1437 |
| 235 | 120 bp | 430 bp | *1138 |  |
| 2419 | 155 bp, 250 bp | 430 bp | *110601, 110602, 1121, 1139, 1143, 1147, 1150, 1157, 1181 | *0317, 0320, 0337, 1358, 1381, 1429, 1459, 1486 |
| 255 | 110 bp | 515 bp | *110404, 1109-111002, 111202-111302, 1116, 1117, 1120, 1122, 112802, 1140, 1146, 1152, 1158, 1159, 1183 | *0308, 1204, 1411, DRB3*0218 |
| 265,7,20 | 120 bp, 160 bp | 430 bp | *1109-111002, 111301, 111302, 1116, 1117, 1120, 1140, 1146, 1152, 1159, 1162, 1178, 1180, 1183 | *0308, 1411 |
| 275 | 85 bp | 430 bp | *111101, 111102, 111401, 111402, 1120, 1168, 1173 | *0103, 0414, 130201-130303, 1323, 1329, 1331, 1333, 1334, 1336-1339, 1341, 1345, 1363, 1365-1367, 1373, 1374, 1385 |
| 285 | 85 bp | 430 bp | *110201-1103, 1116, 1121, 1136, 1140, 1141, 1148, 1155, 1159, 1163, 116501, 116502, 1170, 1176 | *0402, 0437, 130101-130105, 1304, 1308, 1310, 1315, 1317, 1319, 1320, 1322, 1324, 1327, 1328, 1332, 1335, 1340, 1343, 1348, 1351-1354, 1357, 1359, 1361, 1364, 1368-1372, 1375, 1376, 1378-1381, 1383, 1384, 1387, 1391-1393, 1416, 1457, 1510 |
| 295 | 100 bp | 515 bp | *111401, 111402, 111901-1120, 1131, 1145, 1173, 1182 | *0103, 0115, 0124, 0414, 0438, 0482, 0486, 07010101-070103, 0703, 0704, 0706-0717, 080302, 0814, 0815, 0818, 0819, 0823, 0825, 0827, 0829-083002, 0832-0838, 130201-130303, 1312, 1313, 1323, 1330, 1331, 133301-1334, 1336-1339, 1341, 1345, 1365-1367, 1373, 1374, 1382, 1385, 1390, 1424, 1463, 1467, 150201-150207, 1511, 1514, 1519, 1526, 1529, 1530, 1538, 1539, 160501, 160502, 1607, DRB3*0211, DRB5*0107, 0111, DRB5*0203 |
| 305 | 100 bp | 515 bp | *110101-110111, 1103-110602, 1109-111202, 1115, 1122-1125, 112701-1130, 1132, 1133, 1135, 1138-1140, 1143, 1144, 1146, 1147, 1149, 1151, 115401-1155, 1158-1163, 1166, 1167, 1172, 1174-1178, 1180, 1181, 1184 | *0817, 0828, 120201-120204, 1213, 1215, 1216, 1218-1220, 130501, 130502, 131101, 131102, 131401-131403, 1318, 1321, 1324, 1342, 1346, 135001, 135002, 1354, 1362, 1363, 1375, 1453, 1609 |
| 315 | 85 bp | 515 bp | *110401-110405, 110601, 110602, 1118, 1125, 1135, 1138, 1142-1144, 1146, 1147, 1150, 115401, 115402, 1156-1158, 1160, 1167, 1177, 1178, 1183, 1184 | *0123, 0412, 0415, 0418, 0425, 0436, 0453, 0458, 080401, 080404, 0806, 0810, 0812, 0820, 0822, 0828, 0831, 120101-1215, 1217-1220, 1306, 131101, 131102, 1318, 1342, 1358, 1377, 1412, 1415, 1473, 1474, 1478, 1484, 1521, 1615 |
| 325,9 | 75 bp | 430 bp | *111301, 111302, 1117, 1151, 1152 | *1411 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*11 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*11 subtyping.

In addition, wells number 5, 6, 11, 18, 20, 25 and 29 to 31 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, other DRB1 alleles are amplified by primer mixes 1 to 5, 7 to 11, 13 to 22 and 24 to 32. The DRB3*0114 allele is amplified by primer mixes 14 and 16, the DRB3*0211, DRB5*0107, DRB5*0111 and DRB5*0203 alleles are amplified by primer mix 29, the DRB3*0218 allele is amplified by primer mixes 7, 10, 12 and 18 and 25 and the DRB5*0113 allele is amplified by primer mixes 3, 5 and 7.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity than longer PCR bands. In order to have as little co-amplification as possible of non-DRB1*11 alleles by the primer pairs of the DRB1*11 subtyping kit, the DRB1*11 characteristic sequence motif in codon 58, ‘Ag’, has been used in either the 5’-primer or the 3’-primer in many of the DRB1*11 primer mixes and consequently many of the specific PCR products are shorter than 100 base pairs.

6Primer mix 4 may faintly amplify the DRB1*1329 allele.

7Primer mixes 5, 6, 9, 11, 13, 14, 18 to 20 and 26 may give rise to nonspecific amplification and/or a primer oligomer artifact.

8In primer mix 17, the yield of specific PCR product may be lower than for the other DRB1*11 primer mixes.

9Primer mix 32 might generate a long unspecific PCR product. This band should be disregarded when interpreting the DRB1*11 typings.

10Primer mix 9: Specific PCR fragment of 85 bp in the DRB1*1107, 1123, 1125, 1132, 1145, 1155, 1164, 1167 and 1170 and the DRB1*0308 alleles. Specific PCR fragment of 115 bp in the DRB1*1177 allele.

11Primer mix 10: Specific PCR fragment of 65 bp in the DRB1*111301, 111302, 1117, 1126, 1134, 1152 and 1179, the DRB1*0308, 0461, 1411 and the DRB3*0218 alleles. Specific PCR fragment of 85 bp in the DRB1*1107, 1123, 1125, 1145, 1155, 1164 and 1167 alleles. Specific PCR fragment of 120 bp in the DRB1*1174 allele.

12Primer mix 11: Specific PCR fragment of 155 bp in the DRB1*1175 allele. Specific PCR fragment of 200 bp in the DRB1*110801, 110802, 1136, 1142, 1157, 1164 and 1168 and the DRB1*1320, 1325, 1329, 1343, 1356, 1360, 1371, 1378, 1388, 140301, 140302, 1412, 1440, 1469, 1474, 1477, 1484 and1485 alleles. Specific PCR fragment of 255 bp in the DRB1*110601, 110602, 1121, 1147 and 1157 and the DRB1*0320, 1358, 1381 and 1429 alleles.

13Primer mix 12: Specific PCR fragment of 95 bp in the DRB1*1161 allele. Specific PCR fragment of 185 bp in the DRB1*1130 and the DRB3*0218 alleles.

14Primer mix 13: Specific PCR fragment of 70 bp in the DRB1*1160 allele. Specific PCR fragment of 105 bp in the DRB1*1109, 1116, 1120, 112801, 112802, 1140, 1158, 1159, 1183 and the DRB1*0308 alleles.

15Primer mix 15: Specific PCR fragment of 60 bp in the DRB1*1166 allele. Specific PCR fragment of 105 bp in the DRB1*111001, 111002, 111201-111302, 1117 and 1152 and the DRB1*1411 alleles.

16Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*1131, 1152 and the 0310, 1343, 1345, 140101-140103, 140701-1408, 1416, 1422, 1425, 1426, 143201-143202, 1434, 1435, 1438, 1439, 1449, 1453-1455, 1458, 1460, 1462, 1469, 1470, 1472, 1474, 1475, 1482, 1486-1488, 1490 and 1492N alleles.

Specific PCR fragment of 225 bp in the DRB1*1123, 1125, 1132, 1145, 1155, 1164, 1170 and the 0820, 1313, 1318, 1347, 1355, 140301, 140302, 1412, 1427, 1440, 1455, 1463, 1467, 1477, 1478, 1484, 1485, 1489 alleles.

17Primer mix 21: Specific PCR fragment  of 135 bp in the DRB1*1159, 1172, 1180, 1183 and the DRB1*0308 alleles. Specific PCR fragment of 175 bp in the DRB1*1122 and the 0415 and 0461 alleles.

18Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*1115 and the 1362 alleles. Specific PCR fragment of 180 bp in the DRB1*1133, 1135 and 116501-116502 alleles. Specific PCR fragment of 215 bp in DRB1*1169, 1182 and DRB1*1309, 1424 and 1437 alleles. Specific PCR fragment of 245 bp in DRB1*1181 allele.

19Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*1139, 1143, 1150 and the DRB1*0317, 0337, 1459 and 1486 alleles. Specific PCR fragment of 250 bp in the DRB1*110601-110602, 1121, 1147, 1157 and 1181 and the DRB1*0320, 1358, 1381 and 1429 alleles.

20Primer mix 26: Specific PCR fragment of 120 bp in the DRB1*1109-111002, 111301, 111302, 1116, 1117, 1120, 1140, 1146, 1152, 1159, 1162, 1180 and 1183 and the DRB1*0308 and 1411 alleles. Specific PCR fragment of 160 bp in the DRB1*1178 allele.

1 The DRB1*1171 has been shown to be in error.

2Primer mix 9: Specific PCR fragment of 85 bp in the DRB1*1107, 1123, 1125, 1132, 1145, 1155, 1164, 1167 and 1170 and the DRB1*0308 alleles. Specific PCR fragment of 115 bp in the DRB1*1177 allele.

Primer mix 10: Specific PCR fragment of 65 bp in the DRB1*111301, 111302, 1117, 1126, 1134, 1152 and 1179, the DRB1*0308, 0461, 1411 and the DRB3*0218 alleles. Specific PCR fragment of 85 bp in the DRB1*1107, 1123, 1125, 1145, 1155, 1164 and 1167 alleles. Specific PCR fragment of 120 bp in the DRB1*1174 allele.

Primer mix 11: Specific PCR fragment of 155 bp in the DRB1*1175 allele. Specific PCR fragment of 200 bp in the DRB1*110801, 110802, 1136, 1142, 1157, 1164 and 1168 and the DRB1*1320, 1325, 1329, 1343, 1356, 1360, 1371, 1378, 1388, 140301, 140302, 1412, 1440, 1469, 1474, 1477, 1484 and1485 alleles. Specific PCR fragment of 255 bp in the DRB1*110601, 110602, 1121, 1147 and 1157 and the DRB1*0320, 1358, 1381 and 1429 alleles.

Primer mix 12: Specific PCR fragment of 95 bp in the DRB1*1161 allele. Specific PCR fragment of 185 bp in the DRB1*1130 and the DRB3*0218 alleles.

Primer mix 13: Specific PCR fragment of 70 bp in the DRB1*1160 allele. Specific PCR fragment of 105 bp in the DRB1*1109, 1116, 1120, 112801, 112802, 1140, 1158, 1159, 1183 and the DRB1*0308 alleles.

Primer mix 15: Specific PCR fragment of 60 bp in the DRB1*1166 allele. Specific PCR fragment of 105 bp in the DRB1*111001, 111002, 111201-111302, 1117 and 1152 and the DRB1*1411 alleles.

Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*1131, 1152 and the 0310, 1343, 1345, 140101-140103, 140701-1408, 1416, 1422, 1425, 1426, 143201-143202, 1434, 1435, 1438, 1439, 1449, 1453-1455, 1458, 1460, 1462, 1469, 1470, 1472, 1474, 1475, 1482, 1486-1488, 1490 and 1492N alleles.

Specific PCR fragment of 225 bp in the DRB1*1123, 1125, 1132, 1145, 1155, 1164, 1170 and the 0820, 1313, 1318, 1347, 1355, 140301, 140302, 1412, 1427, 1440, 1455, 1463, 1467, 1477, 1478, 1484, 1485, 1489 alleles.

Primer mix 21: Specific PCR fragment  of 135 bp in the DRB1*1159, 1172, 1180, 1183 and the DRB1*0308 alleles. Specific PCR fragment of 175 bp in the DRB1*1122 and the 0415 and 0461 alleles.

Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*1115 and the 1362 alleles. Specific PCR fragment of 180 bp in the DRB1*1133, 1135 and 116501-116502 alleles. Specific PCR fragment of 215 bp in DRB1*1169, 1182 and DRB1*1309, 1424 and 1437 alleles. Specific PCR fragment of 245 bp in DRB1*1181 allele.

Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*1139, 1143, 1150 and the DRB1*0317, 0337, 1459 and 1486 alleles. Specific PCR fragment of 250 bp in the DRB1*110601-110602, 1121, 1147, 1157 and 1181 and the DRB1*0320, 1358, 1381 and 1429 alleles.

Primer mix 26: Specific PCR fragment of 120 bp in the DRB1*1109-111002, 111301, 111302, 1116, 1117, 1120, 1140, 1146, 1152, 1159, 1162, 1180 and 1183 and the DRB1*0308 and 1411 alleles. Specific PCR fragment of 160 bp in the DRB1*1178 allele.