PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*11

Prod. No:	101.115-24/24u

101.115-03/03u

Lot No:    	07R

Expiry Date:    	2015-March-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 40 primer mixes used for DRB1*11 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified DRB1*11 alleles3 | Other amplified DRB alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 175 bp | 515 bp | *11:01:01-11:04:08, 11:06:01-11:21, 11:23-11:29, 11:31-11:66, 11:68-11:70, 11:72-11:103, 11:105-11:126 | *03:08, 03:65 |
| 2 | 175 bp | 430 bp | *11:05, 11:67 | *08:31, 08:41, 12:04, 14:11 |
| 35 | 65 bp | 430 bp | *11:01:01-11:01:16, 11:03-11:06:02, 11:09-11:12:02, 11:15, 11:22-11:25, 11:27:01-11:30, 11:32-11:33, 11:35, 11:37:01-11:41, 11:43-11:44, 11:46:01-11:47, 11:49:01-11:51, 11:54:01-11:56, 11:58:01-11:63, 11:66-11:67, 11:72, 11:74:01-11:78, 11:80-11:81, 11:83-11:85, 11:87-11:88, 11:90-11:97, 11:99-11:102:02, 11:104, 11:106, 11:108-11:109, 11:111-11:117, 11:120-11:123, 11:126 | *04:15, 08:31, 08:41, DRB5*01:13 |
| 46,8 | 200 bp | 430 bp | *11:02:01-11:02:02, 11:14:01-11:14:02, 11:16, 11:18-11:21, 11:31, 11:45, 11:48, 11:65:01-11:65:02, 11:70, 11:73, 11:86, 11:98, 11:118-11:119, 11:124 | *13:01:01-13:01:06, 13:01:08-13:04, 13:06, 13:08, 13:10, 13:12:01-13:13, 13:15-13:16, 13:19, 13:22-13:23:02, 13:27-13:28, 13:30-13:41, 13:45, 13:48, 13:51-13:53, 13:57-13:59, 13:61:01-13:61:02, 13:64-13:66:02, 13:68-13:70, 13:72-13:73, 13:76-13:77, 13:79-13:85, 13:87, 13:89-13:99, 13:101-13:102, 13:104-13:107, 13:109-13:110, 13:112-13:113N, 13:115, 13:117-13:118, 13:120-13:128, 13:130-13:131, 13:135, 13:137N-13:139, 14:16, 14:63, 14:67, 14:78 |
| 55 | 85 bp | 515 bp | *11:01:01-11:01:16, 11:04:01-11:06:02, 11:08:01-11:10:02, 11:12:01-11:13:02, 11:15, 11:17-11:19:03, 11:22-11:35, 11:37:01-11:39, 11:42-11:47, 11:49:01-11:50, 11:52, 11:54:01-11:58:02, 11:60-11:62, 11:64, 11:66-11:67, 11:72-11:75, 11:77-11:78-11:81, 11:83-11:84:02, 11:87-11:97, 11:99-11:102:02, 11:104, 11:106, 11:108-11:109, 11:111-11:117, 11:119-11:121, 11:123, 11:126 | *04:15, 04:61, 08:31, 08:41, 12:04, 14:11, 16:16, DRB5*01:13 |
| 65,6 | 80 bp | 430 bp | *11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-11:41, 11:48, 11:55, 11:59, 11:63, 11:65:01-11:65:02, 11:68, 11:70, 11:76, 11:85-11:86, 11:118, 11:122, 11:124 |  |
| 75 | 125 bp | 430 bp | *11:01:01-11:01:16, 11:05, 11:08:01-11:12:02, 11:14:01-11:15, 11:19:01-11:20, 11:22-11:24, 11:26-11:33, 11:37:01-11:37:02, 11:39, 11:45, 11:49:01-11:49:02, 11:51, 11:53, 11:61-11:62, 11:64, 11:66, 11:68-11:69, 11:72-11:75, 11:79-11:82, 11:87, 11:90-11:91, 11:93-11:97, 11:99-11:101, 11:103, 11:105-11:109, 11:111-11:112, 11:114, 11:117, 11:120-11:122, 11:125-11:126 | *04:61, 08:41, 16:16, DRB3*02:18, DRB5*01:13 |
| 85 | 125 bp | 430 bp | *11:02:01-11:04:08, 11:06:01-11:07, 11:13:01-11:13:02, 11:16-11:18, 11:21, 11:25, 11:34-11:36, 11:38, 11:40-11:44, 11:46:01-11:48, 11:50, 11:52, 11:54:01-11:60, 11:63, 11:65:01-11:65:02, 11:67, 11:70, 11:76-11:78, 11:83-11:85, 11:88-11:89, 11:92, 11:98, 11:102:01-11:102:02, 11:104, 11:113, 11:115-11:116, 11:118-11:119, 11:123-11:124 | *03:08, 03:65, 04:15, 08:31, 12:04, 14:11 |
| 95,8,12 | 85 bp, 115 bp | 430 bp | *11:07, 11:23, 11:25, 11:32, 11:45, 11:55, 11:64, 11:67, 11:70, 11:77, 11:103, 11:105, 11:107, 11:119, 11:125 | *03:08, 03:65 |
| 105,7,13 | 65 bp, 85 bp, 120 bp | 430 bp | *11:07, 11:13:01-11:13:02, 11:17, 11:23, 11:25-11:26, 11:34, 11:45, 11:52, 11:55, 11:64, 11:67, 11:74:01-11:74:02, 11:79, 11:103, 11:105, 11:107, 11:119, 11:125 | *03:08, 03:65, 04:61, 14:11, DRB3*02:18 |
| 1114 | 155 bp, 200 bp, 255 bp | 515 bp | *11:06:01-11:06:02, 11:08:01-11:08:02, 11:21, 11:36, 11:42, 11:47, 11:57, 11:64, 11:68, 11:75 | *03:20, 13:20, 13:25, 13:29, 13:43, 13:56, 13:58, 13:60, 13:71, 13:78, 13:81, 13:88, 13:119, 13:134, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:29, 14:40, 14:69, 14:74, 14:77, 14:84-14:85, 14:98-14:99, 14:102, 14:115-14:116 |
| 125,15 | 95 bp, 145 bp, 185 bp | 430 bp | *11:30, 11:61, 11:90 | DRB3*02:18 |
| 135,8,9,16 | 70 bp, 105 bp | 430 bp | *11:09, 11:16, 11:20, 11:28:01-11:28:02, 11:40, 11:58:01-11:60, 11:83, 11:87, 11:109, 11:113, 11:115, 11:122 | *03:08, 03:65 |
| 14 | 145 bp | 430 bp | *11:17, 11:37:01-11:37:02, 11:41, 11:48, 11:50, 11:52, 11:56, 11:83, 11:107 | *03:02:01-03:03, 03:06, 03:17, 03:26-03:27, 03:31, 03:35, 03:38, 03:41, 03:60, 03:74-03:75, 03:80, 08:20, 13:03:01-13:03:06, 13:07:01-13:08, 13:12:01-13:13, 13:19, 13:26:01, 13:32-13:33:03, 13:36-13:38, 13:40, 13:47-13:49, 13:53, 13:55, 13:58, 13:60, 13:65, 13:70, 13:72, 13:76, 13:81, 13:84-13:85, 13:88-13:90, 13:94-13:95, 13:101, 13:108, 13:115, 13:118, 13:120, 13:122, 13:133-13:135, 14:01:01-14:03:01, 14:05:01-14:09, 14:12:01-14:14, 14:16, 14:18-14:20, 14:22-14:27, 14:29, 14:32:01-14:32:02, 14:34, 14:36-14:41, 14:43-14:45, 14:47-14:49, 14:51, 14:54:01-14:56, 14:58-14:60, 14:62-14:63, 14:67, 14:69-14:70, 14:74-14:75, 14:77-14:78, 14:80-14:92N, 14:94, 14:96-14:106, 14:108-14:119, 14:121-14:125, DRB3*01:14 |
| 155,17 | 60 bp, 105 bp | 430 bp | *11:10:01-11:10:02, 11:12:01-11:13:02, 11:17, 11:52, 11:66, 11:114 | *14:11 |
| 165,18 | 110 bp, 255 bp | 430 bp | *11:15, 11:24, 11:44, 11:49:01-11:50, 11:76, 11:85, 11:91, 11:114 | *03:26, 03:33, 03:43-03:44, 03:71, 13:34, 13:62, 13:64, 13:68, 13:106, 13:136, 14:41, 14:77, 14:97, 14:110, DRB3*01:14 |
| 175,10 | 90 bp | 430 bp | *11:17, 11:23, 11:25, 11:45, 11:52-11:55, 11:64, 11:67, 11:89, 11:119 | *14:11 |
| 185 | 95 bp | 430 bp | *11:07, 11:27:01-11:27:03, 11:84:01-11:84:02, 11:103, 11:107, 11:124-11:125 | *03:08, 03:65, DRB3*02:18 |
| 196,19 | 180 bp, 225 bp, 260 bp | 430 bp | *11:23, 11:25, 11:31-11:32, 11:45, 11:52, 11:55, 11:64, 11:70, 11:86, 11:96, 11:119 | *03:10, 08:20, 13:13, 13:16, 13:18, 13:43, 13:45, 13:47, 13:55, 13:119, 14:01:01-14:01:03, 14:03:01-14:03:02, 14:07:01-14:08, 14:12:01-14:12:02, 14:16, 14:22, 14:25-14:27, 14:32:01-14:32:02, 14:34-14:35, 14:38-14:40, 14:49, 14:53-14:55, 14:58, 14:60, 14:62-14:63, 14:67, 14:69-14:70, 14:72, 14:74-14:75, 14:77-14:78, 14:82, 14:84-14:90, 14:92N, 14:97, 14:99, 14:102, 14:104-14:105, 14:110-14:119, 14:122, 14:124-14:125 |
| 205,11 | 115 bp | 515 bp | *11:10:01-11:10:02, 11:12:01-11:13:02, 11:17, 11:29, 11:44, 11:47, 11:49:01-11:50, 11:52, 11:63, 11:85 | *03:04:01-03:04:02, 03:21, 03:26, 03:30, 03:33, 03:35, 03:44, 13:08, 13:19, 13:57, 13:60, 13:72, 13:76, 13:83-13:84, 13:87, 13:106-13:107, 13:135, 14:01:01-14:01:03, 14:05:01-14:05:03, 14:07:01-14:08, 14:14, 14:16, 14:20, 14:22-14:23:03, 14:26, 14:32:01-14:32:02, 14:34-14:40, 14:43-14:45, 14:49, 14:54:01-14:56, 14:60, 14:62, 14:65, 14:70, 14:72, 14:74-14:75, 14:82, 14:84, 14:86-14:88, 14:90-14:92N, 14:94, 14:96-14:97, 14:99-14:101, 14:103-14:105, 14:111-14:114, 14:117-14:119, 14:122-14:125 |
| 2120 | 135 bp, 175 bp | 430 bp | *11:22, 11:59, 11:72, 11:80, 11:83, 11:87, 11:92 | *03:08, 03:65, 04:15, 04:61 |
| 225,21 | 115 bp, 180 bp, 215 bp, 245 bp | 430 bp | *11:15, 11:33, 11:35, 11:65:01-11:65:02, 11:69, 11:81-11:82 | *13:09, 13:62, 13:129, 14:24, 14:37 |
| 235 | 120 bp | 430 bp | *11:38 |  |
| 2422 | 155 bp, 250 bp | 430 bp | *11:06:01-11:06:02, 11:21, 11:39, 11:43, 11:47, 11:50, 11:57, 11:81, 11:115 | *03:17, 03:20, 03:37, 03:80, 13:58, 13:81, 13:118, 13:131, 13:135, 14:29, 14:59, 14:86, 14:96, 14:99 |
| 255,8 | 110 bp | 515 bp | *11:01:12, 11:04:04, 11:09-11:10:02, 11:12:02-11:13:02, 11:16-11:17, 11:20, 11:22, 11:28:01-11:28:02, 11:40, 11:46:01, 11:52, 11:58:01-11:59, 11:83, 11:87, 11:113, 11:115, 11:118, 11:122 | *03:08, 03:65, 12:04, 14:11, 16:16, DRB3*02:18 |
| 265,23 | 120 bp, 160 bp | 430 bp | *11:09-11:10:02, 11:13:01-11:13:02, 11:16-11:17, 11:20, 11:40, 11:46:01-11:46:02, 11:52, 11:59, 11:62, 11:78, 11:80, 11:83, 11:87, 11:113, 11:118, 11:122 | *03:08, 03:65, 14:11 |
| 275 | 85 bp | 430 bp | *11:11:01-11:11:02, 11:14:01-11:14:02, 11:20, 11:68, 11:73, 11:93, 11:122 | *01:03, 01:39N, 04:14, 04:106, 08:40, 13:02:01-13:02:03, 13:02:05-13:03:06, 13:23:01-13:23:02, 13:29, 13:31, 13:33:01-13:34, 13:36-13:39, 13:41, 13:45, 13:63, 13:65-13:67, 13:73-13:74, 13:85, 13:88, 13:90, 13:95-13:97:01, 13:99, 13:101, 13:103-13:104, 13:107, 13:115, 13:123-13:124, 13:126, 13:128, 13:135, 13:139 |
| 285 | 85 bp | 430 bp | *11:02:01-11:03, 11:16, 11:21, 11:36, 11:40-11:41, 11:48, 11:55, 11:59, 11:63, 11:65:01-11:65:02, 11:70, 11:76, 11:85, 11:118, 11:124 | *01:42, 04:02:01-04:02:03, 04:37, 13:01:01-13:01:06, 13:01:08-13:01:11, 13:04, 13:08, 13:10, 13:15, 13:17, 13:19-13:20, 13:22, 13:24, 13:27-13:28, 13:32, 13:35, 13:40, 13:43, 13:48, 13:51-13:54, 13:57, 13:59, 13:61:01-13:61:02, 13:64, 13:68-13:72, 13:75-13:76, 13:78-13:81, 13:83-13:84, 13:87, 13:89, 13:91-13:94, 13:98, 13:102, 13:105-13:106, 13:110-13:114, 13:116-13:117, 13:121, 13:125, 13:127, 13:130-13:131, 13:138, 14:16, 14:57, 15:10 |
| 295 | 100 bp | 515 bp | *11:14:01-11:14:02, 11:19:01-11:20, 11:31, 11:45, 11:73, 11:82 | *01:03, 01:15, 01:24, 01:39N, 04:14, 04:38, 04:82, 04:86, 04:105:01-04:107, 07:01:01:01-07:01:04, 07:03-07:04, 07:06-07:19, 07:21-07:22, 08:03:02-08:03:03, 08:14-08:15, 08:18-08:19, 08:23, 08:25, 08:27, 08:29-08:30:03, 08:32-08:38, 08:40, 08:45-08:47, 08:49, 12:22, 13:02:01-13:02:03, 13:02:05-13:03:06, 13:12:01-13:13, 13:23:01-13:23:02, 13:30-13:31, 13:33:01-13:34, 13:36-13:39, 13:41, 13:45, 13:65-13:67, 13:73-13:74, 13:82, 13:85, 13:90, 13:95-13:97:01, 13:99, 13:101, 13:103-13:104, 13:107, 13:115, 13:118, 13:120, 13:122-13:124, 13:126, 13:128, 13:133, 13:135, 13:139, 14:24, 14:63, 14:67, 15:02:01-15:02:08, 15:11, 15:14, 15:19, 15:26, 15:29-15:30, 15:38-15:39, 15:44, 15:47, 15:58, 15:60, 15:68, 16:05:01-16:05:02, 16:07, DRB3*02:11, DRB5*01:07, DRB5*01:11, DRB5*02:03 |
| 305 | 100 bp | 515 bp | *11:01:01-11:01:16, 11:03-11:06:02, 11:09-11:12:02, 11:15, 11:22-11:25, 11:27:01-11:30, 11:32-11:33, 11:35, 11:38-11:40, 11:43-11:44, 11:46:01-11:47, 11:49:01-11:49:02, 11:51, 11:54:01-11:55, 11:58:01-11:63, 11:66-11:67, 11:72, 11:74:01-11:78, 11:80-11:81, 11:84:01-11:85, 11:87, 11:90-11:97, 11:99-11:102:02, 11:104, 11:106, 11:108-11:117, 11:120-11:123, 11:126 | *08:17, 08:28, 12:02:01-12:02:05, 12:13, 12:15-12:16:01, 12:18-12:20, 12:23, 12:26-12:27, 12:31N-12:33, 13:05:01-13:05:02, 13:11:01-13:11:02, 13:14:01-13:14:03, 13:18, 13:21:01-13:21:02, 13:24, 13:42, 13:46, 13:50:01-13:50:02, 13:54, 13:62-13:63, 13:75, 13:100, 13:111, 13:132, 13:136, 14:53, 16:09 |
| 315 | 85 bp | 515 bp | *11:04:01-11:04:08, 11:06:01-11:06:02, 11:18, 11:25, 11:35, 11:38, 11:42-11:44, 11:46:01-11:47, 11:50, 11:54:01-11:54:02, 11:56-11:58:02, 11:60, 11:67, 11:77-11:78, 11:83-11:84:02, 11:88, 11:92, 11:102:01-11:102:02, 11:104, 11:113, 11:115-11:116, 11:119, 11:123 | *01:23, 04:12, 04:15, 04:18, 04:25, 04:36, 04:53, 04:58, 04:99, 08:04:01, 08:04:04-08:04:07, 08:06, 08:10, 08:12, 08:20, 08:22, 08:28, 08:31, 12:01:01-12:15, 12:17-12:21, 12:23-12:35, 13:06, 13:11:01-13:11:02, 13:18, 13:42, 13:58, 13:77, 13:108, 14:12:01-14:12:02, 14:15, 14:73-14:74, 14:78, 14:84, 14:105, 15:21, 16:15 |
| 325,11 | 75 bp | 430 bp | *11:13:01-11:13:02, 11:17, 11:51-11:52, 11:89 | *14:11 |
| 338 | 155 bp | 430 bp | *11:88, 11:94 |  |
| 345,24 | 80 bp, 290 bp | 515 bp | *11:31, 11:96-11:97 | *03:10, 13:43, 13:45, 14:35, 14:53, 14:72 |
| 355 | 85 bp | 430 bp | *11:98 |  |
| 365,25 | 115 bp, 170 bp | 430 bp | *11:99, 11:102:01-11:102:02, 11:110 | *03:49 |
| 37 | 150 bp | 430 bp | *11:101 |  |
| 38 | 155 bp | 430 bp | *11:78, 11:106 |  |
| 395 | 110 bp | 430 bp | *11:108 |  |
| 40 | 180 bp | 430 bp | *11:116, 11:120 | *03:13:01-03:13:02, 03:38, 03:79, 13:39, 13:54, 13:77-13:79, 13:100, 14:48, 14:64, 14:109, 14:115, DRB3*01:14 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*11 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*11 subtyping.

In addition, wells number 5, 11, 20, 25, 29 to 31 and 34 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence.

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, other DRB1 alleles are amplified by primer mixes 1 to 5, 7 to 11, 13 to 22, 24 to 32, 34, 36 and 40. In addition, the DRB3*01:14 allele is amplified by primer mixes 14, 16 and 40, the DRB3*02:11, DRB5*01:07, DRB5*01:11 and DRB5*02:03 alleles are amplified by primer mix 29, the DRB3*02:18 allele is amplified by primer mixes 7, 10, 12, 18 and 25 and the DRB5*01:13 allele is amplified by primer mixes 3, 5 and 7.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands. In order to have as little co-amplification as possible of non-DRB1*11 alleles by the primer pairs of the DRB1*11 subtyping kit, the DRB1*11 characteristic sequence motif in codon 58, ‘Ag’, has been used in either the 5’-primer or the 3’-primer in many of the DRB1*11 primer mixes and consequently many of the specific PCR products are shorter than 100 base pairs.

6Primer mix 4 may faintly amplify the DRB1*13:29 allele.

7Primer mix 10 may faintly amplify the DRB1*11:53 allele.

8Primer mixes 4, 9, 13, 25 and 33 may give rise to a primer oligomer artifact.

9Primer mixes 6, 13 and 19 may give rise to nonspecific amplification.

10In primer mix 17 the yield of HLA-specific PCR product may be lower than for other DRB1*11 primer mixes

11Primer mixes 20 and 32 might generate a long unspecific PCR product. This band should be disregarded when interpreting the DRB1*11 typings.

12Primer mix 9: Specific PCR fragment of 85 bp in the DRB1*11:07, 11:23, 11:25, 11:32, 11:45, 11:55, 11:64, 11:67, 11:70, 11:103, 11:105, 11:107, 11:119 and 11:125 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 115 bp in the DRB1*11:77 allele.

13Primer mix 10: Specific PCR fragment of 65 bp in the DRB1*11:13:01-11:13:02, 11:17, 11:26, 11:34, 11:52 and 11:79 and the DRB1*04:61 and 14:11 and in the DRB3*02:18 alleles. Specific PCR fragment of 85 bp in the DRB1*11:23, 11:25, 11:45, 11:55, 11:64, 11:67, 11:119 and 11:125 and the DRB1*03:65 alleles. Specific PCR fragment of 120 bp in the DRB1*11:74 allele.PCR fragment of 65 and 85 bp in the DRB1*11:07 11:103, 11:105 11:107 and in the DRB1*03:08 alleles.

14Primer mix 11: Specific PCR fragment of 155 bp in the DRB1*11:75 allele. Specific PCR fragment of 200 bp in the DRB1*11:08:01-11:08:02, 11:36, 11:42, 11:64 and 11:68 and the DRB1*13:20, 13:25, 13:29, 13:43, 13:56, 13:60, 13:71, 13:78, 13:88, 13:119, 13:134, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:40, 14:69, 14:74, 14:77, 14:84-14:85, 14:98, 14:102 and 14:115-14:116 alleles. Specific PCR fragment of 255 bp in the DRB1*11:06:01-11:06:02, 11:21 and 11:47 and the DRB1*03:20, 13:58, 13:81, 14:29 and 14:99 alleles. Specific PCR fragment of 200 and 255 bp in the 11:57 allele.

15Primer mix 12: Specific PCR fragment of 95 bp in the DRB1*11:61 allele. Specific PCR fragment of 145 bp in the DRB1*11:90 allele. Specific PCR fragment of 185 bp in the DRB1*11:30 and the DRB3*02:18 alleles.

16Primer mix 13: Specific PCR fragment of 70 bp in the DRB1*11:60 and 11:109 alleles. Specific PCR fragment of 105 bp in the DRB1*11:09, 11:16, 11:20, 11:28:01-11:28:02, 11:40, 11:58:01-11:59, 11:83, 11:87, 11:113, 11:115, 11:122 and the DRB1*03:08 and 03:65 alleles.

17Primer mix 15: Specific PCR fragment of 60 bp in the DRB1*11:66 allele. Specific PCR fragment of 105 bp in the DRB1*11:10:01-11:10:02, 11:12:01-11:13:02, 11:17, 11:52 and 11:114 and the DRB1*14:11 alleles.

18Primer mix 16: Specific PCR fragment of 110 bp in the DRB1*11:15, 11:24, 11:44, 11:49:01-11:50, 11:76, 11:85 and 11:114 and the DRB1*03:26, 03:33, 03:43-03:44, 03:71, 13:34, 13:62, 13:64, 13:68, 13:106, 13:136, 14:41, 14:77, 14:97 and 14:110 and in the DRB3*01:14 alleles. Specific PCR fragment of 255 bp in the DRB1*11:91 allele.

19Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*11:31, 11:52 and 11:96 and the DRB1*03:10, 13:43, 13:45, 14:01:01-14:01:03, 14:07:01-14:08, 14:16, 14:22, 14:25-14:26, 14:32:01-14:32:02, 14:34-14:35, 14:38-14:39, 14:49, 14:53-14:54, 14:58, 14:60, 14:62, 14:69-14:70, 14:72, 14:74-14:75, 14:82, 14:86-14:88, 14:90, 14:92N, 14:97, 14:99, 14:104-14:105, 14:110-14:114, 14:117-14:119, 14:122 and 14:124-14:125 alleles. Specific PCR fragment of 225 bp in the DRB1*11:23, 11:25, 11:32, 11:45, 11:55, 11:64, 11:70 and 11:119 and the DRB1*08:20, 13:13, 13:18, 13:47, 13:55, 13:119, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:27, 14:40, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89, 14:102 and 14:115-14:116 alleles. Specific PCR fragment of 260 bp in the DRB1*11:86 and the DRB1*13:16 alleles. Specific PCR fragment of 180 and 225 bp in the 14:55 allele.

20Primer mix 21: Specific PCR fragment of 135 bp in the DRB1*11:59, 11:72, 11:80, 11:83, 11:87 and 11:92 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 175 bp in the DRB1*11:22 and in the DRB1*04:15 and 04:61 alleles.

21Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*11:15 and the DRB1*13:62 alleles. Specific PCR fragment of 180 bp in the DRB1*11:33, 11:35 and 11:65:01-11:65:02 alleles. Specific PCR fragment of 215 bp in DRB1*11:69, 11:82 and DRB1*13:09, 13:129, 14:24 and 14:37 alleles. Specific PCR fragment of 245 bp in DRB1*11:81 allele.

22Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*11:39, 11:43, 11:50 and 11:115 and the DRB1*03:17, 03:37, 03:80, 13:118, 13:131, 13:135, 14:59, 14:86, 14:96 alleles. Specific PCR fragment of 250 bp in the DRB1*11:06:01-11:06:02, 11:21, 11:47, 11:57, 11:81 and the DRB1*03:20, 13:58, 13:81, 14:29 and 14:99 alleles.

23Primer mix 26: Specific PCR fragment of 120 bp in the DRB1*11:09-11:10:02, 11:13:01-11:13:02, 11:16-11:17, 11:20, 11:40, 11:46:01-11:46:02, 11:52, 11:59, 11:62, 11:80, 11:83, 11:87, 11:113, 11:118 and 11:122 and in the DRB1*03:08, 03:65 and 14:11 alleles. Specific PCR fragment of 160 bp in the DRB1*11:78 allele.

24Primer mix 34: Specific PCR fragment of 80 bp in the DRB1*11:31 and 11:96 and in the DRB1*03:10, 13:43, 13:45, 14:35, 14:53 and 14:72 alleles. Specific PCR fragment of 290 bp in the DRB1*11:97 allele.

25Primer mix 36: Specific PCR fragment of 115 bp in the DRB1*11:99 and 11:102:01-11:102:02 and the DRB1*03:49 alleles. Specific PCR fragment of 170 bp in the DRB1*11:110 allele.

‘w’, might be weakly amplified.

1The DRB1*1171 has been shown to be identical to corrected DRB1*11:02:01.

2DRB1*11 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.8.0, April 2012.

3The DRB1*11:24 and 11:91 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 16.

4The DRB1*11:61 and 11:90 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 12.

5Primer mix 9: Specific PCR fragment of 85 bp in the DRB1*11:07, 11:23, 11:25, 11:32, 11:45, 11:55, 11:64, 11:67, 11:70, 11:103, 11:105, 11:107, 11:119 and 11:125 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 115 bp in the DRB1*11:77 allele.

Primer mix 10: Specific PCR fragment of 65 bp in the DRB1*11:13:01-11:13:02, 11:17, 11:26, 11:34, 11:52 and 11:79 and the DRB1*04:61 and 14:11 and in the DRB3*02:18 alleles. Specific PCR fragment of 85 bp in the DRB1*11:23, 11:25, 11:45, 11:55, 11:64, 11:67, 11:119 and 11:125 and the DRB1*03:65 alleles. Specific PCR fragment of 120 bp in the DRB1*11:74 allele.PCR fragment of 65 and 85 bp in the DRB1*11:07 11:103, 11:105 11:107 and in the DRB1*03:08 alleles.

Primer mix 11: Specific PCR fragment of 155 bp in the DRB1*11:75 allele. Specific PCR fragment of 200 bp in the DRB1*11:08:01-11:08:02, 11:36, 11:42, 11:64 and 11:68 and the DRB1*13:20, 13:25, 13:29, 13:43, 13:56, 13:60, 13:71, 13:78, 13:88, 13:119, 13:134, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:40, 14:69, 14:74, 14:77, 14:84-14:85, 14:98, 14:102 and 14:115-14:116 alleles. Specific PCR fragment of 255 bp in the DRB1*11:06:01-11:06:02, 11:21 and 11:47 and the DRB1*03:20, 13:58, 13:81, 14:29 and 14:99 alleles. Specific PCR fragment of 200 and 255 bp in the 11:57 allele.

Primer mix 12: Specific PCR fragment of 95 bp in the DRB1*11:61 allele. Specific PCR fragment of 145 bp in the DRB1*11:90 allele. Specific PCR fragment of 185 bp in the DRB1*11:30 and the DRB3*02:18 alleles.

Primer mix 13: Specific PCR fragment of 70 bp in the DRB1*11:60 and 11:109 alleles. Specific PCR fragment of 105 bp in the DRB1*11:09, 11:16, 11:20, 11:28:01-11:28:02, 11:40, 11:58:01-11:59, 11:83, 11:87, 11:113, 11:115, 11:122 and the DRB1*03:08 and 03:65 alleles.

Primer mix 15: Specific PCR fragment of 60 bp in the DRB1*11:66 allele. Specific PCR fragment of 105 bp in the DRB1*11:10:01-11:10:02, 11:12:01-11:13:02, 11:17, 11:52 and 11:114 and the DRB1*14:11 alleles.

Primer mix 16: Specific PCR fragment of 110 bp in the DRB1*11:15, 11:24, 11:44, 11:49:01-11:50, 11:76, 11:85 and 11:114 and the DRB1*03:26, 03:33, 03:43-03:44, 03:71, 13:34, 13:62, 13:64, 13:68, 13:106, 13:136, 14:41, 14:77, 14:97 and 14:110 and in the DRB3*01:14 alleles. Specific PCR fragment of 255 bp in the DRB1*11:91 allele.

Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*11:31, 11:52 and 11:96 and the DRB1*03:10, 13:43, 13:45, 14:01:01-14:01:03, 14:07:01-14:08, 14:16, 14:22, 14:25-14:26, 14:32:01-14:32:02, 14:34-14:35, 14:38-14:39, 14:49, 14:53-14:54, 14:58, 14:60, 14:62, 14:69-14:70, 14:72, 14:74-14:75, 14:82, 14:86-14:88, 14:90, 14:92N, 14:97, 14:99, 14:104-14:105, 14:110-14:114, 14:117-14:119, 14:122 and 14:124-14:125 alleles. Specific PCR fragment of 225 bp in the DRB1*11:23, 11:25, 11:32, 11:45, 11:55, 11:64, 11:70 and 11:119 and the DRB1*08:20, 13:13, 13:18, 13:47, 13:55, 13:119, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:27, 14:40, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89, 14:102 and 14:115-14:116 alleles. Specific PCR fragment of 260 bp in the DRB1*11:86 and the DRB1*13:16 alleles. Specific PCR fragment of 180 and 225 bp in the 14:55 allele.

Primer mix 21: Specific PCR fragment of 135 bp in the DRB1*11:59, 11:72, 11:80, 11:83, 11:87 and 11:92 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 175 bp in the DRB1*11:22 and in the DRB1*04:15 and 04:61 alleles.

Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*11:15 and the DRB1*13:62 alleles. Specific PCR fragment of 180 bp in the DRB1*11:33, 11:35 and 11:65:01-11:65:02 alleles. Specific PCR fragment of 215 bp in DRB1*11:69, 11:82 and DRB1*13:09, 13:129, 14:24 and 14:37 alleles. Specific PCR fragment of 245 bp in DRB1*11:81 allele.

Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*11:39, 11:43, 11:50 and 11:115 and the DRB1*03:17, 03:37, 03:80, 13:118, 13:131, 13:135, 14:59, 14:86, 14:96 alleles. Specific PCR fragment of 250 bp in the DRB1*11:06:01-11:06:02, 11:21, 11:47, 11:57, 11:81 and the DRB1*03:20, 13:58, 13:81, 14:29 and 14:99 alleles.

Primer mix 26: Specific PCR fragment of 120 bp in the DRB1*11:09-11:10:02, 11:13:01-11:13:02, 11:16-11:17, 11:20, 11:40, 11:46:01-11:46:02, 11:52, 11:59, 11:62, 11:80, 11:83, 11:87, 11:113, 11:118 and 11:122 and in the DRB1*03:08, 03:65 and 14:11 alleles. Specific PCR fragment of 160 bp in the DRB1*11:78 allele.

Primer mix 34: Specific PCR fragment of 80 bp in the DRB1*11:31 and 11:96 and in the DRB1*03:10, 13:43, 13:45, 14:35, 14:53 and 14:72 alleles. Specific PCR fragment of 290 bp in the DRB1*11:97 allele.

Primer mix 36: Specific PCR fragment of 115 bp in the DRB1*11:99 and 11:102:01-11:102:02 and the DRB1*03:49 alleles. Specific PCR fragment of 170 bp in the DRB1*11:110 allele.

‘w’, might be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 5 amplifies the DRB1*11:73, 11:79 and 11:93 alleles. This has been corrected in the Specificity and Interpretation tables.