Gel Documentation Form and Worksheet

DRB1*11 (101.115-24/03, -24u/03u) 	Lot No: 38X	Expiry Date: 2017-July-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                          Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

In order to have as little co-amplification as possible of non-DRB1*11 alleles by the primer pairs of the DRB1*11 subtyping kit, the DRB1*11 characteristic sequence motif in codon 58, ‘Ag’, has been used in either the 5’-primer or the 3’-primer in many of the DRB1*11 primer mixes and consequently many of the specific PCR products are shorter than 100 base pairs.

Primer mix 4 may faintly amplify the DRB1*13:29 allele.

Primer mix 10 may faintly amplify the DRB1*11:53 allele.

Primer mixes 4, 13, 25, 30, 33, 41 and 46 have a tendency to giving rise to primer oligomer formation.

Primer mixes 6, 13, 19, 32 and 44 may have tendencies of unspecific amplifications.

Primer mix 17 may give rise to a lower yield of HLA-specific PCR product than the other DRB1*11 primer mixes.

Primer mix 47 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1DRB1*11 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.17.0, July 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 9: Specific PCR fragment of 85 bp in the DRB1*11:07, 11:103, 11:105, 11:107 and 11:125 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 115 bp in the DRB1*11:77 allele.

Primer mix 10: Specific PCR fragment of 85 bp in the DRB1*11:07, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:67, 11:103, 11:105, 11:107, 11:119 and 11:125 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 120 bp in the DRB1*11:74:01-11:74:02 alleles.

Primer mix 11: Specific PCR fragment of 155 bp in the DRB1*11:75 allele. Specific PCR fragment of 200 bp in the DRB1*11:08:01-11:08:03, 11:36, 11:42, 11:64, 11:68 and 11:130 and the DRB1*13:20, 13:25, 13:29, 13:43, 13:56, 13:60, 13:71, 13:78, 13:88, 13:119, 13:134, 13:149, 13:180, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:40, 14:69, 14:74, 14:77, 14:84-14:85, 14:98, 14:102, 14:115-14:116, 14:135, 14:144 and 14:156 alleles. Specific PCR fragment of 255 bp in the DRB1*11:06:01-11:06:03, 11:21, 11:47 and 11:129 and the DRB1*03:20, 13:58, 13:81, 13:163, 14:29 and 14:99 alleles. Specific PCR fragment of 200 and 255 bp in the 11:57 allele.

Primer mix 12: Specific PCR fragment of 145 bp in the DRB1*11:90 allele. Specific PCR fragment of 185 bp in the DRB1*11:30 and in the DRB3*02:18 alleles. Primer mix 12 may amplify the DRB1*11:60 and 11:109 and the DRB1*15:68 alleles.

Primer mix 16: Specific PCR fragment of 110 bp in the DRB1*11:15, 11:24, 11:44, 11:49:01-11:50, 11:76, 11:85, 11:114 and 11:127 and the DRB1*03:26, 03:33, 03:43-03:44, 03:71, 13:34, 13:62, 13:64, 13:68, 13:106, 13:136, 13:151, 13:174, 14:41, 14:77, 14:97 and 14:110 and in the DRB3*01:14 alleles. Specific PCR fragment of 255 bp in the DRB1*11:91 allele.

Primer mix 19: Specific PCR fragment of 180 bp in the DRB1*11:31, 11:52, 11:96 and 11:148 and the DRB1*03:10, 13:43, 13:45, 13:159, 13:179, 14:01:01-14:01:04, 14:07:01-14:08, 14:16, 14:22, 14:25-14:26, 14:32:01-14:32:03, 14:34-14:35, 14:38-14:39, 14:49, 14:53-14:54, 14:58, 14:60, 14:62, 14:69-14:70, 14:72, 14:74-14:75, 14:82, 14:86-14:88, 14:90, 14:92N, 14:97, 14:99, 14:104-14:105, 14:110-14:114, 14:117-14:119, 14:122, 14:124-14:125, 14:128-14:133, 14:137N, 14:139-14:140, 14:142, 14:146-14:147, 14:149-14:151, 14:153 and 14:157 alleles. Specific PCR fragment of 225 bp in the DRB1*11:23:01-11:23:02, 11:25, 11:32, 11:45, 11:55, 11:64, 11:70 and 11:119 and the DRB1*08:20, 13:13, 13:18, 13:47, 13:55, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158, 13:164, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:27, 14:40, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89, 14:102, 14:115-14:116, 14:135-14:136, 14:144 and 14:156 alleles. Specific PCR fragment of 260 bp in the DRB1*11:86 and the DRB1*13:16 alleles. Specific PCR fragment of 180 and 225 bp in the 14:55 allele.

Primer mix 21: Specific PCR fragment of 135 bp in the DRB1*11:59, 11:80, 11:83, 11:87, 11:135 and 11:142 and the DRB1*03:08 and 03:65 alleles. Specific PCR fragment of 175 bp in the DRB1*11:22 and the DRB1*04:15, 04:61, 04:138, 04:145, 04:154 and 04:170 alleles.

Primer mix 22: Specific PCR fragment of 115 bp in the DRB1*11:15 and the DRB1*13:62 and 13:151 alleles. Specific PCR fragment of 180 bp in the DRB1*11:33, 11:35, 11:65:01-11:65:02 and 11:130 alleles. Specific PCR fragment of 215 bp in DRB1*11:69 and 11:82 and the DRB1*13:09, 13:129, 14:24 and 14:37 alleles. Specific PCR fragment of 245 bp in DRB1*11:81 allele.

Primer mix 24: Specific PCR fragment of 155 bp in the DRB1*11:39, 11:43, 11:50, 11:115, 11:145 and 11:151 and the DRB1*03:17, 03:37, 03:80, 13:118, 13:131, 13:135, 13:150, 14:59, 14:86, 14:96 and 14:153-14:154 alleles. Specific PCR fragment of 250 bp in the DRB1*11:06:01-11:06:03, 11:21, 11:47, 11:57, 11:81 and 11:129 and the DRB1*03:20, 13:58, 13:81, 13:163, 14:29 and 14:99 alleles.

Primer mix 26: Specific PCR fragment of 120 bp in the DRB1*11:09-11:10:02, 11:13:01-11:13:02, 11:16-11:17, 11:20, 11:40, 11:46:01-11:46:02, 11:52, 11:59, 11:62:01-11:62:02, 11:80, 11:83, 11:87, 11:113, 11:118, 11:122 and 11:149 and the DRB1*03:08, 03:65, 14:11 and 14:148 alleles. Specific PCR fragment of 160 bp in the DRB1*11:78 and 11:152 alleles.

Primer mix 34: Specific PCR fragment of 80 bp in the DRB1*11:31, 11:96 and 11:148 and the DRB1*03:10, 13:43, 13:45, 13:179, 14:35, 14:53, 14:72, 14:128, 14:132 and 14:138 alleles. Specific PCR fragment of 290 bp in the DRB1*11:97 allele.

Primer mix 36: Specific PCR fragment of 110 bp in the DRB1*11:99 and 11:102:01-11:102:02 and the DRB1*03:49 alleles. Specific PCR fragment of 170 bp in the DRB1*11:110 allele. Specific PCR fragment of 255 bp in the DRB1*11:128 alleles.

Primer mix 42: Specific PCR fragment of 135 bp in the DRB1*11:72 and 11:92 alleles. Specific PCR fragment of 180 bp in the DRB1*11:104 alleles.

4The following DRB1*11 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| DRB1*11:24, 11:91 | 16 |
| DRB1*11:99, 11:128 | 36 |

The DRB1*11 subtyping kit cannot distinguish the silent mutations in the DRB1*11:01:01-11:01:11 and 11:01:13-11:01:24, the DRB1*11:02:01-11:02:02, the DRB1*11:04:01-11:04:03 and 11:04:05-11:04:11, the DRB1*11:06:01-11:06:03, the DRB1*11:08:01-11:08:03, the DRB1*11:10:01-11:10:02, the DRB1*11:11:01-11:11:03, the DRB1*11:13:01-11:13:02, the DRB1*11:14:01-11:14:02, the DRB1*11:19:01-11:19:03, the DRB1*11:23:01-11:23:02, the DRB1*11:27:01-11:27:03, the DRB1*11:28:01-11:28:02, the DRB1*11:37:01-11:37:02, the DRB1*11:49:01-11:49:02, the DRB1*11:54:01-11:54:02, the DRB1*11:58:01-11:58:02,  the DRB1*11:65:01-11:65:02, the DRB1*11:74:01-11:74:02,  the DRB1*11:84:01-11:84:03 or the DRB1*11:102:01-11:102:02 alleles.

5This lot of the DRB1*11 subtyping kit cannot distinguish the DRB1*11:01:01-11:01:11 and 11:01:13-11:01:24 alleles and the DRB1*11:66 allele, or the DRB1*11:01:01-11:01:11 and 11:01:13-11:01:24 alleles and the DRB1*11:61 allele.

Change in revision R01 compared to R00:

- Primer mix 5 amplifies the DRB1*11:73, 11:79, 11:93, 11:131, 11:132 and 11:153 and the DRB1*04:145 alleles. This has been corrected in the Specificity and Interpretation tables.

Change in revision R02 compared to R01:

- Primer mix 44 may have tendencies of unspecific amplifications. A footnote has been added in the Specificity Table

Change in revision R03 compared to R02:

- Primer mix 15 does not amplify the DRB1*11:66 allele. This has been corrected in the Interpretation and Specificity Tables. Thus, Thus, this lot of the DRB1*11 subtyping kit cannot distinguish the DRB1*11:01:01-11:01:11 and 11:01:13-11:01:24 alleles and the DRB1*11:66 allele.

Change in revision R04 compared to R03:

- Primer mix 12 does not amplify the DRB1*11:61 allele. This has been corrected in the Interpretation and Specificity Tables. Thus, this lot of the DRB1*11 subtyping kit cannot distinguish the DRB1*11:01:01-11:01:11 and 11:01:13-11:01:24 alleles and the DRB1*11:61 allele.