PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*12 Add-on

Prod. No:	101.815-12/12u

Lot No:    	31N

Expiry Date:    	2014-August-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 2 primer mixes used for DRB1*12 Add-on SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of controlband2 | Amplified DRB1*12:01/12:10 alleles3 | Other amplified DRB alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 215 bp | 515 bp | *12:01:01-12:01:04, 12:10 | 12:02:01-12:02:05, 12:04-12:07, 12:09, 12:11-12:12, 12:13w, 12:14-12:15, 12:17-12:18, 12:20-12:21, 12:24N-12:26, 12:28-12:35 |
| 24 | 80 bp | 430 bp | *12:10 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*14 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The two different control primer pairs give rise to an internal positive control band of either 430 or 515 base pairs.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*12 Add-on subtyping.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial nucleotide sequences from the second exon are available. In these instances it is not known whether some of the primers of the SSP set are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*01:01 consensus sequence.

4Due to the sharing of sequence motifs many DRB1*12 alleles are amplified by primer mix 1.

5Short specific PCR fragments have a lower intensity than longer PCR bands.

‘w’, might be weakly amplified.

‘w’, may be weakly amplified.