PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® DRB1*14

Prod. No:	101.117-12/12u

101.117-03/03u

Lot No:    	76K

Expiry Date:    	2013-May-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 48 primer mixes used for DRB1*14 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of contr. band2 | Amplified DRB1*14 alleles3 | Other amplified DRB alleles3,4 |
| --- | --- | --- | --- | --- |
| 1 | 210 bp | 515 bp | *14:01:01-14:01:03, 14:05:01-14:05:03, 14:07:01-14:08, 14:14, 14:18, 14:23:01-14:23:03, 14:26, 14:32:01-14:32:02, 14:34-14:36, 14:38-14:39, 14:42-14:45, 14:54, 14:55w, 14:56, 14:58-14:60, 14:62, 14:64-14:65, 14:70, 14:72, 14:75, 14:81-14:82, 14:86-14:88, 14:90-14:92N, 14:95-14:97, 14:99-14:100 | *11:13:01-11:13:02, 11:17, 11:52, 11:89 |
| 28 | 175 bp | 430 bp | *14:01:01-14:01:03, 14:07:01-14:07:02, 14:16, 14:22, 14:25-14:26, 14:32:01-14:32:02, 14:35, 14:38-14:39, 14:49, 14:53-14:55, 14:58, 14:60, 14:62, 14:69-14:70, 14:74-14:75, 14:82, 14:86-14:88, 14:90, 14:92N, 14:97, 14:99 | *03:10, 13:43, 13:45 |
| 35,6 | 140 bp | 430 bp | *14:02, 14:06:01-14:06:02, 14:09, 14:13, 14:17, 14:19, 14:21, 14:29-14:30, 14:33, 14:46-14:48, 14:51-14:52, 14:80, 14:83 | *13:10, 13:85 |
| 45,7 | 85 bp | 430 bp | *14:02-14:03:02, 14:06:01-14:06:02, 14:12:01-14:13, 14:18-14:20, 14:24, 14:27, 14:29, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:67, 14:77-14:78, 14:81, 14:83, 14:85, 14:89, 14:94, 14:98 | *03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 13:15, 13:19, 13:26, 13:53, 13:57, 13:85-13:86 |
| 5 | 260 bp | 430 bp | *14:02-14:03:02, 14:07:01-14:07:02, 14:09, 14:13-14:14, 14:19, 14:22, 14:24-14:25, 14:27, 14:30, 14:36, 14:40-14:42, 14:44:01-14:44:02, 14:47-14:49, 14:51, 14:53, 14:63, 14:67, 14:69, 14:77, 14:85, 14:89, 14:94, 14:98 | *03:02:01-03:02:02, 03:05:01-03:05:03, 03:09, 03:14, 03:17, 03:27, 03:29, 03:35, 03:38, 03:40-03:41, 03:53, 11:01:01-11:01:12, 11:08:01-11:12:02, 11:14:01-11:15, 11:19:01-11:20, 11:23-11:24, 11:26-11:29, 11:31-11:33, 11:37, 11:39, 11:45, 11:49, 11:51, 11:53, 11:61-11:62, 11:64, 11:66, 11:68-11:69, 11:72-11:75, 11:79-11:82, 11:87, 11:90-11:91, 11:93-11:97, 13:02:01-13:02:03, 13:03:01-13:03:06, 13:05:01-13:05:02, 13:07:01-13:07:02, 13:12-13:14:03, 13:21:01-13:21:02, 13:23, 13:25-13:26, 13:29-13:31, 13:33:01-13:34, 13:36-13:39, 13:41, 13:45-13:47, 13:49-13:50:02, 13:55-13:56, 13:60, 13:62-13:63, 13:65-13:66:02, 13:73-13:74, 13:82, 13:85-13:86, 13:88, 13:90, 13:95-13:97, 13:99-13:101, DRB3*01:14 |
| 68 | 260 bp | 515 bp | *14:01:01-14:01:03, 14:05:01-14:06:02, 14:08, 14:12:01-14:12:02, 14:16-14:18, 14:20-14:21, 14:23:01-14:23:03, 14:26, 14:29, 14:32:01-14:35, 14:37-14:39, 14:43, 14:45, 14:54-14:56, 14:58-14:60, 14:62, 14:64-14:65, 14:70, 14:72, 14:74-14:75, 14:78, 14:80-14:84, 14:86-14:88, 14:90-14:92N, 14:95-14:97, 14:99-14:100 | *03:01:01:01-03:01:09, 03:03-03:04, 03:06-03:08, 03:10-03:13:02, 03:15-03:16, 03:18-03:26, 03:28, 03:30-03:34, 03:36-03:37, 03:39, 03:43-03:52, 03:54-03:56, 03:58, 08:20, 11:02:01-11:04:06, 11:06:01-11:07, 11:13:01-11:13:02, 11:16-11:18, 11:21, 11:25, 11:34-11:36, 11:38, 11:40-11:44, 11:46-11:48, 11:50, 11:52, 11:54:01-11:60, 11:63, 11:65:01-11:65:02, 11:70, 11:76-11:78, 11:83-11:85, 11:88-11:89, 11:92, 11:98, 13:01:01-13:01:08, 13:04, 13:06, 13:08-13:11:02, 13:15, 13:18-13:20, 13:22, 13:24, 13:27-13:28, 13:32, 13:35, 13:40, 13:42-13:44, 13:48, 13:51-13:54, 13:57-13:59, 13:61, 13:64, 13:68-13:72, 13:75-13:81, 13:83-13:84, 13:87, 13:89, 13:91-13:94, 13:98, 13:102 |
| 7 | 225 bp | 430 bp | *14:03:01-14:03:02, 14:12:01-14:12:02, 14:27, 14:40, 14:55, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89 | *08:20, 11:23, 11:25, 11:45, 11:55, 11:64, 13:13, 13:18, 13:47, 13:55 |
| 8 | 170 bp | 430 bp | *14:04, 14:28, 14:31, 14:50, 14:61, 14:68, 14:71, 14:73, 14:76, 14:79, 14:93 | *08:08, 08:11w |
| 95,10 | 130 bp, 250 bp | 430 bp | *14:05:01-14:05:03, 14:37, 14:43, 14:45, 14:56, 14:62, 14:65, 14:78, 14:84, 14:91, 14:96, 14:100 | *03:12, 04:10-04:12, 04:67, 04:91, 08:06, 08:10, 08:12, 08:22, 13:04, 13:32, 13:48, 13:58, 13:72, 13:75, 13:81, 13:89, 13:93-13:94, 15:12 |
| 106 | 175 bp | 430 bp | *14:10, 14:39, 14:57 | *04:62, 04:69 |
| 115,11 | 85 bp, 170 bp, 215 bp | 430 bp | *14:11, 14:16, 14:22, 14:25, 14:31-14:32:02, 14:34, 14:49, 14:52-14:53, 14:69, 14:74 | *01:14, 04:62, 08:05, 08:18, 08:24-08:25, 08:31, 08:40-08:41, 11:67, 12:01:01-12:21, 12:23-12:26, 13:17, 13:43, 13:45 |
| 127 | 225 bp | 515 bp | *14:01:01-14:01:03, 14:05:01-14:05:03, 14:07:01-14:08, 14:14, 14:18, 14:23:01-14:23:03, 14:26, 14:33, 14:35-14:36, 14:38-14:39, 14:42-14:45, 14:51, 14:54, 14:56, 14:58-14:60, 14:62, 14:64, 14:70, 14:72, 14:75, 14:82, 14:86, 14:88, 14:90-14:92N, 14:95-14:97, 14:99 | *03:11:01-03:11:02, 03:17, 03:24, 03:27, 03:35, 11:17, 11:52-11:54:02, 11:89, 13:76 |
| 135 | 115 bp | 430 bp | *14:01:01-14:01:03, 14:04, 14:07:01-14:08, 14:10, 14:16, 14:22, 14:26, 14:28, 14:31-14:32:02, 14:34-14:35, 14:38-14:39, 14:49-14:50, 14:54-14:55, 14:60-14:62, 14:68, 14:70-14:75, 14:79, 14:82, 14:86-14:88, 14:90, 14:92N-14:93, 14:97, 14:99 | *03:10, 11:52, 13:43 |
| 145,12 | 60 bp, 130 bp | 515 bp | *14:22, 14:25, 14:27, 14:53, 14:73 | *08:08, 11:09, 11:28:01-11:28:02, 11:58:01-11:58:02, 11:83, 11:87, 13:05:01-13:05:02, 13:18, 13:26, 13:42, 13:50:01-13:50:02, 16:08 |
| 159 | 185 bp | 430 bp | *14:28-14:29, 14:99 | *03:20, 12:05, 12:14-12:15 |
| 165 | 105 bp | 430 bp | *14:05:01-14:05:02, 14:08, 14:11, 14:14, 14:18, 14:23:01-14:23:03, 14:34, 14:36, 14:42-14:44:01, 14:45, 14:56, 14:59, 14:64, 14:72, 14:91, 14:95-14:96, 14:100 | *01:02:01-01:02:05, 01:23, 01:26, 01:34, 11:01:05, 11:01:10, 11:04:03, 11:06:02, 11:13:01-11:13:02, 11:17, 11:52, 11:54:01, 12:04, 12:09, 13:01:03, 13:76 |
| 17 | 220 bp | 430 bp | *14:19, 14:21, 14:43 | *11:73, 11:79, 11:93, 13:03:01-13:03:06, 13:10, 13:33:01-13:33:03, 13:37, 13:66:01-13:66:02, 13:81, 13:85, 13:88-13:90, 13:94-13:95, 13:101 |
| 185,7 | 85 bp | 430 bp | *14:01:01-14:01:03, 14:05:01-14:05:03, 14:07:01-14:09, 14:14, 14:16-14:17, 14:21-14:23:03, 14:25-14:26, 14:30, 14:32:01-14:39, 14:42-14:45, 14:53-14:56, 14:58-14:60, 14:62, 14:64-14:65, 14:69, 14:72, 14:74-14:75, 14:80, 14:82, 14:84, 14:86-14:88, 14:90-14:92N, 14:95-14:97, 14:99-14:100 | *03:01:01:01-03:01:09, 03:04-03:26, 03:28, 03:30-03:37, 03:39-03:41, 03:43-03:52, 03:54-03:58, 08:20, 11:01:01-11:04:06, 11:06:01-11:21, 11:23-11:29, 11:31-11:66, 11:68-11:70, 11:73-11:91, 11:93-11:98, 13:01:01-13:14:03, 13:16, 13:18, 13:20-13:25, 13:27-13:52, 13:54-13:56, 13:58-13:66:02, 13:68-13:84, 13:87-13:100, 13:102, DRB3*01:14 |
| 195 | 145 bp | 430 bp | *14:15-14:16, 14:40, 14:55, 14:57, 14:77, 14:84 | *08:09, 08:21, 08:35, 13:08, 13:19, 13:57, 13:64, 13:72, 13:76, 13:83 |
| 205,13 | 110 bp, 160 bp, 215 bp | 430 bp | *14:04, 14:15, 14:25, 14:28, 14:31, 14:50, 14:53, 14:58, 14:68-14:69, 14:71, 14:73, 14:76, 14:79, 14:93 | *08:01:01-08:04:05, 08:06-08:17, 08:19-08:23, 08:26-08:28, 08:30:01-08:30:02, 08:32-08:39, 11:23, 11:25, 11:45, 11:55, 11:64, 11:67, 13:13, 13:45, 13:47, 13:55 |
| 2114 | 170 bp, 215 bp | 515 bp | *14:13, 14:24, 14:37, 14:63, 14:65, 14:78, 14:85 | *03:12, 11:69, 11:82, 13:03:01-13:04, 13:09, 13:12-13:13, 13:21:01-13:21:02, 13:30, 13:32-13:33:03, 13:38, 13:48-13:49, 13:55, 13:58, 13:65-13:66:02, 13:75, 13:81, 13:88-13:90, 13:93-13:95, 13:101 |
| 225 | 145 bp | 430 bp | *14:01:01-14:03:02, 14:05:01-14:09, 14:12:01-14:14, 14:16, 14:18-14:20, 14:22-14:27, 14:29, 14:32:01-14:32:02, 14:34, 14:36-14:41, 14:43-14:45, 14:47-14:49, 14:51, 14:54-14:56, 14:58-14:60, 14:62-14:63, 14:67, 14:69-14:70, 14:74-14:75, 14:77-14:78, 14:80-14:92N, 14:94, 14:96-14:100 | *03:02:01-03:03, 03:06, 03:17, 03:26-03:27, 03:31, 03:35, 03:38, 03:41, 08:20, 11:17, 11:37, 11:41, 11:48, 11:50, 11:52, 11:56, 11:83, 13:03:01-13:03:06, 13:07:01-13:08, 13:12-13:13, 13:19, 13:26, 13:32-13:33:03, 13:36-13:38, 13:40, 13:47-13:49, 13:53, 13:55, 13:58, 13:60, 13:65, 13:70, 13:72, 13:76, 13:81, 13:84-13:85, 13:88-13:90, 13:94-13:95, 13:101, DRB3*01:14 |
| 235 | 145 bp | 430 bp | *14:17, 14:21, 14:30, 14:33, 14:35, 14:42, 14:53, 14:64-14:65, 14:72, 14:95 | *03:01:01:01-03:01:09, 03:04-03:05:03, 03:07-03:15, 03:18-03:25, 03:28-03:30, 03:32-03:34, 03:36-03:37, 03:39-03:40, 03:43-03:52, 03:54-03:58, 11:01:01-11:01:06, 11:01:08-11:04:06, 11:06:01-11:16, 11:18-11:21, 11:23-11:29, 11:31-11:36, 11:38-11:40, 11:42-11:47, 11:49, 11:51, 11:53-11:55, 11:57-11:66, 11:68-11:70, 11:72-11:82, 11:84-11:87, 11:89-11:98, 13:01:01-13:02:04, 13:04-13:06, 13:09-13:11:02, 13:14:01-13:16, 13:18, 13:20-13:25, 13:27-13:31, 13:34, 13:39, 13:41-13:46, 13:50:01-13:52, 13:54, 13:56-13:57, 13:59, 13:61-13:64, 13:66:01-13:66:02, 13:68-13:69, 13:71, 13:73-13:75, 13:77-13:79, 13:82-13:83, 13:86-13:87, 13:91-13:93, 13:96-13:100, 13:102 |
| 245,15 | 75 bp, 115 bp, 135 bp, 165 bp | 430 bp | *14:26, 14:36, 14:41, 14:77, 14:90 | *03:18, 03:43, 11:76, 13:34, 13:64, 13:92, DRB3*01:14 |
| 255,16 | 120 bp, 210 bp, 250 bp | 430 bp | *14:07:01-14:07:02, 14:22, 14:25, 14:44:01-14:44:02, 14:49, 14:53, 14:56, 14:68-14:69, 14:93 | *01:14, 03:35, 04:62, 04:69, 08:08, 08:15, 13:45, DRB3*02:16 |
| 26 | 250 bp | 515 bp | *14:01:01-14:01:03 |  |
| 276,8 | 220 bp | 515 bp | *14:46 |  |
| 288,17 | 160 bp, 230 bp | 430 bp | *14:38, 14:47, 14:50, 14:88, 14:98 | *03:01:02, 03:02:02, 03:05:03, 03:11:02, 03:13:02, 03:17, 03:24, 03:27, 03:35, 11:27:01, 11:84, 12:22, 13:33:02-13:33:03, 13:61, DRB3*01:14 |
| 295,8,18 | 125 bp, 175 bp | 515 bp | *14:48, 14:60, 14:97 | *03:26, 03:33, 03:38, 03:44, 11:44, 11:49-11:50, 11:85, 13:31, 13:46, 13:54, 13:77, 13:100, DRB3*01:14 |
| 30 | 200 bp | 430 bp | *14:06:01-14:06:02, 14:12:01-14:12:02, 14:17-14:18, 14:21, 14:29, 14:33, 14:52, 14:59, 14:64, 14:70, 14:78, 14:80-14:81, 14:83, 14:95 | *03:01:01:01-03:01:09, 03:03, 03:06-03:08, 03:10-03:13:02, 03:15-03:16, 03:18-03:20, 03:22-03:24, 03:26, 03:28, 03:32-03:34, 03:36-03:37, 03:39, 03:42, 03:45-03:48, 03:50-03:51, 03:54-03:56, 03:58, 11:16, 11:40, 11:58:01, 11:59, 11:83, 13:01:01-13:01:08, 13:06, 13:09-13:10, 13:15, 13:18, 13:20, 13:27-13:28, 13:32, 13:35, 13:40, 13:42-13:43, 13:51, 13:53, 13:59, 13:61, 13:69, 13:71, 13:77-13:80, 13:87, 13:91-13:93, 13:98, 13:102, 15:03:01:01-15:03:01:02, 15:23, 15:32 |
| 315 | 80 bp | 515 bp | *14:35, 14:53, 14:72 | *03:10, 11:31, 11:96, 13:43, 13:45 |
| 325,8,19 | 100 bp, 145 bp | 430 bp | *14:13, 14:19, 14:21, 14:63, 14:78, 14:85 | *03:12, 13:10, 13:32, 13:65, 13:85, 13:93 |
| 33 | 175 bp | 430 bp | *14:01:01-14:01:03, 14:07:01-14:08, 14:16, 14:22, 14:25-14:26, 14:32:01-14:32:02, 14:34-14:35, 14:38-14:39, 14:49, 14:53-14:55, 14:58, 14:60, 14:62, 14:69-14:70, 14:74-14:75, 14:82, 14:86-14:88, 14:90, 14:92N, 14:97, 14:99 | *03:10, 13:43, 13:45 |
| 34 | 205 bp | 515 bp | *14:04, 14:11, 14:28, 14:31, 14:50, 14:68, 14:71, 14:93 |  |
| 35 | 250 bp | 430 bp | *14:04, 14:11, 14:15, 14:28, 14:31, 14:50, 14:52, 14:71, 14:73, 14:76, 14:79 | *08:04:01, 08:04:04-08:04:05, 08:06, 08:10, 08:12, 08:22, 08:28, 08:31, 11:67, 12:01:01-12:15, 12:17-12:21, 12:23-12:26, 13:17 |
| 365 | 130 bp | 515 bp | *14:63, 14:67, 14:78 | *11:16, 11:20, 13:01:01-13:01:06, 13:01:08-13:02:04, 13:06, 13:10, 13:15-13:16, 13:27-13:28, 13:31-13:32, 13:34-13:36, 13:39-13:41, 13:51, 13:53, 13:59, 13:61, 13:65, 13:67, 13:69, 13:73, 13:77, 13:79-13:80, 13:85, 13:87, 13:91-13:93, 13:96, 13:98-13:99, 13:102-13:103 |
| 37 | 175 bp | 430 bp |  | *03:08, 11:01:01-11:04:06, 11:06:01-11:21, 11:23-11:29, 11:31-11:66, 11:68-11:70, 11:72-11:98 |
| 38 | 215 bp | 515 bp | *14:37, 14:100 | *13:72 |
| 39 | 155 bp | 430 bp | *14:18, 14:33, 14:51, 14:59, 14:64, 14:95 | *03:11:01 |
| 40 | 250 bp | 430 bp | *14:71 |  |
| 415,20 | 115 bp, 165 bp, 245 bp | 430 bp | *14:75, 14:90-14:91, 14:94, 14:97 | *03:04, 03:26, 03:33, 03:44, 11:29, 11:44, 11:47, 11:49-11:50, 11:63, 11:85, 13:60, 13:84 |
| 42 | 225 bp | 515 bp | *14:76, 14:79, 14:93 |  |
| 43 | 200 bp | 430 bp | *14:82, 14:95 |  |
| 44 | 220 bp | 515 bp | *14:83 |  |
| 455 | 125 bp | 515 bp | *14:59, 14:86, 14:96 |  |
| 465,21 | 105 bp, 165 bp, 220 bp | 515 bp | *14:87-14:89, 14:92N |  |
| 475,7 | 60 bp | 430 bp | *14:63, 14:67, 14:78 | *04:12, 04:18, 04:86, 08:03:02, 08:10, 08:12, 08:14-08:15, 08:19, 08:23, 08:27, 08:30:01-08:30:02, 08:32-08:38, 11:45, 13:13, 15:21 |
| 485,22 | 65 bp, 190 bp, 225 bp | 430 bp | *14:45, 14:98 | *03:35, 11:27:01, 11:84, 12:22, 13:33:02-13:33:03, 13:61 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*14 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*14 subtyping.

In addition, wells number 6, 12, 14, 21, 26, 27, 29, 31, 34, 36, 38, 42 and 44 to 46 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*01:01 consensus sequence.

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, non-DRB1*13 alleles will be amplified by primer mixes 1 to 25, 28 to 33, 35 to 39, 41, 47 and 48. In addition, the DRB3*01:14 allele will be amplified by primer mixes 5, 18, 22, 24, 28 and 29, and the DRB3*02:16 allele will be amplified by primer mix 25.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 3, 10 and 27 have a tendency of giving rise to nonspecific amplifications.

7Primer mixes 4, 12, 18 and 47 may yield less specific PCR product than the other DRB1*14 primer mixes.

8Primer mixes 2, 6, 27 to 29 and 32 have tendencies of giving rise to primer oligomer formation.

9Primer mix 15 has a tendency of giving rise to a weaker positive control band than the other DRB1*14 primer mixes.

10Primer mix 9: Specific PCR fragment of 130 bp in the DRB1* 14:62, 14:65 and 14:78 and the DRB1*03:12, 04:10-04:12, 04:67, 04:91, 08:06, 08:10, 08:12, 08:22, 13:04, 13:32, 13:48, 13:58, 13:75, 13:81, 13:89, 13:93-13:94 and 15:12 alleles. Specific PCR fragment of 250 bp in the DRB1*14:05:01-14:05:03, 14:37, 14:43, 14:45, 14:56, 14:84, 14:91, 14:96 and 14:100 and the DRB1*13:72 alleles.

11Primer mix 11: Specific PCR fragment of 85 bp in the DRB1*14:16, 14:22, 14:25, 14:32:01-14:32:02, 14:34, 14:49, 14:53, 14:69 and 14:74 and the DRB1*01:14, 04:62, 13:43 and 13:45 alleles. Specific PCR fragment of 170 bp in the DRB1*14:11 and DRB1*11:67 alleles. Specific PCR fragment of 215 bp in the DRB1*14:52 and the DRB1* 08:05, 08:18, 08:24-08:25, 08:40, 12:01:01-12:03, 12:05-12:21, 12:23-12:26 and 13:17 alleles. Specific PCR fragment of 85 and 215 bp in the DRB1*14:31 allele. Specific PCR fragment of 170 and 215 bp in the DRB1*08:31, 08:4 and 12:04 alleles.

12Primer mix 14: Specific PCR fragment of 60 bp in the DRB1*14:22, 14:25, 14:53 and 14:73 and the DRB1*08:08 alleles. Specific PCR fragment of 130 bp in the DRB1*14:27 and the DRB1*11:09, 11:28:01-11:28:02, 11:58:01-11:58:02, 11:83, 11:87, 13:05:01-13:05:02, 13:18, 13:26, 13:42, 13:50:01-13:50:02 and 16:08 alleles.

13Primer mix 20: Specific PCR fragment of 110 bp in the DRB1*14:25, 14:53, 14:58 and 14:69 and DRB1*13:45 alleles. Specific PCR fragment of 160 bp in the DRB1*14:04, 14:28, 14:31, 14:50, 14:68, 14:71, 14:73, 14:76, 14:79, 14:93, 08:20, 11:23, 11:25, 11:45, 11:55, 11:64, 13:13, 13:47, 13:55 alleles. Specific PCR fragment of 215 bp in the DRB1*14:15 and the DRB1*08:03:02, 08:09, 08:21, 08:27, 08:32, 08:35 and 08:38 alleles. Specific PCR fragments of 160 and 215 bp in the DRB1*08:01:01-08:02:03, 08:04:01-08:04:05, 08:06, 08:07, 08:10, 08:12-08:17, 08:19, 08:22, 08:23, 08:26, 08:28, 08:30:01-08:30:02, 08:33, 08:34, 08:36, 08:37, 08:39 and 11:67 alleles. Specific PCR fragments of 110, 160 and 215 bp in the DRB1*08:08 and 08:11 alleles. Not all specific PCR fragments may be obtained in alleles being amplified by more than one primer pair.

14Primer mix 21: Specific PCR fragment of 170 bp in DRB1*14:13, 14:63, 14:65, 14:78 and 14:85 and the DRB1*03:12, 13:03:01-13:04, 13:12-13:13, 13:21:01-13:21:02, 13:30, 13:32-13:33:03, 13:38, 13:48-13:49, 13:55, 13:58, 13:65-13:66:02, 13:75, 13:81, 13:88-13:90, 13:93-13:95, 13:101 alleles. Specific PCR fragment of 215 bp in the DRB1*14:24 and  14:37 and the DRB1*11:69, 11:82 and 13:09 alleles.

15Primer mix 24: Specific PCR fragment of 75 bp in the DRB1*14:26 and the DRB1*13:92 alleles. Specific PCR fragment of 115 bp in the DRB1*14:36, 14:41 and 14:77 and the DRB1*03:18, 03:43, 11:76, 13:34 and 13:64 and in the DRB3*01:14 alleles. Specific PCR fragment of 135 bp in the DRB1*1436 and DRB1*0318 alleles. Specific PCR fragment of 165 bp in the DRB1*14:90 allele.

16Primer mix 25: Specific PCR fragment of 120 bp in the DRB1*14:07:01-14:07:02, 14:22, 14:25, 14:49, 14:53, 14:68-14:69 and 14:93 and the DRB1*01:14, 04:62, 04:69, 08:08, 08:15 and 13:45 and in the  DRB3*02:16 alleles. Specific PCR fragment of 210 bp in the DRB1*14:56 allele. Specific PCR fragment of 250 bp in the DRB1*14:44:01-14:44:02 and the DRB1*03:35 alleles.

17Primer mix 28: Specific PCR fragment of 160 bp in the DRB1*14:88 allele. Specific PCR fragment of 230 bp in the DRB1*14:38, 14:47, 14:50 and 14:98 and the DRB1*03:01:02, 03:02:02, 03:05:03, 03:11:02, 03:13:02, 03:17, 03:24, 03:27, 03:35, 11:27:01, 11:84, 12:22, 13:33:02-13:33:03 and 13:61 and in the  DRB3*01:14 alleles.

18Primer mix 29: Specific PCR fragment of 125 bp in the DRB1*14:60 and 14:97 and the DRB1*03:26, 03:33, 03:44, 11:44, 11:49-11:50 and 11:85 alleles. Specific PCR fragment of175 bp in the DRB1*14:48 and the DRB1*03:38, 13:31, 13:46, 13:54, 13:77 and13:100 and the DRB3*01:14 alleles.

19Primer mix 32: Specific PCR fragment of 100 bp in the DRB1*14:13, 14:63, 14:78 and 14:85 and the DRB1*03:12, 13:32, 13:65 and 13:93 alleles. Specific PCR fragment of 145 bp in the DRB1*14:19 and 14:21 and the DRB1*03:12, 13:10 and 13:85 alleles.

20Primer mix 41: Specific PCR fragment of 115 bp in the DRB1*14:91, 14:94 and 14:97 and in the DRB1*03:04, 03:26, 03:33, 03:44, 11:29, 11:44, 11:47, 11:49-11:50, 11:63, 11:85, 13:60 and 13:84 alleles. Specific PCR fragment of 165 bp in the DRB1*14:90 allele. Specific PCR fragment of 245 bp in the DRB1*14:75 allele.

21Primer mix 46: Specific PCR fragment of 105 bp in the DRB1*14:92N allele. Specific PCR fragment of 165 bp in the DRB1*14:88 allele. Specific PCR fragment of 220 bp in the DRB1*14:87 and 14:89 alleles.

22Primer mix 48: Specific PCR fragment of 65 bp in the DRB1*14:98 and the DRB1*11:27:01, 11:84, 12:22, 13:33:02-13:33:03 and 13:61 alleles. Specific PCR fragment of 190 bp in the DRB1*14:45 allele. Specific PCR fragment of 225 bp in the DRB1*03:35 allele.

‘w’, might be weakly amplified.

1The DRB1*1466 allele has been renamed to DRB1*14:32:02

2Primer mix 9: Specific PCR fragment of 130 bp in the DRB1* 14:62, 14:65 and 14:78 and the DRB1*03:12, 04:10-04:12, 04:67, 04:91, 08:06, 08:10, 08:12, 08:22, 13:04, 13:32, 13:48, 13:58, 13:75, 13:81, 13:89, 13:93-13:94 and 15:12 alleles. Specific PCR fragment of 250 bp in the DRB1*14:05:01-14:05:03, 14:37, 14:43, 14:45, 14:56, 14:84, 14:91, 14:96 and 14:100 and the DRB1*13:72 alleles.

Primer mix 11: Specific PCR fragment of 85 bp in the DRB1*14:16, 14:22, 14:25, 14:32:01-14:32:02, 14:34, 14:49, 14:53, 14:69 and 14:74 and the DRB1*01:14, 04:62, 13:43 and 13:45 alleles. Specific PCR fragment of 170 bp in the DRB1*14:11 and DRB1*11:67 alleles. Specific PCR fragment of 215 bp in the DRB1*14:52 and the DRB1* 08:05, 08:18, 08:24-08:25, 08:40, 12:01:01-12:03, 12:05-12:21, 12:23-12:26 and 13:17 alleles. Specific PCR fragment of 85 and 215 bp in the DRB1*14:31 allele. Specific PCR fragment of 170 and 215 bp in the DRB1*08:31, 08:4 and 12:04 alleles.

Primer mix 14: Specific PCR fragment of 60 bp in the DRB1*14:22, 14:25, 14:53 and 14:73 and the DRB1*08:08 alleles. Specific PCR fragment of 130 bp in the DRB1*14:27 and the DRB1*11:09, 11:28:01-11:28:02, 11:58:01-11:58:02, 11:83, 11:87, 13:05:01-13:05:02, 13:18, 13:26, 13:42, 13:50:01-13:50:02 and 16:08 alleles.

Primer mix 20: Specific PCR fragment of 110 bp in the DRB1*14:25, 14:53, 14:58 and 14:69 and DRB1*13:45 alleles. Specific PCR fragment of 160 bp in the DRB1*14:04, 14:28, 14:31, 14:50, 14:68, 14:71, 14:73, 14:76, 14:79, 14:93, 08:20, 11:23, 11:25, 11:45, 11:55, 11:64, 13:13, 13:47, 13:55 alleles. Specific PCR fragment of 215 bp in the DRB1*14:15 and the DRB1*08:03:02, 08:09, 08:21, 08:27, 08:32, 08:35 and 08:38 alleles. Specific PCR fragments of 160 and 215 bp in the DRB1*08:01:01-08:02:03, 08:04:01-08:04:05, 08:06, 08:07, 08:10, 08:12-08:17, 08:19, 08:22, 08:23, 08:26, 08:28, 08:30:01-08:30:02, 08:33, 08:34, 08:36, 08:37, 08:39 and 11:67 alleles. Specific PCR fragments of 110, 160 and 215 bp in the DRB1*08:08 and 08:11 alleles. Not all specific PCR fragments may be obtained in alleles being amplified by more than one primer pair.

Primer mix 21: Specific PCR fragment of 170 bp in DRB1*14:13, 14:63, 14:65, 14:78 and 14:85 and the DRB1*03:12, 13:03:01-13:04, 13:12-13:13, 13:21:01-13:21:02, 13:30, 13:32-13:33:03, 13:38, 13:48-13:49, 13:55, 13:58, 13:65-13:66:02, 13:75, 13:81, 13:88-13:90, 13:93-13:95, 13:101 alleles. Specific PCR fragment of 215 bp in the DRB1*14:24 and  14:37 and the DRB1*11:69, 11:82 and 13:09 alleles.

Primer mix 24: Specific PCR fragment of 75 bp in the DRB1*14:26 and the DRB1*13:92 alleles. Specific PCR fragment of 115 bp in the DRB1*14:36, 14:41 and 14:77 and the DRB1*03:18, 03:43, 11:76, 13:34 and 13:64 and in the DRB3*01:14 alleles. Specific PCR fragment of 135 bp in the DRB1*1436 and DRB1*0318 alleles. Specific PCR fragment of 165 bp in the DRB1*14:90 allele.

Primer mix 25: Specific PCR fragment of 120 bp in the DRB1*14:07:01-14:07:02, 14:22, 14:25, 14:49, 14:53, 14:68-14:69 and 14:93 and the DRB1*01:14, 04:62, 04:69, 08:08, 08:15 and 13:45 and in the  DRB3*02:16 alleles. Specific PCR fragment of 210 bp in the DRB1*14:56 allele. Specific PCR fragment of 250 bp in the DRB1*14:44:01-14:44:02 and the DRB1*03:35 alleles.

Primer mix 28: Specific PCR fragment of 160 bp in the DRB1*14:88 allele. Specific PCR fragment of 230 bp in the DRB1*14:38, 14:47, 14:50 and 14:98 and the DRB1*03:01:02, 03:02:02, 03:05:03, 03:11:02, 03:13:02, 03:17, 03:24, 03:27, 03:35, 11:27:01, 11:84, 12:22, 13:33:02-13:33:03 and 13:61 and in the  DRB3*01:14 alleles.

Primer mix 29: Specific PCR fragment of 125 bp in the DRB1*14:60 and 14:97 and the DRB1*03:26, 03:33, 03:44, 11:44, 11:49-11:50 and 11:85 alleles. Specific PCR fragment of175 bp in the DRB1*14:48 and the DRB1*03:38, 13:31, 13:46, 13:54, 13:77 and13:100 and the DRB3*01:14 alleles.

Primer mix 32: Specific PCR fragment of 100 bp in the DRB1*14:13, 14:63, 14:78 and 14:85 and the DRB1*03:12, 13:32, 13:65 and 13:93 alleles. Specific PCR fragment of 145 bp in the DRB1*14:19 and 14:21 and the DRB1*03:12, 13:10 and 13:85 alleles.

Primer mix 41: Specific PCR fragment of 115 bp in the DRB1*14:91, 14:94 and 14:97 and in the DRB1*03:04, 03:26, 03:33, 03:44, 11:29, 11:44, 11:47, 11:49-11:50, 11:63, 11:85, 13:60 and 13:84 alleles. Specific PCR fragment of 165 bp in the DRB1*14:90 allele. Specific PCR fragment of 245 bp in the DRB1*14:75 allele.

Primer mix 46: Specific PCR fragment of 105 bp in the DRB1*14:92N allele. Specific PCR fragment of 165 bp in the DRB1*14:88 allele. Specific PCR fragment of 220 bp in the DRB1*14:87 and 14:89 alleles.

Primer mix 48: Specific PCR fragment of 65 bp in the DRB1*14:98 and the DRB1*11:27:01, 11:84, 12:22, 13:33:02-13:33:03 and 13:61 alleles. Specific PCR fragment of 190 bp in the DRB1*14:45 allele. Specific PCR fragment of 225 bp in the DRB1*03:35 allele.

‘w’, might be weakly amplified.