Gel Documentation Form and Worksheet

DRB1*14 (101.117-12/03, -12u/03u) 	Lot No: 06X	Expiry Date: 2017-January-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                           Failed lanes: ___________ 	                               Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 11, 23 and 28 may have tendencies of unspecific amplifications.

Primer mixes 13, 16, 18 and 47 may give rise to a lower yield of HLA-specific PCR product than the other DRB1*14 primer mixes.

Primer mixes 2, 16, 26, 40, 43, 45 and 46 have a tendency of giving rise to primer oligomer formation.

Primer mix 48 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

1DRB1*14 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.16.0, April 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 7: Specific PCR fragment of 200 bp in the DRB1*14:45 and 14:122 alleles. Specific PCR fragment of 225 bp in the DRB1*14:03:01-14:03:02, 14:12:01-14:12:02, 14:27, 14:40, 14:55, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89, 14:102, 14:115-14:116, 14:135-14:136 and 14:144 and the DRB1*08:20, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:119, 13:13, 13:18, 13:47, 13:55, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158 and 13:164 alleles.

Primer mix 9: Specific PCR fragment of 130 bp in the DRB1*14:62, 14:65 and 14:78 and in the DRB1*03:12, 04:10:01-04:12, 04:67, 04:91, 04:126, 04:136, 04:147, 08:06, 08:10, 08:12, 08:22, 08:54, 12:46, 13:04, 13:32, 13:48, 13:58, 13:75, 13:81, 13:89, 13:93-13:94:02, 13:108, 13:164, 13:169 and 15:12 alleles. Specific PCR fragment of 250 bp in the DRB1*14:05:01-14:05:04, 14:37, 14:43, 14:45, 14:56, 14:84, 14:91, 14:96, 14:100, 14:103, 14:127:01-14:127:02, 14:130, 14:136 and the DRB1*13:72 alleles.

Primer mix 14: Specific PCR fragment of 60 bp in the DRB1*14:22, 14:25, 14:53, 14:73, 14:105 and 14:128 and the DRB1*08:08 alleles. Specific PCR fragment of 130 bp in the DRB1*14:27 and the DRB1*11:09, 11:28:02, 11:58:01, 11:83, 11:87, 11:113, 11:115, 13:05:01-13:05:02, 13:18, 13:26:01-13:26:02, 13:42, 13:50:02, 13:136, 13:144, 13:158 and 16:08 alleles.

Primer mix 20: Specific PCR fragment of 110 bp in the DRB1*14:25, 14:53, 14:58, 14:69 and 14:128 and the DRB1*13:45 alleles. Specific PCR fragment of 160 bp in the DRB1*14:04, 14:28, 14:31, 14:50, 14:68, 14:71, 14:73, 14:76, 14:79, 14:93, 14:102, 14:107, 14:116, 14:120, 14:138 and 14:145  and the DRB1*08:20, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:119, 13:13, 13:47, 13:55, 13:154 and 13:164 alleles.

Specific PCR fragment of 215 bp in the DRB1*14:15 and the DRB1*08:01:02, 08:09, 08:21, 08:27, 08:32, 08:35, 08:36:02 and 08:38 alleles. Specific PCR fragments of 160 and 215 bp in the DRB1*08:01:01, 08:01:03-08:04:07, 08:06-08:07, 08:10-08:17, 08:19, 08:22-08:23, 08:26, 08:28, 08:30:01-08:30:03, 08:33-08:34, 08:36:01, 08:37, 08:39, 08:42, 08:44-08:46, 08:48-08:55 and 11:67 alleles. Specific PCR fragments of 110, 160 and 215 bp in the DRB1*08:08 allele.  Not all specific PCR fragments may be obtained in alleles being amplified by more than one primer pair.

Primer mix 21: Specific PCR fragment of 170 bp in DRB1*14:13, 14:63, 14:65, 14:78 and 14:85 and the DRB1*03:12, 13:03:01-13:04, 13:12:01-13:13, 13:21:01-13:21:02, 13:30, 13:32-13:33:03, 13:38, 13:48-13:49, 13:55, 13:58, 13:65-13:66:02, 13:75, 13:81, 13:88-13:90, 13:93-13:95, 13:101, 13:108, 13:115, 13:118, 13:120, 13:122, 13:133-13:134, 13:139, 13:151-13:152, 13:154, 13:161, 13:164, 13:167, 13:169 and 13:174 alleles. Specific PCR fragment of 215 bp in the DRB1*14:24 and 14:37 and the DRB1*11:69, 11:82, 13:09 and 13:129 alleles.

Primer mix 24: Specific PCR fragment of 75 bp in the DRB1*14:26 and the DRB1*13:92 alleles. Specific PCR fragment of 115 bp in the DRB1*14:41, 14:77 and 14:110 and the DRB1*03:43, 03:71, 11:76, 11:114, 13:34, 13:64, 13:136 and 13:174 and in the DRB3*01:14 alleles. Specific PCR fragment of 135 bp in the DRB1*14:36 and the DRB1*03:18 alleles. Specific PCR fragment of 165 bp in the DRB1*14:90 allele.

Primer mix 25: Specific PCR fragment of 85 bp in the DRB1*14:55 allele. Specific PCR fragment of 120 bp in the DRB1*14:07:01-14:07:02, 14:22, 14:25, 14:49, 14:53, 14:68-14:69, 14:93, 14:119 and 14:143 and the DRB1*01:14, 04:62, 04:69, 13:45, 13:159, 13:179, DRB3*02:16 alleles. Specific PCR fragment of 210 bp in the DRB1*14:56, 14:84 and 14:136 allele. Specific PCR fragment of 250 bp in the DRB1*14:44:01-14:44:03 and the DRB1*03:35 alleles. Specific PCR fragment of 85 bp and 120 bp in the DRB1*08:08 and 08:15 alleles.

Primer mix 26: Specific PCR fragment of 125 bp in the DRB1*14:59, 14:86, 14:96 and 14:145 and the DRB1*12:34 alleles. Specific PCR fragment of 250 bp in the DRB1*14:01:01-14:01:04 alleles.

Primer mix 27: Specific PCR fragment of 165 bp in the DRB1*14:141 and in the DRB3*01:07, DRB3*02:01-02:08, 02:11, 02:13-02:29N alleles. Specific PCR fragment of 220 bp in the DRB1*14:46 alleles.

Primer mix 28: Specific PCR fragment of 170 bp in the DRB1*14:11 and 14:148 and the DRB1*08:31, 08:41, 11:67 and 12:04 alleles. Specific PCR fragment of 225 bp in the DRB1*14:50, 14:76, 14:79 and 14:93 and the DRB1*03:92 and 12:22 alleles.

Primer mix 29: Specific PCR fragment of 125 bp in the DRB1*14:60 and 14:97 and the DRB1*03:26, 03:33, 03:44, 11:44, 11:49:01-11:50, 11:85, 11:127 and 13:106 alleles. Specific PCR fragment of 175 bp in the DRB1*14:48 and the DRB1*03:38, 03:79, 03:85, 03:96, 13:31, 13:46, 13:54, 13:77, 13:100 and 13:162 and in the DRB3*01:14 alleles. Specific PCR fragment of 220 bp in the DRB1*14:104, 14:111 and 14:137N and the DRB1*03:14-03:15, 03:90 and 11:105 alleles.

Primer mix 32: Specific PCR fragment of 100 bp in the DRB1*14:13, 14:63, 14:78 and 14:85 and in the DRB1*03:12, 13:32, 13:65, 13:93 and 13:139 alleles. Specific PCR fragment of 145 bp in the DRB1*14:19, 14:21 and 14:109 and in the DRB1*03:76, 13:10, 13:85 and 13:170, 13:176 and 13:178 alleles. Specific PCR fragment of 100 and 145 bp in the DRB1*13:120 allele.

Primer mix 35: Specific PCR fragment of 250 bp in the DRB1*14:04, 14:11, 14:15, 14:28, 14:31, 14:50, 14:52, 14:71, 14:73, 14:76, 14:79, 14:107, 14:120, 14:126, 14:138, 14:145 and 14:148 and the DRB1*03:92, 08:04:01, 08:04:04-08:04:07, 08:06, 08:10, 08:12, 08:22, 08:28, 08:31, 08:54, 11:67, 12:01:01-12:15, 12:17-12:21, 12:23-12:38, 12:40-12:41, 12:43-12:46, 13:17, 13:116 and 13:175 alleles. Specific PCR fragment of 280 bp in the DRB1*14:113 and the DRB1*04:85 alleles.

Primer mix 36: Specific PCR fragment of 120 bp in the DRB1*14:48, 14:67, 14:78, 14:108-14:109, 14:115, 03:13:01-03:13:02, 03:38, 03:58, 03:65, 03:79, 03:88, 11:16, 11:20, 13:01:01-13:01:06, 13:01:08-13:02:09, 13:06, 13:10, 13:15-13:16, 13:27-13:28, 13:31-13:32, 13:34-13:36, 13:39-13:41, 13:51, 13:53, 13:59, 13:61:01-13:61:02, 13:65, 13:67, 13:69, 13:73, 13:77-13:80, 13:85, 13:87, 13:91-13:93, 13:96:01-13:96:02, 13:98-13:99, 13:102-13:105, 13:109-13:110, 13:112-13:113N, 13:117, 13:120-13:121, 13:123-13:128, 13:130-13:131, 13:137N-13:139, 13:141-13:143, 13:145-13:148, 13:153, 13:155-13:157, 13:159-13:160, 13:163, 13:165-13:166, 13:168, 13:170-13:171, 13:173, 13:176-13:177, 13:179, DRB3*01:08 alleles. Specific PCR fragment of 155 bp in the DRB1*14:18, 14:33, 14:51, 14:59, 14:95, 14:106 and 14:132-14:134 and the DRB1*03:11:01 and 03:97 alleles. Specific PCR fragment of 120 bp and 155 bp in the DRB1*14:64 allele.

Primer mix 38: Specific PCR fragment of 165 bp in the DRB1*14:05:01, 14:05:03-14:05:04, 14:43-14:45, 14:56, 14:91, 14:96, 14:103, 14:127:01-14:127:02 and 14:136 and the DRB1*03:35 alleles. Specific PCR fragment of 165 and 215 bp in the DRB1*14:37 and 14:100 and the DRB1*13:72 alleles.

Primer mix 40: Specific PCR fragment of 215 bp in the DRB1*14:107 allele. Specific PCR fragment of 250 bp in the DRB1*14:71 allele.

Primer mix 41: Specific PCR fragment of 115 bp in the DRB1*14:91, 14:94 and 14:97 and the DRB1*03:04:01-03:04:02, 03:26, 03:33, 03:44, 11:29, 11:44, 11:47, 11:49:01-11:50, 11:63, 11:85, 11:127, 13:60, 13:84 and 13:106 alleles. Specific PCR fragment of 140 bp in the DRB1*14:129 allele. Specific PCR fragment of 165 bp in the DRB1*14:90 allele. Specific PCR fragment of 245 bp in the DRB1*14:75 and 14:103 alleles.

Primer mix 43: Specific PCR fragment of 95 bp in the DRB1*14:101 allele. Specific PCR fragment of 200 bp in the DRB1*14:82, 14:95 and 14:132 and the DRB1*03:01:17 alleles.

Primer mix 44: Specific PCR fragment of 215 bp in the DRB1*14:70 and 14:83 and the DRB1*15:03:01:01-15:03:02, 15:23, 15:54, 15:57 and 15:94 alleles. Specific PCR fragment of 280 bp in the DRB1*14:113 and the DRB1*04:85 alleles.

Primer mix 45: Specific PCR fragment of 65 bp in the DRB1*14:98 and the DRB1*11:27:01, 11:84:01, 11:124, 11:138, 12:22, 13:33:02-13:33:03, 13:61:01, 13:94:02 and 13:96:02 alleles. Specific PCR fragment of 190 bp in the DRB1*14:45 allele. Specific PCR fragment of 225 bp in the DRB1*14:127:01 and the DRB1*03:35 alleles.

Primer mix 46: Specific PCR fragment of 105 bp in the DRB1*14:92N allele. Specific PCR fragment of 140 bp in the DRB1*14:129 allele. Specific PCR fragment of 165 bp in the DRB1*14:88 allele. Specific PCR fragment of 220 bp in the DRB1*14:87 and 14:89 alleles.

4The DRB1*14 primer set cannot separate the DRB1*14:21 and the DRB1*03:76 and DRB1*13:178 alleles. These alleles can be distinguished by the DR low resolution and/or DRB1*03 and DRb1*13 kits.

5The following DRB1*14 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| DRB1*14:01:02-14:01:04, 14:86 | 26 |
| DRB1*14:26, 14:110 | 24 |

The DRB1*14 subtyping kit cannot distinguish the silent mutations in the DRB1*14:01:01-14:01:04 alleles, the DRB1*14:03:01-14:03:02 alleles, the DRB1*14:06:02-14:06:03 alleles, the DRB1*14:12:01-14:12:02 alleles, the DRB1*14:23:01-14:23:04 alleles, the DRB1*14:32:01-14:32:02 alleles, the DRB1*14:44:01 and 14:44:03 alleles or the DRB1*14:54:01-14:54:02 alleles.

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.