Gel Documentation Form and Worksheet

DRB3 (101.121-24/04,-24u/04u) Lot No: 7F6 Expiry Date: 2020-06-01

Name:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Sample ID:\_\_\_\_\_\_\_\_\_\_\_\_\_\_

DNA Conc.(ng/ul):\_\_\_\_\_\_\_\_\_

Test Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Tested By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Review Date:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed By:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

*Interpretation:\_\_\_\_\_\_\_\_\_\_\_ Failed lanes*: \_\_\_\_\_\_\_\_\_\_\_\_ *Comments:\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_*

**Gel Picture**

|  |
| --- |
| PHOTO DOCUMENT |





‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

**Notes:**

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 9, 10 and 13 may have tendencies of unspecific amplifications.

Primer mix 7, 11, 19, 22 and 46 have a tendency giving rise to primer oligomer formation.

Primer mixes 4, 10 and 29 may give rise to a lower yield of HLA-specific PCR product than the other DRB3 primer mixes.

Primer mix 46 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.









**1**DRB3 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page [www.ebi.ac.uk/imgt/hla](http://www.ebi.ac.uk/imgt/hla), release 3.29.0, August 2017.

**2**Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page <http://hla.alleles.org/alleles/deleted.html>.

**3**The following DRB3 primer mixes have two or more product sizes:

|  |  |  |  |
| --- | --- | --- | --- |
| Primer Mix | Size of spec. PCR product | Amplified DRB3alleles | Other amplified DRBalleles |
|  **5** | 95 bp 125 bp | \*01:03\*01:10 | **DRB1\*14:46** |
|  **7** | 90 bp 180 bp | \*01:05\*01:07, 01:15,02:42 | **DRB1\*03:42, DRB1\*14:46** |
| **12** | 90 bp 145 bp | \*02:17\*02:01, 02:04, 02:24, 02:48, 02:58 | **DRB1\*03:37, DRB1\*03:80, DRB1\*03:140, DRB1\*04:183, DRB1\*11:43, DRB1\*11:50, DRB1\*11:115, DRB1\*11:151, DRB1\*11:171, DRB1\*11:174, DRB1\*12:34, DRB1\*13:131, DRB1\*13:150, DRB1\*14:59, DRB1\*14:86, DRB1\*14:96, DRB1\*14:145, DRB1\*14:153-14:154, DRB1\*15:65** |
| **14** | 185 bp265 bp | \*01:07, 01:09, 02:02:01:01-02:02:09, 02:05, 02:07-02:13, 02:15-02:19, 02:21-02:23, 02:25-02:36, 02:39-02:47, 02:49, 02:53-02:57, 02:60, 02:62-02:65\*01:02 | **DRB1\*10:01:01:01w-10:01:09w, DRB1\*10:03w-10:11w, DRB1\*10:13w, DRB1\*10:15w-10:16w, DRB1\*10:18w-10:22w, DRB1\*13:216, DRB1\*14:141, DRB1\*16:30** |
| **15** | 120 bp 180 bp 210 bp | \*02:03, 2:38, 02:50, 03:01:01-03:09, 03:11, 03:13\*02:16\*02:17 | **DRB1\*12:57, DRB1\*13:195** |
| **17** | 120 bp 180 bp | \*01:08, 02:06, 02:20, 03:10\*02:16, 02:23 | **DRB1\*03:42, DRB1\*03:87, DRB1\*13:67, DRB1\*14:46** |
| **18** | 100 bp170 bp | \*01:09, 02:07, 02:09, 02:21, 02:44, 02:56\*01:18 |  |
| **19** | 180 bp 270 bp | \*01:42, 02:08, 02:18, 02:23\*02:14 | **DRB1\*11:30** |
| **21** | 120 bp 155 bp210 bp | \*01:13\*02:43\*01:19, 01:31, 02:11 |  |
| **22** | 185 bp 245 bp  | \*01:11\*01:02, 02:12 |  |
| **23** | 195 bp240 bp | \*02:13\*01:02 |  |
| **25** | 150 bp220 bp 260 bp | \*02:60\*01:19, 02:26\*01:12, 02:31:01 |  |
| **28** | 235 bp260 bp | \*02:27\*01:12, 02:31:01, 02:55N | **DRB1\*11:30, DRB1\*12:57, DRB1\*13:67, DRB1\*13:195, DRB1\*14:46** |
| **31** | 70 bp95 bp | \*01:21, 02:19, 02:21, 02:25, 02:48, 02:56\*01:16, 02:29N | **DRB1\*12:34****DRB1\*01:77w** |
| **33** | 150 bp180 bp | \*01:32, 01:39, 01:41, 02:37-02:38, 02:50, 03:01:01-03:02, 03:05-03:09, 03:11-03:13\*01:20 | **DRB1\*03:35, DRB1\*07:04, DRB1\*14:38:01-14:38:02, DRB1\*14:50, DRB1\*14:93, DRB1\*14:127:01-14:127:02** |
| **34** | 135 bp175 bp | \*01:43, 03:05\*01:24, 02:36 |  |
| **35** | 115 bp250 bp | \*02:51\*01:17, 02:33 |  |
| **36** | 105 bp155 bp185 bp | \*01:22\*02:43\*01:35, 02:35 |  |
| **37** | 155 bp215 bp | \*02:30\*01:31, 01:45, 02:11, 02:47 |  |
| **38** | 70 bp205 bp | \*02:32\*01:07?, 02:05?, 02:07?-02:08?, 02:34 | **DRB1\*03:44, DRB1\*10:13, DRB1\*13:106** |
| **39** | 190 bp260 bp | \*02:39\*01:40N |  |
| **40** | 130 bp190 bp | \*03:13\*01:07, 01:15, 01:30, 01:42, 02:42 | **DRB1\*03:42, DRB1\*14:46** |
| **42** | 165 bp205 bp | \*02:61Q\*01:37, 03:08 |  |
| **44** | 100 bp260 bp | \*01:26N\*02:55N |  |

**4**The following DRB3 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

|  |  |
| --- | --- |
| Alleles | Primer mix |
| DRB3\*02:32, 02:34 | 38 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.