PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*11

Prod. No:	101. 416-12/12u

101.416-04/04u

Lot No:    	62N

Expiry Date:    	2014-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 56 primer mixes used for HLA-A*11 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*11 alleles3 | Other amplified HLA-A alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 195 bp | 800 bp | *11:01:01-11:01:35, 11:03-11:13, 11:15:01-11:15:02, 11:20-11:27, 11:29-11:37, 11:39, 11:41-11:44, 11:46-11:52Q, 11:54-11:56, 11:58-11:70, 11:72-11:76, 11:79-11:97, 11:99N-11:100, 11:102-11:105, 11:107-11:109N, 11:112 | *01:13, 02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:78, 02:103, 03:01:01:01-03:01:08, 03:01:10-03:01:22, 03:01:24-03:07, 03:09-03:31, 03:33-03:35, 03:37-03:40, 03:42-03:56, 03:58-03:88, 03:90-03:106, 03:109-03:110, 03:112-03:133, 03:135, 24:19, 30:18, 30:55, 34:01:01-34:09, 66:01-66:02, 66:04, 66:06-66:10, 66:12-66:14, 66:16, 68:01:01:01-68:02:04, 68:06-68:14, 68:16-68:19, 68:21:01-68:30, 68:32-68:35, 68:37-68:45, 68:47-68:56, 68:58-68:83, 68:85, 69:01, 74:13, C*03:82 |
| 2 | 270 bp | 1070 bp | *11:02:01-11:02:03, 11:14, 11:16, 11:38, 11:57, 11:77, 11:101, 11:110 |  |
| 337 | 190 bp | 1070 bp | *11:03, 11:33:01-11:33:02 | *01:61, 03:22:01-03:22:02, 03:42, 03:135, 30:01:01-30:04:02, 30:06-30:20, 30:22-30:27N, 30:29-30:34, 30:36-30:46, 30:48-30:51, 30:53, 30:55-30:58, 31:03-31:04, 33:49, 34:02:01w-34:02:02w, 34:03, 34:04w, 34:07w-34:09w |
| 4 | 235 bp | 1070 bp | *11:04, 11:27, 11:35 | *01:12, 01:14, 01:19, 02:156, 03:01:01:01-03:01:29, 03:01:31-03:17, 03:19-03:39, 03:41-03:74, 03:76-03:94, 03:96-03:97, 03:99-03:104, 03:106-03:133, 24:92, 30:01:01-30:04:02, 30:06-30:13, 30:15-30:20, 30:22-30:27N, 30:29-30:42, 30:44-30:46, 30:48-30:51, 30:53-30:58, 31:03-31:04, 32:04, 33:49, 34:02:01-34:04, 34:07-34:09, 36:01-36:02, 36:05 |
| 59 | 180 bp, 210 bp | 800 bp | *11:05, 11:110 | *02:294, 66:08 |
| 610 | 195 bp, 550 bp | 1070 bp | *11:06-11:07, 11:18, 11:22 | *01:79w, 01:104w, 02:38, 02:101:01, 02:135, 02:154, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:36 |
| 77 | 215 bp | 1070 bp | *11:08, 11:44, 11:50Q, 11:94 | *01:01:01:01-01:02, 01:04N, 01:07-01:11N, 01:13-01:18N, 01:21-01:24, 01:26-01:33, 01:35-01:40, 01:42-01:62, 01:64, 01:67-01:72, 01:74-01:88, 01:90-01:104, 03:18, 03:50, 03:66, 03:135, 30:04:01-30:04:02, 30:06, 30:17, 30:29, 30:46, 36:01-36:02, 36:04-36:05 |
| 85 | 100 bp | 800 bp | *11:09, 11:34 |  |
| 95,6,8 | 85 bp | 800 bp | *11:27, 11:38-11:39, 11:94 | *01:26, 24:59 |
| 105,11 | 120 bp, 280 bp | 1070 bp | *11:10, 11:43 | *01:51, 25:01:01-25:12N, 25:14-25:16, 26:01:01-26:06, 26:08-26:18, 26:20-26:29, 26:32-26:43:02, 26:45-26:64, 26:66-26:71N, 31:03, 33:13, 34:01:01-34:08, 66:01, 66:04-66:11, 66:13-66:15, B*07:102, B*08:18, B*35:178, B*73:01-73:02 |
| 11 | 215 bp | 800 bp | *11:01:01-11:07, 11:09-11:22, 11:27, 11:29-11:30, 11:32-11:34, 11:36-11:43, 11:45-11:47, 11:49, 11:51-11:52Q, 11:54-11:93, 11:95-11:100, 11:102-11:112 | *01:12, 01:19, 01:25w, 03:02:01-03:02:03, 03:10, 03:31-03:32, 03:65, 03:69N, 03:73, 03:76, 03:82, 03:90, 03:106, 03:113, 24:92 |
| 12 | 235 bp | 1070 bp | *11:01:01-11:01:20, 11:01:22-11:03, 11:05-11:25, 11:29-11:34, 11:36-11:47, 11:49, 11:51-11:52Q, 11:54-11:89, 11:91-11:100, 11:102-11:112 | *01:01:01:01-01:01:22, 01:01:24-01:02, 01:04N, 01:06-01:07, 01:09-01:11N, 01:13, 01:16N-01:18N, 01:21-01:33, 01:35-01:40, 01:42-01:62, 01:64-01:65, 01:67-01:72, 01:74-01:88, 01:90-01:104, 03:18, 03:135, 36:04 |
| 138 | 195 bp | 800 bp | *11:11 | *43:01 |
| 145,7 | 100 bp | 1070 bp | *11:01:01-11:11, 11:13-11:16, 11:20-11:27, 11:29-11:39, 11:41-11:52Q, 11:54-11:60, 11:61w, 11:62-11:95, 11:97, 11:99N-11:105, 11:107-11:112 | *01:13, 01:17, 03:63, 03:88, 25:02, 26:13, 26:19, 26:33, 34:01:01w-34:01:02w, 34:02:01-34:04, 34:05w, 34:06, 34:08, 66:01, 66:04, 66:06-66:11, 66:13-66:14 |
| 155,12 | 100 bp, 215 bp | 800 bp | *11:12, 11:24:02-11:25, 11:31, 11:35 | *01:06, 02:34w-02:35:03w, 02:56:01-02:56:02, 02:62, 02:78, 02:103, 02:156, 03:01:01:01-03:31, 03:33-03:112, 03:113w, 03:114-03:133, 03:135, 24:24, 29:01:01:01-29:03, 29:04w, 29:05-29:18, 29:21-29:29, 29:31-29:32, 30:01:01-30:02:04, 30:03, 30:07-30:20, 30:22-30:27N, 30:30-30:45, 30:48-30:58, 31:03-31:04, 32:04, 32:17, 33:49, 34:02:01, 34:03-34:04, 34:07-34:09, 66:02, 66:12, 66:16, 68:01:01:01-68:02:04, 68:06-68:14, 68:16-68:19, 68:21:01-68:23, 68:25-68:30, 68:32-68:39, 68:41-68:71, 68:73-68:84, 69:01, 74:13 |
| 165,8,13 | 95 bp, 205 bp | 1070 bp | *11:07, 11:13, 11:54 |  |
| 17 | 205 bp | 1070 bp | *11:14, 11:50Q | *01:10, 30:01:01-30:03, 30:07-30:08, 30:10-30:16, 30:18-30:20, 30:22, 30:24-30:26, 30:30-30:45, 30:48-30:51, 30:53-30:58, 80:01-80:02 |
| 18 | 235 bp | 1070 bp | *11:15:01, 11:47 | *30:39 |
| 1914 | 175 bp, 265 bp | 1070 bp | *11:42, 11:46 |  |
| 2015 | 175 bp, 240 bp | 800 bp | *11:17, 11:36, 11:40, 11:42 | *01:43, 26:01:01-26:02, 26:04, 26:07:01-26:15, 26:17-26:19, 26:22-26:29, 26:31-26:34, 26:36-26:43:02, 26:45-26:56, 26:58-26:72, 43:01 |
| 215,7 | 100 bp | 1070 bp | *11:17, 11:19, 11:40, 11:98 | *01:01:01:01-01:01:27, 01:01:29-01:04N, 01:06, 01:08-01:12, 01:14-01:16N, 01:18N-01:33, 01:35-01:70, 01:72-01:99, 01:101-01:104, 03:41, 24:44w, 24:109w, 25:01:01-25:01:05, 25:03-25:12N, 25:14, 25:16, 26:01:01-26:01:21, 26:02-26:12, 26:14-26:18, 26:20-26:29, 26:31-26:32, 26:34-26:43:02, 26:45-26:63, 26:66-26:72, 31:03, 33:13, 36:01-36:05, 43:01, 66:05, 66:15, 74:10, 80:01-80:02 |
| 2216 | 150 bp, 230 bp | 1070 bp | *11:06, 11:18, 11:37 | *02:05:01-02:06:04, 02:06:06-02:06:10, 02:08, 02:10, 02:14, 02:21, 02:28, 02:41, 02:44, 02:51, 02:54, 02:57, 02:61, 02:72, 02:79:01-02:79:02, 02:84, 02:91, 02:99, 02:106, 02:108, 02:122, 02:126-02:127, 02:137, 02:142-02:144, 02:154, 02:169-02:170, 02:172, 02:178-02:180, 02:229, 02:232, 02:244, 02:248, 02:259, 02:271, 02:278, 02:286, 02:290, 02:295, 02:300, 02:310, 02:324, 02:328, 02:330, 02:333, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 68:15 |
| 235,6 | 80 bp | 800 bp | *11:03, 11:08, 11:20, 11:25 | *02:03:01w-02:03:04w, 02:13, 02:26, 02:38, 02:99, 02:117w, 02:148w, 02:171:01-02:171:02, 02:253w, 02:258w, 02:264w, 02:281w, 68:42w, 68:54w, 68:61w, 68:63 |
| 24 | 295 bp | 1070 bp | *11:21N, 11:48 | *01:04N, 01:41, 03:21N, 23:07N, 24:11N |
| 257 | 220 bp | 1070 bp | *11:24:01, 11:47 | *03:01:02, 30:02:05, 34:02:02, 80:01-80:02 |
| 2617 | 305 bp, 370 bp | 1070 bp | *11:26, 11:41 | *01:03, 26:32, 26:70, 33:13, 36:03, 74:10 |
| 2718 | 210 bp, 260 bp | 1070 bp | *11:17, 11:27, 11:38-11:39, 11:55, 11:94 | *01:01:01:01-01:02, 01:04N, 01:06-01:19, 01:21-01:33, 01:35-01:40, 01:42-01:62, 01:64-01:65, 01:67-01:71, 01:74-01:88, 01:90-01:101, 01:103-01:104, 03:18, 03:97, 03:122, 03:135, 24:92, 26:62, 26:72, 80:01-80:02 |
| 2819 | 240 bp, 280 bp | 800 bp | *11:15:02, 11:49 | *26:35 |
| 295 | 80 bp | 1070 bp | *11:10 | *02:55, 03:24, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 33:01:01-33:01:06, 33:03:01-33:07, 33:10-33:12, 33:14-33:16, 33:18-33:20, 33:22-33:47, 33:49-33:52, 33:54, 34:01:01-34:09, 66:01-66:10, 66:12-66:16, 68:01:01:01-68:09, 68:11N-68:12, 68:15-68:29, 68:31-68:35, 68:37-68:45, 68:47-68:65, 68:67-68:74, 68:76-68:83, 68:85, 69:01 |
| 3020 | 235 bp, 270 bp, | 1070 bp | *11:29, 11:45, 11:55 | *66:11 |
| 319,21 | 230 bp, 285 bp | 800 bp | *11:30, 11:43, 11:49 | *01:101, 03:87, 31:03, 33:13 |
| 3222 | 205 bp, 295 bp | 1070 bp | *11:31-11:32, 11:35, 11:50Q | *01:21, 02:156, 03:07, 03:42, 03:133, 30:04:01-30:04:02, 30:06, 30:09, 30:17, 30:29, 30:46, 30:51, 31:03-31:04, 33:49 |
| 335,23 | 110 bp, 280 bp, 305 bp | 1070 bp | *11:56, 11:61, 11:74 |  |
| 345,24 | 85 bp, 180 bp, 265 bp | 1070 bp | *11:46, 11:58, 11:67 |  |
| 355,25 | 105 bp, 290 bp | 1070 bp | *11:62, 11:68 |  |
| 365, 26 | 65 bp, 380 bp | 1070 bp | *11:64-11:65 | B*44:36 |
| 375,27 | 105 bp, 290 bp | 1070 bp | *11:66, 11:72 |  |
| 385,28 | 120 bp, 225 bp | 1070 bp | *11:75, 11:78N | B*07:117 |
| 39 | 280 bp | 1070 bp | *11:52Q, 11:74 |  |
| 405,29 | 90 bp, 185 bp | 1070 bp | *11:57, 11:59 | *02:156 |
| 41 | 205 bp | 1070 bp | *11:31, 11:35, 11:60 | *01:12, 01:19, 01:21, 02:156, 03:02:01-03:02:03, 03:07, 03:10, 03:31-03:32, 03:42, 03:73, 03:76, 03:82, 03:90, 03:106, 03:113, 03:133, 24:92, 30:04:01w-30:04:02w, 30:06w, 30:09, 30:17w, 30:29w, 30:46w, 31:03-31:04, 33:49 |
| 4230 | 175 bp, 560 bp | 1070 bp | *11:63, 11:69N | *01:56N |
| 438 | 445 bp | 1070 bp | *11:70 | *01:83, 02:65, 02:80, 02:114, 02:117, 02:152, 02:246, 02:279, 02:298, 02:304, 03:23:01-03:23:02, 03:89, 33:08-33:09, 68:30, 68:75, 74:04, B*57:04, B*57:32 |
| 4431 | 145 bp, 215 bp | 800 bp | *11:76, 11:106 |  |
| 45 | 210 bp | 1070 bp | *11:79 |  |
| 465 | 120 bp | 1070 bp | *11:80 |  |
| 475 | 55 bp | 1070 bp | *11:44 |  |
| 48 | 200 bp | 1070 bp | *11:71 |  |
| 4932 | 185 bp, 220 bp | 800 bp | *11:16, 11:23, 11:35, 11:57, 11:73, 11:109N | *02:156, 03:123 |
| 505,33 | 90 bp, 310 bp | 800 bp | *11:34, 11:86 |  |
| 51 | 265 bp | 1070 bp | *11:51 |  |
| 5234 | 230 bp, 545 bp | 1070 bp | *11:81, 11:108 | *01:45, 03:78 |
| 53 | 195 bp | 1070 bp | *11:100 |  |
| 5435 | 190 bp, 265 bp | 1070 bp | *11:89, 11:109N |  |
| 555,36 | 120 bp, 235 bp | 1070 bp | *11:90, 11:102 | *80:01-80:02 |
| 56 | 230 bp | 1070 bp | *11:99N |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*11 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*11 subtyping.

In addition, wells number 5, 8, 9, 11, 13, 15, 20, 23, 28, 31, 44, 49 and 50 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-A alleles 1st, 4th or 5th exon or intron nucleotide sequences are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the 1st, 4th or 5th exon or in the introns are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*11 alleles will be amplified by primer mixes 1, 3 to 7, 9 to 15, 17, 18, 20 to 32, 40 to 43, 49, 52 and 55. In addition, the B*07:102, B*08:18, B*73:01 and B*73:02 alleles will be amplified by primer mix 10, the B*44:36 allele will be amplified by primer mix 36, the B*07:117 allele will be amplified by primer mix 38, the B*57:04 and B*57:32 alleles will be amplified by primer mix 43 and the C*03:82 allele will be amplified by primer mix 1.

5Short specific PCR fragments are less intense and not as sharp as longer specific bands.

6Primer mixes 9 and 23 may give rise to nonspecific amplifications.

7Primer mixes 7, 14, 21 and 25 may give a lower yield of HLA-specific PCR product than the other A*11 primer mixes.

8Primer mixes 9, 13, 16 and 43 may have a tendency of giving rise to primer oligomer formation.

9Primer mix 5: Specific PCR product of 180 bp in the A*11:05 allele. Specific PCR product of 210 bp in the A*11:110 and the A*02:294 and 66:08 alleles.

10Primer mix 6: Specific PCR product of 195 bp in the A*11:07 and 11:22 and the A*01:79w alleles. Specific PCR product of 550 bp in the A*11:06 and 11:18 and the A*01:104w, 02:38, 02:101:01, 02:135, 02:154, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30 and 26:36 alleles.

11Primer mix 10: Specific PCR fragment of 120 bp in the A*11:10, 01:51, 25:01:01-25:12N, 25:14-25:16, 26:01:01-26:06, 26:08-26:18, 26:20-26:29, 26:32-34, 26:36-26:43:02, 26:45-26:64, 26:66-26:71N, 34:01:01-34:08, 66:01, 66:04-66:11, 66:13-66:15, B*07:102, B*08:18, B*35:178, B*73:01-73:02 alleles. Specific PCR fragment of 280 bp in the A*11:43 and the A*31:03 alleles. Specific PCR fragment of 120 and 280 bp in the A*26:35 and 33:13 alleles.

12Primer mix 15: Specific PCR fragment of 100 bp in the A*11:12 and the A*02:34w-02:35:03w, 02:56:01-02:56:02, 02:62, 02:78, 02:103, 03:01:02, 03:02, 03:06, 03:10, 03:18, 03:31, 03:50, 03:65-03:66, 03:69, 03:73, 03:75-03:76, 03:82, 03:90, 03:95, 03:106, 03:113w, 03:135, 24:24, 29:01:01:01-29:03, 29:04w, 29:05-29:18, 29:21-29:29, 29:31-29:32, 30:17, 30:52, 32:17, 66:02, 66:12, 66:16, 68:01:01:01-68:02:04, 68:06-68:14, 68:16-68:19, 68:21:01-68:23, 68:25-68:30, 68:32-68:39, 68:41-68:71, 68:73-68:84, 69:01, 74:13 alleles. Specific PCR fragment of 215 bp in the A*11:24:02-11:25, 11:31 and 11:35 and the A*01:06, 02:156, 03:01:10, 03:08, 03:29-03:30, 03:34, 03:36, 03:41, 03:62-03:63, 03:88-03:89, 30:02:01-30:02:04, 30:03, 30:07, 30:09, 30:10, 30:12-30:13, 30:22, 30:27N, 30:31-30:34, 30:44-30:45, 30:50-30:51, 30:57, 31:03-31:04, 32:04, 34:02:01, 34:03-34:04 and 34:07-34:08 alleles. Specific PCR fragment of 100 and 215 bp in the A*03:01:01:01-03:01:01:03, 03:01:03-03:01:05, 03:01:07-03:01:09, 03:01:11-03:11:31, 03:03N-03:05, 03:07, 03:09, 03:11N-03:17, 03:19-03:28, 03:33, 03:35, 03:37-03:40, 03:42-03:49, 03:51-03:61, 03:64, 03:67-03:68N, 03:70-03:72, 03:74, 03:77-03:81, 03:83-03:87, 03:91N-03:94, 03:96-03:105, 03:109-03:112, 03:114-03:133, 30:01:01-30:01:04, 30:08, 30:11:01-30:11:02, 30:14L-30:16, 30:18-30:20, 30:23-30:26, 30:30, 30:35-30:43, 30:48-30:49, 30:53-30:56, 30:58, 33:49,  and 34:09  alleles. Both PCR fragments may not always be obtained.

13Primer mix 16: Specific PCR fragment of 95 bp in the A*11:13 allele. Specific PCR fragment of 205 bp in the A*11:07 and 11:54 alleles.

14Primer mix 19: Specific PCR product of 175 bp in the A*11:42 allele. Specific PCR product of 265 bp in the A*11:46 allele.

15Primer mix 20: Specific PCR fragment of 175 bp in the A*11:42 allele. Specific PCR fragment of 240 bp in the A*11:17, 11:36 and 11:40 and the A*01:43, 26:01:01-26:02, 26:04, 26:07:01-26:15, 26:17-26:19, 26:22-26:29, 26:31-26:34, 26:36-26:43:02, 26:45-26:56 and 43:01 alleles.

16Primer mix 22: Specific PCR fragment of 150 bp in the A*11:37 allele. Specific PCR fragment of 230 bp in the A*11:17, 11:36 and 11:40 and the A*01:43, 26:01:01-26:02, 26:04, 26:07:01-26:15, 26:17-26:19, 26:22-26:29, 26:31-26:34, 26:36-26:43:02, 26:45-26:56, 26:58-26:72 and 43:01 alleles.

17Primer mix 26: Specific PCR fragment of 305 bp in the A*11:26 and the A*01:03, 26:32, 33:13, 36:03 and 74:10 alleles. Specific PCR fragment of 370 bp in the A*11:41 allele.

18Primer mix 27: Specific PCR fragment of 210 bp in the A*11:17 and the A*26:62 and 26:72 alleles. Specific PCR fragment of 260 bp in the A*11:27, 11:38-11:39, 11:55, 11:94, 01:01:01:01-01:02, 01:04N, 01:06-01:19, 01:21-01:33, 01:35-01:40, 01:42, 01:44-01:62, 01:64-01:65, 01:67-01:71, 01:74-01:88, 01:90-01:101, 01:103-01:104, 03:18, 03:97, 03:122, 03:135, 24:92, 80:01-80:02 alleles. Specific PCR fragment of 210 and 260 bp in the A*01:43 allele.

19Primer mix 28: Specific PCR fragment of 240 bp in the A*11:15:02 and 11:49 alleles. Specific PCR fragment of 280 bp in the  A*26:35 allele.

20Primer mix 30: Specific PCR fragment of 235 bp in the A*11:45 and the A*66:11 alleles. Specific PCR fragment of 265 bp in the A*11:29 and 11:55 alleles.

21Primer mix 31: Specific PCR fragment of 230 bp in the A*11:49 allele. Specific PCR fragment of 285 bp in the A*11:30 and 11:43 and the A*01:101, 03:87, 31:03 and 33:13 alleles.

22Primer mix 32: Specific PCR fragment of 205 bp in the A*11:31 and 11:35 and the A*01:21, 02:156, 03:07, 03:42, 03:133, 30:04:01-30:04:02, 30:06, 30:09, 30:17, 30:29, 30:46, 31:03-31:04 and 33:49 alleles. Specific PCR fragment of 295 bp in the A*11:32 and 11:50Q and the A*30:51 alleles.

23Primer mix 33: Specific PCR fragment of 110 bp in the A*11:61 allele. Specific PCR fragment of 280 bp in the A*11:74 allele. Specific PCR fragment of 305 bp in the A*11:56 allele.

24Primer mix 34: Specific PCR fragment of 85 bp in the A*11:58 allele. Specific PCR fragment of 180 bp in the A*11:67 allele. Specific PCR fragment of 265 bp in the A*11:46 allele.

25Primer mix 35: Specific PCR fragment of 105 bp in the A*11:62 allele. Specific PCR fragment of 290 bp in the A*11:68 allele.

26Primer mix 36: Specific PCR fragment of 65 bp in the A*11:65 and in the B*44:36 alleles. Specific PCR fragment of 380 bp in the A*11:64 allele.

27Primer mix 37: Specific PCR fragment of 105 bp in the A*11:66 allele. Specific PCR fragment of 290 bp in the A*11:72 allele.

28Primer mix 38: Specific PCR fragment of 120 bp in the A*11:75 and in the B*07:117 alleles. Specific PCR fragment of 225 bp in the A*11:78N allele.

29Primer mix 40: Specific PCR fragment of 90 bp in the A*11:59 allele. Specific PCR fragment of 185 bp in the A*11:57 and the A*02:156 alleles.

30Primer mix 42: Specific PCR fragment of 175 bp in the A*11:63 allele. Specific PCR fragment of 560 bp in the A*11:69N and the A*01:56N alleles.

31Primer mix 44: Specific PCR fragment of 145 bp in the A*11:76 allele. Specific PCR fragment of 215 bp in the A*11:106 allele.

32Primer mix 49: Specific PCR fragment of 185 bp in the A*11:16, 11:35, 11:57, 11:73 and 11:109N and the A*02:156 and 03:123 alleles. Specific PCR fragment of 220 bp in the A*11:23 allele.

33Primer mix 50: Specific PCR fragment of 90 bp in the A*11:34 alleles. Specific PCR fragment of 310 bp in the A*11:86 allele.

34Primer mix 52: Specific PCR fragment of 230 bp in the A*11:81 allele. Specific PCR fragment of 545 bp in the A*11:108 and the A*01:45 and 03:78 alleles.

35Primer mix 54: Specific PCR fragment of 190 bp in the A*11:109N allele. Specific PCR fragment of 265 bp in the A*11:89 allele.

36Primer mix 55: Specific PCR fragment of 120 bp in the A*11:102 allele. Specific PCR fragment of 235 bp in the A*11:90 and the A*80:01-80:02 alleles.

37Primer mix 3 may faintly amplify the A*03:66 and 11:08 alleles.

‘w’, may be weakly amplified.

1The A*1128 allele has been renamed to A*11:15:02. The A*11:53 allele has been shown to be indentical to the corrected sequence of A*11:02:01.

2HLA-A*11 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.7.0, January 2012.

3The A*11:13 and A*11:54 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 16.

4The A*11:58 and A*11:67 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 34.

5The A*11:62 and A*11:68 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 35.

6The A*11:63 and A*11:69N alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 42.

7The A*11:64 and A*11:65 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 36.

8The A*11:66 and A*11:72 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 37.

9The A*11:81 and 11:108 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 52.

10Primer mix 5: Specific PCR product of 180 bp in the A*11:05 allele. Specific PCR product of 210 bp in the A*11:110 and the A*02:294 and 66:08 alleles.

Primer mix 6: Specific PCR product of 195 bp in the A*11:07 and 11:22 and the A*01:79w alleles. Specific PCR product of 550 bp in the A*11:06 and 11:18 and the A*01:104w, 02:38, 02:101:01, 02:135, 02:154, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30 and 26:36 alleles.

Primer mix 10: Specific PCR fragment of 120 bp in the A*11:10, 01:51, 25:01:01-25:12N, 25:14-25:16, 26:01:01-26:06, 26:08-26:18, 26:20-26:29, 26:32-34, 26:36-26:43:02, 26:45-26:64, 26:66-26:71N, 34:01:01-34:08, 66:01, 66:04-66:11, 66:13-66:15, B*07:102, B*08:18, B*35:178, B*73:01-73:02 alleles. Specific PCR fragment of 280 bp in the A*11:43 and the A*31:03 alleles. Specific PCR fragment of 120 and 280 bp in the A*26:35 and 33:13 alleles.

Primer mix 15: Specific PCR fragment of 100 bp in the A*11:12 and the A*02:34w-02:35:03w, 02:56:01-02:56:02, 02:62, 02:78, 02:103, 03:01:02, 03:02, 03:06, 03:10, 03:18, 03:31, 03:50, 03:65-03:66, 03:69, 03:73, 03:75-03:76, 03:82, 03:90, 03:95, 03:106, 03:113w, 03:135, 24:24, 29:01:01:01-29:03, 29:04w, 29:05-29:18, 29:21-29:29, 29:31-29:32, 30:17, 30:52, 32:17, 66:02, 66:12, 66:16, 68:01:01:01-68:02:04, 68:06-68:14, 68:16-68:19, 68:21:01-68:23, 68:25-68:30, 68:32-68:39, 68:41-68:71, 68:73-68:84, 69:01, 74:13 alleles. Specific PCR fragment of 215 bp in the A*11:24:02-11:25, 11:31 and 11:35 and the A*01:06, 02:156, 03:01:10, 03:08, 03:29-03:30, 03:34, 03:36, 03:41, 03:62-03:63, 03:88-03:89, 30:02:01-30:02:04, 30:03, 30:07, 30:09, 30:10, 30:12-30:13, 30:22, 30:27N, 30:31-30:34, 30:44-30:45, 30:50-30:51, 30:57, 31:03-31:04, 32:04, 34:02:01, 34:03-34:04 and 34:07-34:08 alleles. Specific PCR fragment of 100 and 215 bp in the A*03:01:01:01-03:01:01:03, 03:01:03-03:01:05, 03:01:07-03:01:09, 03:01:11-03:11:31, 03:03N-03:05, 03:07, 03:09, 03:11N-03:17, 03:19-03:28, 03:33, 03:35, 03:37-03:40, 03:42-03:49, 03:51-03:61, 03:64, 03:67-03:68N, 03:70-03:72, 03:74, 03:77-03:81, 03:83-03:87, 03:91N-03:94, 03:96-03:105, 03:109-03:112, 03:114-03:133, 30:01:01-30:01:04, 30:08, 30:11:01-30:11:02, 30:14L-30:16, 30:18-30:20, 30:23-30:26, 30:30, 30:35-30:43, 30:48-30:49, 30:53-30:56, 30:58, 33:49,  and 34:09  alleles. Both PCR fragments may not always be obtained.

Primer mix 16: Specific PCR fragment of 95 bp in the A*11:13 allele. Specific PCR fragment of 205 bp in the A*11:07 and 11:54 alleles.

Primer mix 19: Specific PCR product of 175 bp in the A*11:42 allele. Specific PCR product of 265 bp in the A*11:46 allele.

Primer mix 20: Specific PCR fragment of 175 bp in the A*11:42 allele. Specific PCR fragment of 240 bp in the A*11:17, 11:36 and 11:40 and the A*01:43, 26:01:01-26:02, 26:04, 26:07:01-26:15, 26:17-26:19, 26:22-26:29, 26:31-26:34, 26:36-26:43:02, 26:45-26:56 and 43:01 alleles.

Primer mix 22: Specific PCR fragment of 150 bp in the A*11:37 allele. Specific PCR fragment of 230 bp in the A*11:17, 11:36 and 11:40 and the A*01:43, 26:01:01-26:02, 26:04, 26:07:01-26:15, 26:17-26:19, 26:22-26:29, 26:31-26:34, 26:36-26:43:02, 26:45-26:56, 26:58-26:72 and 43:01 alleles.

Primer mix 26: Specific PCR fragment of 305 bp in the A*11:26 and the A*01:03, 26:32, 33:13, 36:03 and 74:10 alleles. Specific PCR fragment of 370 bp in the A*11:41 allele.

Primer mix 27: Specific PCR fragment of 210 bp in the A*11:17 and the A*26:62 and 26:72 alleles. Specific PCR fragment of 260 bp in the A*11:27, 11:38-11:39, 11:55, 11:94, 01:01:01:01-01:02, 01:04N, 01:06-01:19, 01:21-01:33, 01:35-01:40, 01:42, 01:44-01:62, 01:64-01:65, 01:67-01:71, 01:74-01:88, 01:90-01:101, 01:103-01:104, 03:18, 03:97, 03:122, 03:135, 24:92, 80:01-80:02 alleles. Specific PCR fragment of 210 and 260 bp in the A*01:43 allele.

Primer mix 28: Specific PCR fragment of 240 bp in the A*11:15:02 and 11:49 alleles. Specific PCR fragment of 280 bp in the  A*26:35 allele.

Primer mix 30: Specific PCR fragment of 235 bp in the A*11:45 and the A*66:11 alleles. Specific PCR fragment of 265 bp in the A*11:29 and 11:55 alleles.

Primer mix 31: Specific PCR fragment of 230 bp in the A*11:49 allele. Specific PCR fragment of 285 bp in the A*11:30 and 11:43 and the A*01:101, 03:87, 31:03 and 33:13 alleles.

Primer mix 32: Specific PCR fragment of 205 bp in the A*11:31 and 11:35 and the A*01:21, 02:156, 03:07, 03:42, 03:133, 30:04:01-30:04:02, 30:06, 30:09, 30:17, 30:29, 30:46, 31:03-31:04 and 33:49 alleles. Specific PCR fragment of 295 bp in the A*11:32 and 11:50Q and the A*30:51 alleles.

Primer mix 33: Specific PCR fragment of 110 bp in the A*11:61 allele. Specific PCR fragment of 280 bp in the A*11:74 allele. Specific PCR fragment of 305 bp in the A*11:56 allele.

Primer mix 34: Specific PCR fragment of 85 bp in the A*11:58 allele. Specific PCR fragment of 180 bp in the A*11:67 allele. Specific PCR fragment of 265 bp in the A*11:46 allele.

Primer mix 35: Specific PCR fragment of 105 bp in the A*11:62 allele. Specific PCR fragment of 290 bp in the A*11:68 allele.

Primer mix 36: Specific PCR fragment of 65 bp in the A*11:65 and in the B*44:36 alleles. Specific PCR fragment of 380 bp in the A*11:64 allele.

Primer mix 37: Specific PCR fragment of 105 bp in the A*11:66 allele. Specific PCR fragment of 290 bp in the A*11:72 allele.

Primer mix 38: Specific PCR fragment of 120 bp in the A*11:75 and in the B*07:117 alleles. Specific PCR fragment of 225 bp in the A*11:78N allele.

Primer mix 40: Specific PCR fragment of 90 bp in the A*11:59 allele. Specific PCR fragment of 185 bp in the A*11:57 and the A*02:156 alleles.

Primer mix 42: Specific PCR fragment of 175 bp in the A*11:63 allele. Specific PCR fragment of 560 bp in the A*11:69N and the A*01:56N alleles.

Primer mix 44: Specific PCR fragment of 145 bp in the A*11:76 allele. Specific PCR fragment of 215 bp in the A*11:106 allele.

Primer mix 49: Specific PCR fragment of 185 bp in the A*11:16, 11:35, 11:57, 11:73 and 11:109N and the A*02:156 and 03:123 alleles. Specific PCR fragment of 220 bp in the A*11:23 allele.

Primer mix 50: Specific PCR fragment of 90 bp in the A*11:34 alleles. Specific PCR fragment of 310 bp in the A*11:86 allele.

Primer mix 52: Specific PCR fragment of 230 bp in the A*11:81 allele. Specific PCR fragment of 545 bp in the A*11:108 and the A*01:45 and 03:78 alleles.

Primer mix 54: Specific PCR fragment of 190 bp in the A*11:109N allele. Specific PCR fragment of 265 bp in the A*11:89 allele.

Primer mix 55: Specific PCR fragment of 120 bp in the A*11:102 allele. Specific PCR fragment of 235 bp in the A*11:90 and the A*80:01-80:02 alleles.

12This lot of the HLA-A*03 subtyping kit cannot distinguish the A*11:77 and the A*11:02:01-11:02:03 alleles.

‘w’, may be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 43 does not amplify the A*11:77 and the A*03:26 alleles. Thus, this lot of the HLA-A*03 subtyping kit cannot distinguish the A*11:77 and the A*11:02:01-11:02:03 alleles.

Change in revision R02 compared to R01:

- Primer mix 3 may faintly amplify the A*03:66 and 11:08 alleles. A footnote has been added in the Specificity Table.