Olerup SSP® ## PHOTO DOCUMENT

Olerup SSP® HLA-A*02

Prod. No:	101.412-24/24u

101.412-04/04u

Lot No:    	29S

Expiry Date:    	2015-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 96 primer mixes used for HLA-A*02 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*02 alleles3 | Other amplified HLA-A alleles4 |
| --- | --- | --- | --- | --- |
| 18,76 | 360 bp | 800 bp | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:01:81, 02:01:83-02:06:13, 02:08-02:09, 02:12-02:14, 02:16, 02:19-02:22:02, 02:24:01-02:32N, 02:34, 02:36-02:38, 02:40:01-02:46, 02:49-02:68, 02:70-02:77, 02:79:01-02:80, 02:82N-02:86, 02:88N-02:89, 02:91-02:97:02, 02:99, 02:101:01-02:102, 02:104-02:107, 02:109, 02:111, 02:113N-02:123, 02:125N-02:127, 02:131-02:135, 02:137-02:145, 02:147, 02:149-02:190, 02:192-02:211:01, 02:212-02:218, 02:220-02:241, 02:243, 02:245-02:246, 02:248-02:249, 02:251-02:254, 02:256-02:260, 02:262-02:264, 02:266-02:267, 02:269-02:281, 02:283-02:296, 02:299, 02:301N-02:302, 02:304-02:307, 02:309-02:316, 02:318, 02:320-02:330, 02:332-02:334, 02:336-02:346, 02:348-02:350N, 02:352-02:360, 02:362-02:368, 02:370-02:392, 02:394-02:397, 02:399-02:402 | *01:104, 03:09, 03:89, 03:108, 11:06, 11:18, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:36, 26:78, 29:19, 30:13, 30:16, 30:44, 30:46, 32:08, 33:24, 68:05, 68:15, 68:20, 74:06, 80:01:01:01w |
| 2 | 240 bp | 1070 bp | *02:01:01:01-02:01:02, 02:01:04-02:01:15, 02:01:17-02:01:19, 02:01:21-02:01:81, 02:01:83-02:04, 02:07:01-02:07:02, 02:09, 02:12-02:13, 02:15N-02:20:02, 02:22:01-02:22:02, 02:24:01-02:27, 02:29-02:34, 02:36-02:40:02, 02:42-02:43N, 02:45-02:47, 02:49-02:50, 02:52-02:53N, 02:55-02:56:02, 02:58-02:60:02, 02:62-02:68, 02:70-02:71, 02:73-02:75, 02:77, 02:80, 02:82N-02:83N, 02:85-02:86, 02:88N-02:89, 02:92-02:97:02, 02:101:01-02:105, 02:107, 02:109-02:111, 02:113N-02:121, 02:123, 02:125N, 02:130-02:135, 02:138-02:141, 02:145, 02:147-02:153, 02:155-02:168, 02:171:01-02:171:02, 02:173-02:177, 02:181-02:228, 02:230-02:231, 02:233-02:242, 02:245-02:246, 02:249-02:258, 02:260-02:270, 02:272-02:277, 02:279-02:285, 02:287-02:289, 02:291-02:294, 02:296, 02:299, 02:301N-02:307, 02:309, 02:311-02:318, 02:320-02:323, 02:325-02:327, 02:329, 02:332, 02:334-02:336, 02:338-02:343, 02:345-02:354, 02:356N-02:357, 02:360-02:365, 02:367-02:372, 02:374-02:375, 02:377-02:381, 02:383-02:386, 02:388-02:392, 02:395N-02:397, 02:399-02:403 | *03:09, 03:89, 03:108, 74:06 |
| 36 | 175 bp | 1070 bp | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:01:81, 02:01:83-02:03:04, 02:05:01-02:16, 02:18, 02:20:01-02:22:02, 02:24:01-02:38, 02:40:01-02:43N, 02:45-02:56:02, 02:58-02:64, 02:66-02:78, 02:80-02:85, 02:87-02:97:02, 02:99, 02:101:01-02:107, 02:109, 02:111-02:134, 02:136-02:148, 02:150-02:151, 02:153-02:155, 02:157-02:189, 02:191-02:193, 02:195-02:223N, 02:225N-02:267, 02:269-02:299, 02:301N-02:302, 02:304-02:308, 02:310-02:333, 02:335-02:337, 02:339-02:370, 02:372-02:392, 02:394-02:399, 02:401-02:403 | *68:02:01:01-68:02:05, 68:15, 68:18N, 68:23, 68:27-68:28, 68:31, 68:34, 68:40, 68:48, 68:51, 68:53-68:54, 68:60-68:62, 68:64, 68:67, 68:74, 68:77-68:78, 68:80-68:82, 68:86, 68:92, 68:97, 69:01 |
| 45 | 80 bp | 1070 bp | *02:01:01:01-02:01:15, 02:01:18-02:01:19, 02:01:21-02:01:51, 02:01:53-02:01:62, 02:01:64-02:01:73, 02:01:75-02:01:81, 02:01:83-02:01:86, 02:03:01-02:04, 02:06:01-02:06:11, 02:06:13-02:07:02, 02:09-02:13, 02:15N-02:22:02, 02:24:01-02:35:01, 02:35:03-02:44, 02:45w, 02:46, 02:49, 02:51-02:54, 02:56:01w, 02:57-02:62, 02:64-02:77, 02:78w, 02:79:01-02:97:02, 02:99, 02:101:01-02:102, 02:104-02:114, 02:116-02:121, 02:123-02:128, 02:130-02:143, 02:145-02:153, 02:156-02:168, 02:170-02:171:02, 02:173-02:175, 02:177-02:178, 02:180-02:184, 02:187-02:194, 02:196-02:204, 02:207-02:208, 02:210-02:228, 02:230-02:231, 02:233-02:270, 02:272-02:285, 02:287-02:314N, 02:316-02:319, 02:322-02:323, 02:325-02:336, 02:338-02:343, 02:345-02:358, 02:360-02:371, 02:374-02:375, 02:377, 02:379-02:381, 02:383-02:392, 02:394-02:403 | *01:69:02w, 01:83:02, 03:01:03w, 03:23:01, 03:89, 11:121, 26:07:02, 29:48, 33:01:02, 33:08-33:09, 68:30, 74:04 |
| 55 | 125 bp | 1070 bp | *02:01:01:01-02:01:15, 02:01:17-02:01:19, 02:01:21-02:01:81, 02:01:83-02:01:86, 02:04, 02:06:01-02:07:02, 02:09-02:11:04, 02:14-02:18, 02:20:01-02:21, 02:25-02:26, 02:28-02:36, 02:39-02:43N, 02:45-02:46, 02:48, 02:51-02:53N, 02:55-02:62, 02:64, 02:66-02:67, 02:68w, 02:69-02:97:02, 02:99, 02:101:01-02:101:02, 02:103, 02:105-02:114, 02:116, 02:118-02:121, 02:123-02:134, 02:138-02:141, 02:143-02:147, 02:149-02:151, 02:153, 02:156-02:171:02, 02:173-02:178, 02:180-02:181, 02:183-02:185, 02:187-02:190, 02:192-02:208, 02:210-02:221, 02:223N-02:225N, 02:227N-02:228, 02:231, 02:233-02:252, 02:254-02:257, 02:259-02:263, 02:265-02:266, 02:268-02:270, 02:272-02:273, 02:275-02:279, 02:282-02:285, 02:287-02:308, 02:310-02:314N, 02:316-02:319, 02:321N-02:322, 02:325-02:336, 02:338-02:343, 02:346-02:354, 02:356N-02:358, 02:360-02:369, 02:371-02:372, 02:374-02:375, 02:377-02:381, 02:383-02:392, 02:394-02:401 | *03:17:01, 23:01:01-23:01:11, 23:04-23:56, 24:13:01-24:13:02, 24:18, 24:24, 24:94, 24:188, 24:207, 24:228, 30:61, 31:67-31:68, 32:28, 33:32:01, 68:08:01-68:08:02, 68:63, 69:01, B*51:136, C*07:204 |
| 65,6 | 75 bp | 800 bp | *02:02-02:03:04, 02:05:01-02:05:06, 02:08, 02:22:01-02:22:02, 02:47, 02:49-02:50, 02:63, 02:102, 02:104, 02:115, 02:117, 02:122, 02:136, 02:148, 02:154-02:155, 02:172, 02:179, 02:186, 02:191, 02:209, 02:229-02:230, 02:232, 02:253, 02:258, 02:264, 02:267, 02:281, 02:286, 02:315, 02:320, 02:323-02:324, 02:337, 02:344-02:345, 02:355, 02:359, 02:370, 02:373N, 02:376, 02:382 | *24:138, 68:01:01:01-68:01:13, 68:01:15-68:07, 68:10-68:25, 68:27-68:47, 68:49N, 68:51-68:62, 68:64-68:99 |
| 77 | 175 bp | 1070 bp | *02:02, 02:05:01-02:05:06, 02:08, 02:14, 02:47, 02:63, 02:102, 02:115, 02:154-02:155, 02:172, 02:179, 02:186, 02:209, 02:229, 02:232, 02:286, 02:320, 02:324, 02:337, 02:344, 02:359, 02:373N, 02:376 |  |
| 811 | 415 bp, 505 bp | 1070 bp | *02:02, 02:05:01-02:05:06, 02:14, 02:32N, 02:47, 02:63, 02:102, 02:115, 02:154-02:155, 02:172, 02:179, 02:186, 02:209, 02:229, 02:232, 02:271, 02:286, 02:320, 02:324, 02:337, 02:344, 02:359, 02:373N, 02:376 |  |
| 95 | 105 bp | 1070 bp | *02:03:01, 02:03:02w, 02:03:03-02:03:04, 02:25, 02:38, 02:117, 02:148, 02:171:01-02:171:02, 02:230, 02:253, 02:258, 02:264, 02:267, 02:280-02:281, 02:315, 02:345, 02:355, 02:370 | *24:99, 26:10 |
| 105 | 115 bp | 800 bp | *02:03:01, 02:03:02w, 02:03:03-02:03:04, 02:13, 02:26, 02:38, 02:99, 02:117, 02:148, 02:171:01-02:171:02, 02:230, 02:253, 02:258, 02:264, 02:280-02:281, 02:315, 02:323, 02:345, 02:355, 02:370, 02:402 | *03:17:01, 24:18, 24:56, 24:177, 26:10, 68:42, 68:54, 68:61, 68:63, B*51:136, C*12:37 |
| 118,9,12 | 225 bp, 350 bp | 1070 bp | *02:04, 02:17:01-02:17:02, 02:57, 02:65, 02:89, 02:108, 02:110, 02:152, 02:268, 02:300, 02:303, 02:334 | *23:12, 24:28, 24:30, 24:42, 24:89, 29:19, 29:48, 32:08, 33:24, 68:05, 68:20, 74:06 |
| 1213 | 195 bp, 235 bp | 1070 bp | *02:05:01-02:06:13, 02:08, 02:10, 02:14, 02:21, 02:28, 02:41, 02:44, 02:51, 02:54, 02:57, 02:61, 02:72, 02:79:01-02:79:02, 02:84-02:85, 02:91, 02:99, 02:106, 02:108, 02:122, 02:126-02:127, 02:137, 02:142-02:144, 02:154, 02:169-02:170, 02:172, 02:178-02:180, 02:229, 02:232, 02:244, 02:248, 02:259, 02:271, 02:278, 02:286, 02:290, 02:295, 02:300, 02:310, 02:324, 02:328, 02:330, 02:333, 02:337, 02:344, 02:355, 02:358-02:359, 02:373N, 02:376, 02:382, 02:387, 02:398 | *11:06, 11:18, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:78, 68:05, 68:15, 68:20 |
| 138 | 145 bp | 800 bp | *02:05:01-02:05:06, 02:08, 02:14, 02:84, 02:154, 02:172, 02:179, 02:229, 02:232, 02:271, 02:286, 02:324, 02:337, 02:344, 02:359, 02:373N, 02:376, 02:382 |  |
| 145,8,14 | 95 bp, 170 bp | 800 bp | *02:07:01-02:07:02, 02:15N, 02:18, 02:91, 02:103, 02:112, 02:130, 02:191, 02:219, 02:255, 02:261, 02:265, 02:282, 02:319, 02:322, 02:335, 02:369, 02:403 |  |
| 155,8,15 | 125 bp, 265 bp, 305 bp | 800 bp | *02:21, 02:87, 02:96, 02:112, 02:129, 02:136, 02:186 |  |
| 165,16 | 110 bp, 155 bp | 800 bp | *02:09, 02:49-02:50, 02:73, 02:93, 02:122, 02:156, 02:172, 02:279 | *11:119, 23:31, 23:45, 23:55, 24:15, 24:41, 24:51, 24:92, 26:10, 32:28, 68:02:01:01-68:02:05, 68:15, 68:18N, 68:25, 68:27-68:28, 68:31, 68:34, 68:40, 68:44, 68:48-68:49N, 68:51, 68:53-68:54, 68:60-68:62, 68:64, 68:67, 68:74, 68:77-68:78, 68:80-68:82, 68:86, 68:92, 68:97, B*44:59:02, B*44:136, B*51:136, C*07:204, C*12:37 |
| 1717 | 205 bp, 360 bp | 1070 bp | *02:10, 02:17:01-02:17:02, 02:39, 02:83N, 02:108, 02:110, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303, 02:398 | *23:12, 24:28, 24:30, 24:42, 24:89 |
| 1810 | 175 bp | 1070 bp | *02:11:01, 02:11:03-02:11:04, 02:29, 02:69, 02:128, 02:297-02:298, 02:308, 02:395N |  |
| 195,18 | 125 bp, 165 bp | 1070 bp | *02:12-02:13, 02:19, 02:27, 02:37-02:38, 02:44, 02:54, 02:142, 02:271, 02:280, 02:283 | *11:119, 24:02:01:01-24:02:15, 24:02:17-24:02:41, 24:02:43-24:05, 24:07-24:11N, 24:14-24:15, 24:17, 24:19-24:20, 24:23, 24:25-24:53, 24:55-24:64, 24:66-24:86N, 24:88-24:90N, 24:92-24:93, 24:95-24:106, 24:108-24:109, 24:111-24:124, 24:126-24:132N, 24:134-24:137, 24:139-24:157, 24:159, 24:161-24:166, 24:168-24:187, 24:189-24:203, 24:205-24:206, 24:209-24:210, 24:212, 24:214-24:226, 24:229-24:230, 33:19, 68:09, 68:26 |
| 2019 | 135 bp, 295 bp | 1070 bp | *02:49, 02:200 |  |
| 2120 | 220 bp, 255 bp | 1070 bp | *02:15N, 02:291, 02:356N |  |
| 22 | 160 bp | 1070 bp | *02:19, 02:36-02:37, 02:54, 02:255 | *23:01:01-23:02, 23:05-23:42, 23:44-23:56, 24:02:01:01-24:02:41, 24:02:43-24:02:64, 24:04-24:09N, 24:11N, 24:13:01-24:15, 24:17, 24:19-24:20, 24:23w, 24:24-24:32, 24:34-24:64, 24:66-24:74, 24:76-24:93, 24:95-24:109, 24:111-24:124, 24:126-24:137, 24:139-24:157, 24:159-24:166, 24:168-24:203, 24:205-24:206, 24:209, 24:212-24:218, 24:220-24:230, 33:19, 68:26, 68:65 |
| 2321 | 135 bp, 180 bp | 800 bp | *02:17:01-02:17:02, 02:31, 02:108, 02:110, 02:161, 02:268, 02:300, 02:303, 02:387 | *24:94, 24:138, 24:188, 24:228 |
| 245,22 | 75 bp, 115 bp, 160 bp | 1070 bp | *02:27, 02:230, 02:233 | *11:119, 24:59, 24:190, 24:210, 24:219, 24:229, C*07:204 |
| 2523 | 125bp, 145 bp, 165 bp | 1070 bp | *02:16, 02:131, 02:226N, 02:283 | *24:55 |
| 265,24 | 125 bp, 165 bp, 260 bp | 800 bp | *02:33, 02:52, 02:198 |  |
| 275,25 | 95 bp, 145 bp | 1070 bp | *02:10, 02:28, 02:50, 02:52, 02:73, 02:93, 02:95, 02:110, 02:114, 02:117, 02:122, 02:155-02:156, 02:185, 02:242, 02:244, 02:279, 02:282, 02:300, 02:304, 02:339 | *03:123, 11:16, 11:35, 11:57, 11:73, 24:131, 24:138, 24:188, 24:218, 30:13, 30:16, 30:44, 30:46, 68:01:01:01-68:11N, 68:13:01-68:48, 68:51-68:99 |
| 28 | 235 bp | 800 bp | *02:08, 02:20:01-02:20:02, 02:55-02:56:02, 02:62, 02:78, 02:103, 02:128, 02:169, 02:195 |  |
| 2926 | 220 bp, 300 bp | 1070 bp | *02:45-02:46, 02:48, 02:56:01w-02:56:02w, 02:78w, 02:92, 02:103w, 02:129, 02:169w, 02:180, 02:195w, 02:358, 02:369 |  |
| 3027 | 130 bp, 160 bp | 800 bp | *02:43N, 02:104, 02:163 |  |
| 315,6,28 | 95 bp, 220 bp | 800 bp | *02:45, 02:56:01-02:56:02, 02:78, 02:82N, 02:103, 02:169, 02:195 | *23:08N |
| 3229 | 205 bp, 230 bp, 255 bp | 1070 bp | *02:46-02:48, 02:70, 02:129, 02:176 |  |
| 337,8,30 | 230 bp, 325 bp | 1070 bp | *02:19, 02:36-02:37, 02:54, 02:165, 02:168, 02:255, 02:400 |  |
| 345,6,9,31 | 120 bp, 165 bp, 180 bp | 800 bp | *02:34-02:35:02, 02:56:01-02:56:02, 02:62, 02:78, 02:88N, 02:103, 02:395N | *30:01:01-30:01:07, 30:08, 30:11:01-30:11:02, 30:14L-30:20, 30:23-30:24, 30:26, 30:30-30:31, 30:35-30:43, 30:48-30:49, 30:52-30:56, 30:58-30:60, 30:62-30:63, 30:65, 68:01:11, C*03:82 |
| 355,8,32 | 110 bp, 155 bp | 1070 bp | *02:40:01-02:40:02, 02:51, 02:77, 02:130 | *23:01:01-23:01:11, 23:02w, 23:04-23:23, 23:25-23:33, 23:35-23:56, 24:24, 24:71, 31:67-31:68, 32:28, 33:32:01, 68:51w |
| 365,33 | 85 bp, 445 bp | 1070 bp | *02:24:01-02:24:02, 02:65, 02:94N, 02:135, 02:137, 02:152, 02:309 | *01:104, 03:09, 03:89, 03:108, 11:06, 11:18, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:36, 26:78, 29:19, 29:48, 30:13, 30:16, 30:44, 30:46, 32:08, 33:24, 74:06, 80:01:01:01w |
| 37 | 140 bp | 1070 bp | *02:38, 02:68, 02:101:01-02:101:02, 02:154 | *11:119, 23:10w, 24:10:01-24:10:02, 24:46w, 24:91, 24:210, 26:10, 68:61 |
| 385,34 | 125 bp, 225 bp | 800 bp | *02:41, 02:80, 02:117, 02:289, 02:304, 02:351 |  |
| 395,6,35 | 170 bp, 260 bp 300 bp | 1070 bp | *02:18, 02:159, 02:170, 02:293Q, 02:364 |  |
| 405,8,36 | 90 bp, 210 bp | 1070 bp | *02:40:01-02:40:02, 02:51, 02:67, 02:130 | *29:22, 33:22 |
| 418 | 185 bp | 1070 bp | *02:02, 02:05:01-02:05:06, 02:08, 02:14, 02:17:01-02:17:02, 02:47, 02:57-02:58, 02:63, 02:75, 02:102, 02:108, 02:110, 02:115, 02:154-02:155, 02:172, 02:179, 02:186, 02:209, 02:229, 02:232, 02:271, 02:286, 02:300, 02:303, 02:320, 02:324, 02:337, 02:344, 02:359, 02:373N, 02:376 | *24:94, 24:138, 24:188, 24:228, C*12:94, C*14:48 |
| 427 | 225 bp | 800 bp | *02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:103 | *01:13, 01:17, 03:01:01:01-03:07, 03:09-03:11N, 03:13-03:31, 03:33-03:35, 03:37-03:40, 03:42-03:58, 03:60-03:71, 03:73-03:87, 03:90-03:106, 03:109-03:141, 03:143-03:151, 03:153-03:162N, 30:55, 34:08, 68:71, 74:13 |
| 4337 | 180 bp, 225 bp | 800 bp | *02:03:01-02:03:04, 02:22:01-02:22:02, 02:49, 02:71, 02:104, 02:117, 02:136, 02:148, 02:191, 02:230, 02:253, 02:258, 02:264, 02:267, 02:281, 02:315, 02:323, 02:345, 02:355, 02:370, 02:382, 02:402 | *26:22, 66:09, 68:83 |
| 445,38 | 125 bp, 155 bp, 205 bp | 800 bp | *02:59, 02:203, 02:222N |  |
| 455,39 | 105 bp, 185 bp | 1070 bp | *02:19, 02:39, 02:44, 02:60:01-02:60:02, 02:79:01-02:79:02, 02:86, 02:254, 02:400 | *01:20, 01:66, 03:95, 24:14, 24:93 |
| 465,7,40 | 70 bp, 205 bp | 1070 bp | *02:61, 02:66 |  |
| 4741 | 165 bp, 205 bp, 255 bp, 305 bp | 800 bp | *02:35:01-02:35:03, 02:48, 02:63, 02:78, 02:90, 02:144, 02:205, 02:207, 02:331 |  |
| 485 | 115 bp | 1070 bp | *02:64, 02:177-02:178 |  |
| 495.42 | 115 bp, 145 bp | 1070 bp | *02:27, 02:38, 02:101:01-02:101:02, 02:154, 02:267 | *11:119, 23:10w, 24:10:01-24:10:02, 24:46w, 24:59, 24:190, 24:210, 24:229, 26:10, 68:61, C*07:204 |
| 50 | 145 bp | 800 bp | *02:01:01:02L |  |
| 51 | 410 bp | 1070 bp | *02:65, 02:80, 02:114, 02:117, 02:152, 02:246, 02:279, 02:304 | *01:104, 03:09, 03:89, 03:108, 11:06, 11:18, 29:19, 29:48, 32:08, 33:24, 68:05, 68:20, 74:06, 80:01:01:01w |
| 52 | 185 bp | 800 bp | *02:76:01-02:76:02, 02:319 | *30:22 |
| 53 | 225 bp | 1070 bp | *02:14, 02:17:01-02:17:02, 02:57-02:58, 02:108, 02:110, 02:300, 02:303 | *03:75, 23:01:01-23:01:11, 23:03:01-23:13, 23:14:02-23:33, 23:35-23:56, 24:13:01, 24:18, 24:24, 24:94, 24:188, 24:207, 24:228, 29:07, 30:52, 31:29 |
| 547,8,43 | 240 bp, 350 bp | 800 bp | *02:74:01-02:74:02, 02:189-02:190 |  |
| 556 | 200 bp | 1070 bp | *02:01:83, 02:13, 02:26, 02:30, 02:40:01-02:40:02, 02:51, 02:99, 02:130, 02:226N, 02:323, 02:399, 02:402 | *03:95, 29:22, 33:22 |
| 5644 | 150 bp, 275 bp | 1070 bp | *02:53N, 02:81, 02:124, 02:175, 02:181 |  |
| 575,7,45 | 105 bp, 145 bp | 1070 bp | *02:42, 02:160, 02:175, 02:310 |  |
| 58 | 255 bp | 800 bp | *02:11:01-02:11:04, 02:35:01-02:35:03, 02:48, 02:69, 02:78, 02:90, 02:128, 02:297-02:298, 02:308, 02:331 |  |
| 598,46 | 145 bp, 190 bp | 800 bp | *02:97:01-02:97:02, 02:305N |  |
| 608 | 345 bp | 1070 bp | *02:50, 02:65, 02:73, 02:122, 02:135, 02:152, 02:245, 02:334 | *25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:36, 26:78, 29:19, 29:48, 32:08, 33:24, 68:05, 68:15, 68:20, 74:06 |
| 617,8 | 185 bp | 1070 bp | *02:50, 02:73, 02:122, 02:156, 02:245, 02:334, 02:338 | *03:123, 11:16, 11:35, 11:57, 11:73, 68:01:01:01-68:11N, 68:13:01-68:43, 68:45-68:48, 68:51-68:82, 68:84-68:99 |
| 62 | 180 bp | 800 bp | *02:69, 02:179 | *68:01:01:01-68:32, 68:34-68:95, 68:97-68:99 |
| 635,47 | 95 bp, 160 bp | 800 bp | *02:105, 02:259, 02:301N | *03:51, 30:37, 32:46 |
| 645,48 | 125 bp, 175 bp, 260 bp | 1070 bp | *02:106, 02:145, 02:164, 02:187, 02:221 | *01:44, 03:44, 23:41, 24:72, 29:15 |
| 6549 | 145 bp, 185 bp, 250 bp | 1070 bp | *02:107, 02:160, 02:202, 02:251 |  |
| 66 | 170 bp | 800 bp | *02:109 |  |
| 675,50 | 105bp, 130 bp, 180 bp, 300 bp | 800 bp | *02:111, 02:177, 02:330, 02:350N |  |
| 6851 | 210 bp, 260 bp | 1070 bp | *02:113N, 02:184, 02:321N |  |
| 6952 | 170 bp, 225 bp, 335 bp | 1070 bp | *02:114, 02:166, 02:246, 02:260, 02:279 |  |
| 705,53 | 125 bp, 180 bp, 285 bp | 800 bp | *02:53N, 02:115, 02:192, 02:269, 02:367 |  |
| 718 | 260 bp | 800 bp | *02:116 |  |
| 725,54 | 110 bp, 230 bp, 260 bp | 1070 bp | *02:19, 02:44, 02:118, 02:135, 02:149, 02:152, 02:183, 02:189-02:190, 02:309, 02:402 |  |
| 735,55 | 80 bp, 200 bp, 250 bp | 800 bp | *02:119, 02:158, 02:263 |  |
| 7456 | 175 bp, 250 bp | 800 bp | *02:120, 02:187, 02:223N |  |
| 755,8, 57 | 100 bp, 165 bp, 205 bp | 1070 bp | *02:121, 02:166, 02:373N | *23:47, 68:14, 68:49N, 68:81, B*44:59:02 |
| 76 | 220 bp | 800 bp | *02:50, 02:122, 02:143, 02:225N |  |
| 775,58 | 85 bp, 110 bp | 800 bp | *02:123, 02:162, 02:295, 02:344 |  |
| 785,59 | 75 bp, 130 bp, 240 bp | 1070 bp | *02:124, 02:193, 02:213 |  |
| 798,60 | 215 bp, 270 bp | 1070 bp | *02:53N, 02:125N, 02:184, 02:192, 02:227N |  |
| 805,61 | 90 bp, 135 bp, 230 bp | 1070 bp | *02:126, 02:194, 02:214 |  |
| 815,7,62 | 110 bp, 235 bp, 315 bp | 800 bp | *02:127, 02:165, 02:167, 02:183, 02:299, 02:400 |  |
| 8263 | 130 bp, 255 bp, 295 bp | 1070 bp | *02:131, 02:199, 02:291 |  |
| 835,64 | 120 bp, 240 bp | 800 bp | *02:132, 02:215, 02:237 | *03:95 |
| 847,65 | 180 bp, 205 bp, 245 bp | 800 bp | *02:133, 02:164, 02:250N, 02:315 | *01:44, 03:44 |
| 8566 | 170 bp, 260 bp | 800 bp | *02:134, 02:314N, 02:376 | *32:48N |
| 8667 | 180 bp, 255 bp, 310 bp | 800 bp | *02:135, 02:174, 02:228, 02:309 |  |
| 8768 | 155 bp, 210 bp | 800 bp | *02:138, 02:181, 02:284N |  |
| 8869 | 135 bp, 240 bp | 800 bp | *02:139, 02:188, 02:235 | B*15:67, B*35:110, C*12:94, C*14:48 |
| 895,70 | 120 bp, 200 bp | 800 bp | *02:140, 02:182, 02:227N | *24:133, 31:48, 33:15 |
| 905,71 | 100 bp, 130 bp, 190 bp | 800 bp | *02:72, 02:141, 02:161, 02:206, 02:275 |  |
| 915 | 110 bp | 800 bp | *02:50, 02:73, 02:93, 02:122, 02:156, 02:162, 02:172, 02:211:01-02:211:02, 02:279 | *11:119, 23:31, 23:45, 23:55, 24:15, 24:41, 24:51, 24:92, 26:10, 32:28, 68:02:01:01-68:02:05, 68:15, 68:18N, 68:25, 68:27-68:28, 68:31, 68:34, 68:40, 68:44, 68:48-68:49N, 68:51, 68:53-68:54, 68:60-68:62, 68:64, 68:67, 68:74, 68:76-68:78, 68:80-68:82, 68:86, 68:92, 68:97, B*44:59:02, B*44:136, B*51:136, C*07:204, C*12:37 |
| 9272 | 180 bp, 260 bp | 1070 bp | *02:146, 02:173 |  |
| 9373 | 130 bp, 250 bp, 295 bp | 800 bp | *02:147, 02:157, 02:163, 02:339 | *23:52, 24:73, 24:157, B*15:173, B*18:63 |
| 945,7,74 | 110 bp, 210 bp, 360 bp | 800 bp | *02:150, 02:197, 02:325 |  |
| 95 | 135 bp | 800 bp | *02:52, 02:105, 02:147, 02:151, 02:188, 02:339 | *68:06, C*12:94, C*14:48 |
| 967,75 | 135 bp, 255 bp | 1070 bp | *02:153, 02:196, 02:366N |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*02 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*02 subtyping.

In addition, wells number 6, 10, 13-16, 23, 26, 28, 30, 31, 34, 38, 42-44, 47, 50, 52, 54, 58, 59, 62, 63, 66, 67, 70, 71, 73, 74, 76, 77, 81, 83-91 and 93-95 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA-A alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*02 alleles will be amplified by primer mixes 1 to 6, 9 to 12, 16, 17, 19, 22 to 25, 27, 31, 34 to 37, 40 to 43, 45, 49, 51 to 53, 55, 60 to 64, 75, 83 to 85, 89, 91, 93 and 95. In addition, the B*15:67 and B*35:110 alleles will be amplified by primer mix 88, the B*15:173 and B*18:63 alleles will be amplified by primer mix 93, the B*44:59:02 allele will be amplified by primer mixes 16, 75 and 91, the B*44:136 allele will be amplified by primer mixes 16 and 91, the B*51:136 allele will be amplified by primer mixes 5, 10, 16 and 91, the C*03:82 allele will be amplified by primer mix 34, the C*07:204 allele will be amplified by primer mixes 5, 16, 24, 49 and 91, the C*12:37 allele will be amplified by primer mixes 10, 16 and 91 and the C12:94 and C*14:48  alleles will be amplified by primer mixes 41, 88 and 95.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6The specific primers in primer mixes 3, 6, 31, 34, 39 and 55 may give rise to a lower yield of HLA-specific PCR product than the other A*02 primer mixes.

7Primer mixes 7, 33, 42, 46, 54, 57, 61, 81, 84, 94 and 96 have a tendency to giving rise to primer oligomer formation.

8Primer mixes 1, 11, 13, 14, 15, 33, 40, 41, 54, 59, 60, 61, 71, 75 and 79 may have tendencies of unspecific amplifications.

9Primer mixes 11 and 34 may give rise to a long fragment of approx. 600 bp in some HLA-A alleles. This band should not be considered in the interpretatation of HLA-A*02 typings.

10Primer mix 18 may faintly amplify the C*04:01:01:01-04:01:41 alleles.

11Primer mix 8: Specific PCR fragment of 415 bp in the A*02:02, 02:05:01-02:05:06, 02:14, 02:47, 02:63, 02:102, 02:115, 02:154-02:155, 02:172, 02:179, 02:186, 02:209, 02:229, 02:232, 02:271, 02:286, 02:320, 02:324, 02:337, 02:344, 02:359, 02:373N and 02:376 alleles. Specific PCR fragment of 505 bp in the A*02:32N allele.

12Primer mix 11: Specific PCR fragment of 225 bp in the A*02:89 allele. Specific PCR fragment of 350 bp in the A*02:04, 02:17:01-02:17:02, 02:57, 02:65, 02:108, 02:110, 02:152, 02:268, 02:300, 02:303 and 02:334 and the A*23:12, 24:28, 24:30, 24:42, 24:89, 29:19, 29:48, 32:08, 33:24, 68:05, 68:20 and 74:06 alleles.

13Primer mix 12: Specific PCR fragment of 195 bp in A*02:85 allele. Specific PCR fragment of 235 bp in the A*02:05:01-02:06:13, 02:08, 02:10, 02:14, 02:21, 02:28, 02:41, 02:44, 02:51, 02:54, 02:57, 02:61, 02:72, 02:79:01-02:79:02, 02:84, 02:91, 02:99, 02:106, 02:108, 02:122, 02:126-02:127, 02:137, 02:142-02:144, 02:154, 02:169-02:170, 02:172, 02:178-02:180, 02:229, 02:232, 02:244, 02:248, 02:259, 02:271, 02:278, 02:286, 02:290, 02:295, 02:300, 02:310, 02:324, 02:328, 02:330, 02:333, 02:337, 02:344, 02:355, 02:358-02:359, 02:373N, 02:376, 02:382, 02:387 and 02:398 and the A*11:06, 11:18, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:78, 68:05, 68:15 and 68:20 alleles.

14Primer mix 14: Specific PCR fragment of 95 bp in A*02:91 and 02:322 alleles. Specific PCR fragment of 170 bp in the A*02:07:01-02:07:02, 02:15N, 02:18, 02:103, 02:112, 02:130, 02:191, 02:219, 02:255, 02:261, 02:265, 02:282, 02:319, 02:335, 02:369 and 02:403 alleles.

15Primer mix 15: Specific PCR fragment of 125 bp in the A*02:21 and 02:186 alleles. Specific PCR fragment of 265 bp in the A*02:87, 02:112, 02:129 and 02:136 alleles. Specific PCR fragment of 305 bp in the A*02:96 allele.

16Primer mix 16: Specific PCR fragment of 110 bp in the A*02:50, 02:73, 02:93, 02:122, 02:156, 02:172 and 02:279 and the A*11:119, 23:31, 23:45, 23:55, 24:15, 24:41, 24:51, 24:92, 26:10, 32:28, 68:02:01:01-68:02:05, 68:15, 68:18N, 68:25, 68:27-68:28, 68:31, 68:34, 68:40, 68:44, 68:48-68:49N, 68:51, 68:53-68:54, 68:60-68:62, 68:64, 68:67, 68:74, 68:77-68:78, 68:80-68:82, 68:86, 68:92 and 68:97 and in the B*44:59:02, 44:136, B*51:136, C*07:204 and C*12:37 alleles. Specific PCR fragment of 155 bp in the A*02:09 and 02:49 alleles.

17Primer mix 17: Specific PCR fragment of 205 bp in the A*02:83N allele. Specific PCR fragment of 360 bp in the A*02:10, 02:17:01-02:17:02, 02:39, 02:108, 02:110, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303 and 02:398 and the A*23:12, 24:28, 24:30, 24:42 and 24:89 alleles.

18Primer mix 19: Specific PCR fragment of 125 bp in the A*02:12-02:13, 02:19, 02:27, 02:37-02:38, 02:44, 02:54, 02:142, 02:271, 02:280, 11:119, 24:02:01:01-24:02:15, 24:02:17-24:02:41, 24:02:43-24:05, 24:07-24:11N, 24:14-24:15, 24:17, 24:19-24:20, 24:23, 24:25-24:53, 24:55-24:64, 24:66-24:86N, 24:88-24:90N, 24:92-24:93, 24:95-24:106, 24:108-24:109, 24:111-24:124, 24:126-24:132N, 24:134-24:137, 24:139-24:157, 24:159, 24:161-24:166, 24:168-24:187, 24:189-24:203, 24:205-24:206, 24:209-24:210, 24:212, 24:214-24:226, 24:229-24:230, 33:19, 68:09 and 68:26 alleles. Specific PCR fragment of 165 bp in the A*02:283 allele.

19Primer mix 20: Specific PCR fragment of 135 bp in the A*02:200 allele. Specific PCR fragment of 295 bp in the A*02:49 allele.

20Primer mix 21: Specific PCR fragment of 220 bp in the A*02:15N and 02:356N alleles. Specific PCR fragment of 255 bp in the A*02:291 allele.

21Primer mix 23: Specific PCR fragment of 135 bp in the A*02:31, 02:161 and 02:387 alleles. Specific PCR fragment of 180 bp in the A*02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300 and 02:303 and the A* 24:94, 24:138, 24:188 and 24:228 alleles.

22Primer mix 24: Specific PCR fragment of 75 bp in the A*02:230 allele. Specific PCR fragment of 115 bp in the A*02:27 and the A*11:119, 24:59, 24:190, 24:210, 24:229 and in the C*07:204 alleles. Specific PCR fragment of 160bp in the A*02:233 and the A*24:219 alleles.

23Primer mix 25: Specific PCR fragment of 125 bp in the A*02:226N allele. Specific PCR fragment of 145 bp in the A*02:16 and 02:131 and the A*24:55 alleles. Specific PCR fragment of 165 bp in the A*02:283 allele.

24Primer mix 26: Specific PCR fragment of 125 bp in the A*02:33 allele. Specific PCR fragment of 165 bp in the A*02:52 allele. Specific PCR fragment of 260 bp in the A*02:198 allele.

25Primer mix 27: Specific PCR fragment of 95 bp in the A*02:28, 02:155 and 02:185 and the A*30:13, 30:16, 30:44 and 30:46 alleles. Specific PCR fragments of 145 bp in A*02:10, 02:50, 02:52, 02:73, 02:93, 02:95, 02:110, 02:114, 02:117, 02:122, 02:156, 02:242, 02:244, 02:279, 02:282, 02:300, 02:304 and 02:339 and the A*03:123, 11:16, 11:35, 11:57, 11:73, 24:131, 24:138, 24:188, 24:218 68:01:01:01-68:11N, 68:13-68:48 and 68:51-68:99 alleles.

26Primer mix 29: Specific PCR fragment of 220 bp in the A*02:45-02:46, 02:48, 02:56:01w-02:56:02w, 02:78w, 02:92, 02:103w, 02:129, 02:169w, 02:195w, 02:358 and 02:369 alleles. Specific PCR fragment of 300 bp in A*02:180 allele.

27Primer mix 30: Specific PCR fragment of 130 bp in the A*02:163 allele. Specific PCR fragment of 160 bp in A*02:43N and 02:104 alleles.

28Primer mix 31: Specific PCR fragment of 95 bp in A*02:82N and in the A*23:08N alleles. Specific PCR fragments of 220 bp in A*02:45, 02:56:01-02:56:02, 02:78, 02:103, 02:169 and 02:195 alleles.

29Primer mix 32: Specific PCR fragment of 205 bp in the A*02:176 allele. Specific PCR fragment of 230 bp in the A*02:46, 02:48, 02:70 and 02:129 alleles. Specific PCR fragment of 255 bp in the A*02:47 allele.

30Primer mix 33: Specific PCR fragment of 230 bp in A*02:165, 02:168 and 02:400 alleles. Specific PCR fragments of 325 bp in A*02:19, 02:36-02:37, 02:54 and 02:255 alleles.

31Primer mix 34: Specific PCR fragment of 120 bp in the A*02:88N allele. Specific PCR fragment of 165 bp in the A*02:395N allele. Specific PCR fragment of 180 bp in the A*02:34-02:35:02, 02:56:01-02:56:02, 02:62, 02:78 and 02:103 and the A*30:01:01-30:01:07, 30:08, 30:11:01-30:11:02, 30:14L-30:20, 30:23-30:24, 30:26, 30:30-30:31, 30:35-30:43, 30:48-30:49, 30:52-30:56, 30:58-30:60, 30:62-30:63, 30:65 and 68:01:11 and in the C*03:82 alleles.

32Primer mix 35: Specific PCR fragment of 110 bp in A*02:40:01-02:40:02, 02:51 and 02:130 and the A*23:01:01-23:01:11, 23:02w, 23:04-23:23, 23:25-23:33, 23:35-23:56, 24:24, 24:71, 31:67-31:68, 32:28, 33:32:01 and 68:51w alleles. Specific PCR fragment of 155 bp in A*02:77 allele.

33Primer mix 36: Specific PCR fragment of 85 bp in A*02:94N allele. Specific PCR fragments of 445 bp in the A*02:24:01-02:24:02, 02:65, 02:135, 02:137, 02:152 and 02:309 and the A*01:104, 03:09, 03:89, 03:108, 11:06, 11:18, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:36, 26:78, 29:19, 29:48, 30:13, 30:16, 30:44, 30:46, 32:08, 33:24, 74:06 and 80:01:01:01w alleles.

34Primer mix 38: Specific PCR fragment of 125 bp in A*02:41, 02:80, 02:117, 02:289 and 02:304 alleles. Specific PCR fragments of 225 bp in the A*02:351 allele.

35Primer mix 39: Specific PCR fragment of 170 bp in the A*02:18 allele. Specific PCR fragment of 260 bp in the A*02:159 and 02:293Q alleles. Specific PCR fragment of 300 bp in the A*02:170 and 02:364 alleles.

36Primer mix 40: Specific PCR fragment of 90 bp in the A*02:67 allele. Specific PCR fragment of 210 bp in the A*02:40:01-02:40:02, 02:51 and 02:130 and the A*29:22 and 33:22 alleles.

37Primer mix 43: Specific PCR fragment of 180 bp in the A*02:71 allele. Specific PCR fragment of 225 bp in the A*02:03:01-02:03:04, 02:22:01-02:22:02, 02:49, 02:104, 02:117, 02:136, 02:148, 02:191, 02:230, 02:253, 02:258, 02:264, 02:267, 02:281, 02:315, 02:323, 02:345, 02:355, 02:370 02:382 and 02:402 and the A*26:22, 66:09 and 68:83 alleles.

38Primer mix 44: Specific PCR fragment of 125 bp in the A*02:203 allele. Specific PCR fragment of 155 bp in the A*02:59 allele. Specific PCR fragment of 205 bp in the A*02:222N allele.

39Primer mix 45: Specific PCR fragment of 105 bp in the A*02:60:01-02:60:02 and 02:254 alleles. Specific PCR fragment of 185 bp in the A*02:19, 02:39, 02:44, 02:79:01-02:79:02, 02:86 and 02:400 and the A*01:20, 01:66, 03:95, 24:14 and 24:93 alleles.

40Primer mix 46: Specific PCR fragment of 70 bp in the A*02:66 allele. Specific PCR fragment of 205 bp in the A*02:61 allele.

41Primer mix 47: Specific PCR fragment of 165 bp in the A*02:63 allele. Specific PCR fragment of 205 bp in the A*02:144 and 02:205 alleles. Specific PCR fragment of 255 bp in the A*02:35:01-02:35:03, 02:48, 02:78, 02:90 and 02:331 alleles. Specific PCR fragment of 305 bp in the A*02:207 allele.

42Primer mix 49: Specific PCR fragment of 115 bp in the A*02:27 and 02:267 and the A*11:119, 24:59, 24:190, 24:210, 24:229 and in the C*07:204 allele. Specific PCR fragment of 145 bp in the A*02:38, 02:101:01-02:101:02 and 02:154 and the A*11:119, 23:10w, 24:10:01-24:10:02, 24:46w, 24:210, 26:10 and 68:61 alleles.

43Primer mix 54: Specific PCR fragment of 240 bp in the A*02:189 and 02:190 alleles. Specific PCR fragment of 350 bp in A*02:74:01-02:74:02 alleles.

44Primer mix 56: Specific PCR fragment of 150 bp in the A*02:175 and 02:181 alleles. Specific PCR fragment of 275 bp in A*02:53N, 02:81 and 02:124 alleles.

45Primer mix 57: Specific PCR fragment of 105 bp in the A*02:42 and 02:310 alleles. Specific PCR fragment of 145 bp in A*02:160 and 02:175 alleles.

46Primer mix 59: Specific PCR fragment of 145 bp in the A*02:97:01-02:97:02 alleles. Specific PCR fragment of 190 bp in A*02:305N allele.

47Primer mix 63: Specific PCR fragment of 95 bp in the A*02:259 and the A*03:51, 30:37 and 32:46 alleles. Specific PCR fragment of 160 bp in A*02:105 and 02:301N alleles.

48Primer mix 64: Specific PCR fragment of 125 bp in the A*02:106 and 02:145 and the A*24:72 alleles. Specific PCR fragment of 175 bp in A*02:164 and 02:221 and the A*01:44, 03:44 23:41 and 29:15 alleles. Specific PCR fragment of 260 bp in A*02:187 allele.

49Primer mix 65: Specific PCR fragment of 145 bp in the A*02:160 and 02:251 alleles. Specific PCR fragment of 185 bp in A*02:107 allele. Specific PCR fragment of 250 bp in A*02:202 allele.

50Primer mix 67: Specific PCR fragment of 105 bp in the A*02:177 allele. Specific PCR fragment of 130 bp in A*02:111 allele. Specific PCR fragment of 180 bp in A*02:330 allele. Specific PCR fragment of 300 bp in A*02:350N allele.

51Primer mix 68: Specific PCR fragment of 210 bp in the A*02:113N and 02:321N alleles. Specific PCR fragment of 260 bp in A*02:184 allele.

52Primer mix 69: Specific PCR fragment of 170 bp in the A*02:114, 02:246 and 02:279 alleles. Specific PCR fragment of 225 bp in A*02:260 allele. Specific PCR fragment of 335 bp in A*02:166 allele.

53Primer mix 70: Specific PCR fragment of 125 bp in the A*02:115 allele. Specific PCR fragment of 180 bp in A*02:367 allele. Specific PCR fragment of 285 bp in A*02:53N, 02:192 and 02:269 alleles.

54Primer mix 72: Specific PCR fragment of 110 bp in the A*02:183 allele. Specific PCR fragment of 230 bp in the A*02:189 allele. Specific PCR fragment of 260 bp in the A*02:19, 02:44, 02:118, 02:135, 02:149, 02:152, 02:190, 02:309 and 02:402 alleles.

55Primer mix 73: Specific PCR fragment of 80 bp in the A*02:263 allele. Specific PCR fragment of 200 bp in the A*02:119 allele. Specific PCR fragment of 250 bp in A*02:158 allele.

56Primer mix 74: Specific PCR fragment of 175 bp in the A*02:120 allele. Specific PCR fragment of 250 bp in A*02:187 and 02:223N alleles.

57Primer mix 75: Specific PCR fragment of 100 bp in the A*02:373N and the A*68:49N alleles. Specific PCR fragment of 165 bp in the A*02:166 allele. Specific PCR fragment of 205 bp in the A*02:121 and the A*23:47, 68:14 and 68:81 and in the B*44:59:02 alleles.

58Primer mix 77: Specific PCR fragment of 85 bp in the A*02:123, 02:295 and 02:344 alleles. Specific PCR fragment of 110 bp in the A*02:162 allele.

59Primer mix 78: Specific PCR fragment of 75 bp in the A*02:193 allele. Specific PCR fragment of 130 bp in the A*02:213 allele. Specific PCR fragment of 240 bp in the A*02:124 allele.

60Primer mix 79: Specific PCR fragment of 215 bp in the A*02:125N and 02:227N alleles. Specific PCR fragment of 270 bp in A*02:53N, 02:184 and 02:192 alleles.

61Primer mix 80: Specific PCR fragment of 90 bp in the A*02:214 allele. Specific PCR fragment of 135 bp in the A*02:194 allele. Specific PCR fragment of 230 bp in A*02:126 allele.

62Primer mix 81: Specific PCR fragment of 110 bp in the A*02:183 allele. Specific PCR fragment of 235 bp in A*02:165 and 02:400 alleles. Specific PCR fragment of 315 bp in A*02:127, 02:167 and 02:299 alleles.

63Primer mix 82: Specific PCR fragment of 130 bp in the A*02:131 allele. Specific PCR fragment of 255 bp in the A*02:291 allele. Specific PCR fragment of 295 bp in the A*02:199 allele.

64Primer mix 83: Specific PCR fragment of 120 bp in the A*02:132 and 02:215 allele. Specific PCR fragment of 240 bp in the A*02:237 and the A*03:95 alleles.

65Primer mix 84: Specific PCR fragment of 180 bp in the A*02:133 and 02:164 and the A*01:44 and 03:44 alleles. Specific PCR fragment of 205 bp in the A*02:315 allele. Specific PCR fragment of 245 bp in the A*02:250N allele.

66Primer mix 85: Specific PCR fragment of 170 bp in the A*02:314N and 02:376 and the A*32:48N allele. Specific PCR fragment of 260 bp in the A*02:134 allele.

67Primer mix 86: Specific PCR fragment of 180 bp in the A*02:135 and 02:309 alleles. Specific PCR fragment of 255 bp in the A*02:174 allele. Specific PCR fragment of 310 bp in the A*02:228 allele.

68Primer mix 87: Specific PCR fragment of 155 bp in the A*02:138 and 02:181 alleles. Specific PCR fragment of 210 bp in the A*02:284N allele.

69Primer mix 88: Specific PCR fragment of 135 bp in the A*02:188 and 02:235 and in the B*15:67, B*35:110, C*12:94 and C*14:48 alleles. Specific PCR fragment of 240 bp in the A*02:139 allele.

70Primer mix 89: Specific PCR fragment of 120 bp in the A*02:140 and 02:182 and the A*24:133, 31:48 and 33:15 alleles. Specific PCR fragment of 200 bp in the A*02:227N allele.

71Primer mix 90: Specific PCR fragment of 100 bp in the A*02:72, 02:206 and 02:275 alleles. Specific PCR fragment of 130 bp in the A*02:161 allele. Specific PCR fragment of 190 bp in the A*02:141 allele.

72Primer mix 92: Specific PCR fragment of 180 bp in the A*02:173 allele. Specific PCR fragment of 260 bp in the A*02:146 allele.

73Primer mix 93: Specific PCR fragment of 130 bp in the A*02:163 allele. Specific PCR fragment of 250 bp in the A*02:147 and 02:339 and the A*23:52, 24:73 and 24:157 and in the B*15:173 and B*18:63 alleles. Specific PCR fragment of 295 bp in the A*02:157 allele.

74Primer mix 94: Specific PCR fragment of 110 bp in the A*02:325 allele. Specific PCR fragment of 210 bp in the A*02:150 allele. Specific PCR fragment of 360 bp in A*02:197 allele.

75Primer mix 96: Specific PCR fragment of 135 bp in the A*02:196 allele. Specific PCR fragment of 255 bp in A*02:153 and 02:366N alleles.

76Primer mix 1 may amplify the A*02:07:01-02:07:02 alleles.

’w’, might be weakly amplified.

1The sequence of A*020116 has been renamed to A*02:134.

The sequence of A*02:01:20 has been shown to be identical to A*02:01:18

The sequence of A*02:01:82 has been shown to be identical to A*02:01:84.

The sequence of the A*02:23 allele has been shown to be identical to A*02:22:01.

The sequence of the A*02:98 allele has been shown to be identical to A*02:96.

The A*02:100 has never been assigned.

2HLA-A*02 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.11.0, January 2013.

3The A*02:97:01-02:97:02 and the 02:305N alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 59.

4The A*02:107, the 02:202 and the 02:251 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 65.

5The A*02:111 and 02:350N alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 67.

6The A*02:119 and 02:263 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 73.

7The A*02:120 and 02:223N alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 74.

8The A*02:132, the 02:215 and the 02:237 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 83.

9The A*02:134 and 02:314N alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 85.

10The A*02:138 and 02:284N alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 87.

11The A*02:139 and 02:235 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 88.

12The A*02:141 and 02:275 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 90.

13The A*02:150, the 02:197 and the 02:325 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 94.

14The A*02:153 and 02:196 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 96.

15The A*02:180 and 02:358 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 29.

16The A*02:193 and 02:213 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 78.

17The A*02:269 and 02:367 alleles may be distinguished by the different sizes of the specific PCR products generated by primer mix 70.

18Primer mix 8: Specific PCR fragment of 415 bp in the A*02:02, 02:05:01-02:05:06, 02:14, 02:47, 02:63, 02:102, 02:115, 02:154-02:155, 02:172, 02:179, 02:186, 02:209, 02:229, 02:232, 02:271, 02:286, 02:320, 02:324, 02:337, 02:344, 02:359, 02:373N and 02:376 alleles. Specific PCR fragment of 505 bp in the A*02:32N allele.

Primer mix 11: Specific PCR fragment of 225 bp in the A*02:89 allele. Specific PCR fragment of 350 bp in the A*02:04, 02:17:01-02:17:02, 02:57, 02:65, 02:108, 02:110, 02:152, 02:268, 02:300, 02:303 and 02:334 and the A*23:12, 24:28, 24:30, 24:42, 24:89, 29:19, 29:48, 32:08, 33:24, 68:05, 68:20 and 74:06 alleles.

Primer mix 12: Specific PCR fragment of 195 bp in A*02:85 allele. Specific PCR fragment of 235 bp in the A*02:05:01-02:06:13, 02:08, 02:10, 02:14, 02:21, 02:28, 02:41, 02:44, 02:51, 02:54, 02:57, 02:61, 02:72, 02:79:01-02:79:02, 02:84, 02:91, 02:99, 02:106, 02:108, 02:122, 02:126-02:127, 02:137, 02:142-02:144, 02:154, 02:169-02:170, 02:172, 02:178-02:180, 02:229, 02:232, 02:244, 02:248, 02:259, 02:271, 02:278, 02:286, 02:290, 02:295, 02:300, 02:310, 02:324, 02:328, 02:330, 02:333, 02:337, 02:344, 02:355, 02:358-02:359, 02:373N, 02:376, 02:382, 02:387 and 02:398 and the A*11:06, 11:18, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:78, 68:05, 68:15 and 68:20 alleles.

Primer mix 14: Specific PCR fragment of 95 bp in A*02:91 and 02:322 alleles. Specific PCR fragment of 170 bp in the A*02:07:01-02:07:02, 02:15N, 02:18, 02:103, 02:112, 02:130, 02:191, 02:219, 02:255, 02:261, 02:265, 02:282, 02:319, 02:335, 02:369 and 02:403 alleles.

Primer mix 15: Specific PCR fragment of 125 bp in the A*02:21 and 02:186 alleles. Specific PCR fragment of 265 bp in the A*02:87, 02:112, 02:129 and 02:136 alleles. Specific PCR fragment of 305 bp in the A*02:96 allele.

Primer mix 16: Specific PCR fragment of 110 bp in the A*02:50, 02:73, 02:93, 02:122, 02:156, 02:172 and 02:279 and the A*11:119, 23:31, 23:45, 23:55, 24:15, 24:41, 24:51, 24:92, 26:10, 32:28, 68:02:01:01-68:02:05, 68:15, 68:18N, 68:25, 68:27-68:28, 68:31, 68:34, 68:40, 68:44, 68:48-68:49N, 68:51, 68:53-68:54, 68:60-68:62, 68:64, 68:67, 68:74, 68:77-68:78, 68:80-68:82, 68:86, 68:92 and 68:97 and in the B*44:59:02, 44:136, B*51:136, C*07:204 and C*12:37 alleles. Specific PCR fragment of 155 bp in the A*02:09 and 02:49 alleles.

Primer mix 17: Specific PCR fragment of 205 bp in the A*02:83N allele. Specific PCR fragment of 360 bp in the A*02:10, 02:17:01-02:17:02, 02:39, 02:108, 02:110, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303 and 02:398 and the A*23:12, 24:28, 24:30, 24:42 and 24:89 alleles.

Primer mix 19: Specific PCR fragment of 125 bp in the A*02:12-02:13, 02:19, 02:27, 02:37-02:38, 02:44, 02:54, 02:142, 02:271, 02:280, 11:119, 24:02:01:01-24:02:15, 24:02:17-24:02:41, 24:02:43-24:05, 24:07-24:11N, 24:14-24:15, 24:17, 24:19-24:20, 24:23, 24:25-24:53, 24:55-24:64, 24:66-24:86N, 24:88-24:90N, 24:92-24:93, 24:95-24:106, 24:108-24:109, 24:111-24:124, 24:126-24:132N, 24:134-24:137, 24:139-24:157, 24:159, 24:161-24:166, 24:168-24:187, 24:189-24:203, 24:205-24:206, 24:209-24:210, 24:212, 24:214-24:226, 24:229-24:230, 33:19, 68:09 and 68:26 alleles. Specific PCR fragment of 165 bp in the A*02:283 allele.

Primer mix 20: Specific PCR fragment of 135 bp in the A*02:200 allele. Specific PCR fragment of 295 bp in the A*02:49 allele.

Primer mix 21: Specific PCR fragment of 220 bp in the A*02:15N and 02:356N alleles. Specific PCR fragment of 255 bp in the A*02:291 allele.

Primer mix 23: Specific PCR fragment of 135 bp in the A*02:31, 02:161 and 02:387 alleles. Specific PCR fragment of 180 bp in the A*02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300 and 02:303 and the A* 24:94, 24:138, 24:188 and 24:228 alleles.

Primer mix 24: Specific PCR fragment of 75 bp in the A*02:230 allele. Specific PCR fragment of 115 bp in the A*02:27 and the A*11:119, 24:59, 24:190, 24:210, 24:229 and in the C*07:204 alleles. Specific PCR fragment of 160bp in the A*02:233 and the A*24:219 alleles.

Primer mix 25: Specific PCR fragment of 125 bp in the A*02:226N allele. Specific PCR fragment of 145 bp in the A*02:16 and 02:131 and the A*24:55 alleles. Specific PCR fragment of 165 bp in the A*02:283 allele.

Primer mix 26: Specific PCR fragment of 125 bp in the A*02:33 allele. Specific PCR fragment of 165 bp in the A*02:52 allele. Specific PCR fragment of 260 bp in the A*02:198 allele.

Primer mix 27: Specific PCR fragment of 95 bp in the A*02:28, 02:155 and 02:185 and the A*30:13, 30:16, 30:44 and 30:46 alleles. Specific PCR fragments of 145 bp in A*02:10, 02:50, 02:52, 02:73, 02:93, 02:95, 02:110, 02:114, 02:117, 02:122, 02:156, 02:242, 02:244, 02:279, 02:282, 02:300, 02:304 and 02:339 and the A*03:123, 11:16, 11:35, 11:57, 11:73, 24:131, 24:138, 24:188, 24:218 68:01:01:01-68:11N, 68:13-68:48 and 68:51-68:99 alleles.

Primer mix 29: Specific PCR fragment of 220 bp in the A*02:45-02:46, 02:48, 02:56:01w-02:56:02w, 02:78w, 02:92, 02:103w, 02:129, 02:169w, 02:195w, 02:358 and 02:369 alleles. Specific PCR fragment of 300 bp in A*02:180 allele.

Primer mix 30: Specific PCR fragment of 130 bp in the A*02:163 allele. Specific PCR fragment of 160 bp in A*02:43N and 02:104 alleles.

Primer mix 31: Specific PCR fragment of 95 bp in A*02:82N and in the A*23:08N alleles. Specific PCR fragments of 220 bp in A*02:45, 02:56:01-02:56:02, 02:78, 02:103, 02:169 and 02:195 alleles.

Primer mix 32: Specific PCR fragment of 205 bp in the A*02:176 allele. Specific PCR fragment of 230 bp in the A*02:46, 02:48, 02:70 and 02:129 alleles. Specific PCR fragment of 255 bp in the A*02:47 allele.

Primer mix 33: Specific PCR fragment of 230 bp in A*02:165, 02:168 and 02:400 alleles. Specific PCR fragments of 325 bp in A*02:19, 02:36-02:37, 02:54 and 02:255 alleles.

Primer mix 34: Specific PCR fragment of 120 bp in the A*02:88N allele. Specific PCR fragment of 165 bp in the A*02:395N allele. Specific PCR fragment of 180 bp in the A*02:34-02:35:02, 02:56:01-02:56:02, 02:62, 02:78 and 02:103 and the A*30:01:01-30:01:07, 30:08, 30:11:01-30:11:02, 30:14L-30:20, 30:23-30:24, 30:26, 30:30-30:31, 30:35-30:43, 30:48-30:49, 30:52-30:56, 30:58-30:60, 30:62-30:63, 30:65 and 68:01:11 and in the C*03:82 alleles.

Primer mix 35: Specific PCR fragment of 110 bp in A*02:40:01-02:40:02, 02:51 and 02:130 and the A*23:01:01-23:01:11, 23:02w, 23:04-23:23, 23:25-23:33, 23:35-23:56, 24:24, 24:71, 31:67-31:68, 32:28, 33:32:01 and 68:51w alleles. Specific PCR fragment of 155 bp in A*02:77 allele.

Primer mix 36: Specific PCR fragment of 85 bp in A*02:94N allele. Specific PCR fragments of 445 bp in the A*02:24:01-02:24:02, 02:65, 02:135, 02:137, 02:152 and 02:309 and the A*01:104, 03:09, 03:89, 03:108, 11:06, 11:18, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 26:36, 26:78, 29:19, 29:48, 30:13, 30:16, 30:44, 30:46, 32:08, 33:24, 74:06 and 80:01:01:01w alleles.

Primer mix 38: Specific PCR fragment of 125 bp in A*02:41, 02:80, 02:117, 02:289 and 02:304 alleles. Specific PCR fragments of 225 bp in the A*02:351 allele.

Primer mix 39: Specific PCR fragment of 170 bp in the A*02:18 allele. Specific PCR fragment of 260 bp in the A*02:159 and 02:293Q alleles. Specific PCR fragment of 300 bp in the A*02:170 and 02:364 alleles.

Primer mix 40: Specific PCR fragment of 90 bp in the A*02:67 allele. Specific PCR fragment of 210 bp in the A*02:40:01-02:40:02, 02:51 and 02:130 and the A*29:22 and 33:22 alleles.

Primer mix 43: Specific PCR fragment of 180 bp in the A*02:71 allele. Specific PCR fragment of 225 bp in the A*02:03:01-02:03:04, 02:22:01-02:22:02, 02:49, 02:104, 02:117, 02:136, 02:148, 02:191, 02:230, 02:253, 02:258, 02:264, 02:267, 02:281, 02:315, 02:323, 02:345, 02:355, 02:370 02:382 and 02:402 and the A*26:22, 66:09 and 68:83 alleles.

Primer mix 44: Specific PCR fragment of 125 bp in the A*02:203 allele. Specific PCR fragment of 155 bp in the A*02:59 allele. Specific PCR fragment of 205 bp in the A*02:222N allele.

Primer mix 45: Specific PCR fragment of 105 bp in the A*02:60:01-02:60:02 and 02:254 alleles. Specific PCR fragment of 185 bp in the A*02:19, 02:39, 02:44, 02:79:01-02:79:02, 02:86 and 02:400 and the A*01:20, 01:66, 03:95, 24:14 and 24:93 alleles.

Primer mix 46: Specific PCR fragment of 70 bp in the A*02:66 allele. Specific PCR fragment of 205 bp in the A*02:61 allele.

Primer mix 47: Specific PCR fragment of 165 bp in the A*02:63 allele. Specific PCR fragment of 205 bp in the A*02:144 and 02:205 alleles. Specific PCR fragment of 255 bp in the A*02:35:01-02:35:03, 02:48, 02:78, 02:90 and 02:331 alleles. Specific PCR fragment of 305 bp in the A*02:207 allele.

Primer mix 49: Specific PCR fragment of 115 bp in the A*02:27 and 02:267 and the A*11:119, 24:59, 24:190, 24:210, 24:229 and in the C*07:204 allele. Specific PCR fragment of 145 bp in the A*02:38, 02:101:01-02:101:02 and 02:154 and the A*11:119, 23:10w, 24:10:01-24:10:02, 24:46w, 24:210, 26:10 and 68:61 alleles.

Primer mix 54: Specific PCR fragment of 240 bp in the A*02:189 and 02:190 alleles. Specific PCR fragment of 350 bp in A*02:74:01-02:74:02 alleles.

Primer mix 56: Specific PCR fragment of 150 bp in the A*02:175 and 02:181 alleles. Specific PCR fragment of 275 bp in A*02:53N, 02:81 and 02:124 alleles.

Primer mix 57: Specific PCR fragment of 105 bp in the A*02:42 and 02:310 alleles. Specific PCR fragment of 145 bp in A*02:160 and 02:175 alleles.

Primer mix 59: Specific PCR fragment of 145 bp in the A*02:97:01-02:97:02 alleles. Specific PCR fragment of 190 bp in A*02:305N allele.

Primer mix 63: Specific PCR fragment of 95 bp in the A*02:259 and the A*03:51, 30:37 and 32:46 alleles. Specific PCR fragment of 160 bp in A*02:105 and 02:301N alleles.

Primer mix 64: Specific PCR fragment of 125 bp in the A*02:106 and 02:145 and the A*24:72 alleles. Specific PCR fragment of 175 bp in A*02:164 and 02:221 and the A*01:44, 03:44 23:41 and 29:15 alleles. Specific PCR fragment of 260 bp in A*02:187 allele.

Primer mix 65: Specific PCR fragment of 145 bp in the A*02:160 and 02:251 alleles. Specific PCR fragment of 185 bp in A*02:107 allele. Specific PCR fragment of 250 bp in A*02:202 allele.

Primer mix 67: Specific PCR fragment of 105 bp in the A*02:177 allele. Specific PCR fragment of 130 bp in A*02:111 allele. Specific PCR fragment of 180 bp in A*02:330 allele. Specific PCR fragment of 300 bp in A*02:350N allele.

Primer mix 68: Specific PCR fragment of 210 bp in the A*02:113N and 02:321N alleles. Specific PCR fragment of 260 bp in A*02:184 allele.

Primer mix 69: Specific PCR fragment of 170 bp in the A*02:114, 02:246 and 02:279 alleles. Specific PCR fragment of 225 bp in A*02:260 allele. Specific PCR fragment of 335 bp in A*02:166 allele.

Primer mix 70: Specific PCR fragment of 125 bp in the A*02:115 allele. Specific PCR fragment of 180 bp in A*02:367 allele. Specific PCR fragment of 285 bp in A*02:53N, 02:192 and 02:269 alleles.

Primer mix 72: Specific PCR fragment of 110 bp in the A*02:183 allele. Specific PCR fragment of 230 bp in the A*02:189 allele. Specific PCR fragment of 260 bp in the A*02:19, 02:44, 02:118, 02:135, 02:149, 02:152, 02:190, 02:309 and 02:402 alleles.

Primer mix 73: Specific PCR fragment of 80 bp in the A*02:263 allele. Specific PCR fragment of 200 bp in the A*02:119 allele. Specific PCR fragment of 250 bp in A*02:158 allele.

Primer mix 74: Specific PCR fragment of 175 bp in the A*02:120 allele. Specific PCR fragment of 250 bp in A*02:187 and 02:223N alleles.

Primer mix 75: Specific PCR fragment of 100 bp in the A*02:373N and the A*68:49N alleles. Specific PCR fragment of 165 bp in the A*02:166 allele. Specific PCR fragment of 205 bp in the A*02:121 and the A*23:47, 68:14 and 68:81 and in the B*44:59:02 alleles.

Primer mix 77: Specific PCR fragment of 85 bp in the A*02:123, 02:295 and 02:344 alleles. Specific PCR fragment of 110 bp in the A*02:162 allele.

Primer mix 78: Specific PCR fragment of 75 bp in the A*02:193 allele. Specific PCR fragment of 130 bp in the A*02:213 allele. Specific PCR fragment of 240 bp in the A*02:124 allele.

Primer mix 79: Specific PCR fragment of 215 bp in the A*02:125N and 02:227N alleles. Specific PCR fragment of 270 bp in A*02:53N, 02:184 and 02:192 alleles.

Primer mix 80: Specific PCR fragment of 90 bp in the A*02:214 allele. Specific PCR fragment of 135 bp in the A*02:194 allele. Specific PCR fragment of 230 bp in A*02:126 allele.

Primer mix 81: Specific PCR fragment of 110 bp in the A*02:183 allele. Specific PCR fragment of 235 bp in A*02:165 and 02:400 alleles. Specific PCR fragment of 315 bp in A*02:127, 02:167 and 02:299 alleles.

Primer mix 82: Specific PCR fragment of 130 bp in the A*02:131 allele. Specific PCR fragment of 255 bp in the A*02:291 allele. Specific PCR fragment of 295 bp in the A*02:199 allele.

Primer mix 83: Specific PCR fragment of 120 bp in the A*02:132 and 02:215 allele. Specific PCR fragment of 240 bp in the A*02:237 and the A*03:95 alleles.

Primer mix 84: Specific PCR fragment of 180 bp in the A*02:133 and 02:164 and the A*01:44 and 03:44 alleles. Specific PCR fragment of 205 bp in the A*02:315 allele. Specific PCR fragment of 245 bp in the A*02:250N allele.

Primer mix 85: Specific PCR fragment of 170 bp in the A*02:314N and 02:376 and the A*32:48N allele. Specific PCR fragment of 260 bp in the A*02:134 allele.

Primer mix 86: Specific PCR fragment of 180 bp in the A*02:135 and 02:309 alleles. Specific PCR fragment of 255 bp in the A*02:174 allele. Specific PCR fragment of 310 bp in the A*02:228 allele.

Primer mix 87: Specific PCR fragment of 155 bp in the A*02:138 and 02:181 alleles. Specific PCR fragment of 210 bp in the A*02:284N allele.

Primer mix 88: Specific PCR fragment of 135 bp in the A*02:188 and 02:235 and in the B*15:67, B*35:110, C*12:94 and C*14:48 alleles. Specific PCR fragment of 240 bp in the A*02:139 allele.

Primer mix 89: Specific PCR fragment of 120 bp in the A*02:140 and 02:182 and the A*24:133, 31:48 and 33:15 alleles. Specific PCR fragment of 200 bp in the A*02:227N allele.

Primer mix 90: Specific PCR fragment of 100 bp in the A*02:72, 02:206 and 02:275 alleles. Specific PCR fragment of 130 bp in the A*02:161 allele. Specific PCR fragment of 190 bp in the A*02:141 allele.

Primer mix 92: Specific PCR fragment of 180 bp in the A*02:173 allele. Specific PCR fragment of 260 bp in the A*02:146 allele.

Primer mix 93: Specific PCR fragment of 130 bp in the A*02:163 allele. Specific PCR fragment of 250 bp in the A*02:147 and 02:339 and the A*23:52, 24:73 and 24:157 and in the B*15:173 and B*18:63 alleles. Specific PCR fragment of 295 bp in the A*02:157 allele.

Primer mix 94: Specific PCR fragment of 110 bp in the A*02:325 allele. Specific PCR fragment of 210 bp in the A*02:150 allele. Specific PCR fragment of 360 bp in A*02:197 allele.

Primer mix 96: Specific PCR fragment of 135 bp in the A*02:196 allele. Specific PCR fragment of 255 bp in A*02:153 and 02:366N alleles.

’w’, might be weakly amplified.

Change in revision R01 compared to R00:

- Primer mix 1 may amplify the A*02:07:01-02:07:02 alleles. A footnote has been added in the Specificity Table.

Change in revision R02 compared to R01:

- Primer mix 1 does amplify the A*02:158 allele. This has been corrected in the Specificity and Interpretation tables.