PHOTO DOCUMENT

Olerup SSP AB

Olerup SSPTM HLA-A*24

Prod. No:	101.422-24/03 – including Taq pol.

101.422-24u/03u – without Taq pol.

Lot No:    	X96

Expiry Date:    	2009-June-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 48 primer mixes used for HLA-A*24 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*24 alleles | Other amplified HLA Class I alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 210 bp | 800 bp | 24020101-240212, 240213weakly, 240214, 240301-2411N, 2413-2415, 2417-2432, 2434-2475, 2477-2479 | 2301-2316, 8001weakly |
| 2 | 245 bp | 1070 bp | 240301-240302, 2410, 2418, 2422, 2433 | 021701-021702, 9208, 9210, 2304, 2907 |
| 3 | 210 bp | 800 bp | 24020101-2405, 2407-2411N, 2417, 2419-2421, 2423, 2425-2450, 2455, 2456, 2458-2463, 2466-2479 |  |
| 45 | 220 bp | 1070 bp | 2406, 2413, 2418, 2422, 2424 | 021701-021702, 9208, 9210, 2301-2313, 2315, 2316, 2907 |
| 5 | 175 bp | 1070 bp | 24020101-2404, 2406-2411N, 2413, 2417-2423, 2425-2450, 2454-2456, 2458-2463, 2465-2479 | 021701-021702weakly, 2314 |
| 65 | 175 bp | 1070 bp | 2405, 2424 | 2301-2313, 2315, 2316, 2907 |
| 7 | 330 bp | 1070 bp | 24020101-240302, 2405-2411N, 2413, 2417, 2418, 2420- 2427, 2429-2443, 2445N-2450, 2452, 2454-2456, 2458-2479 | 2301-2316 |
| 84 | 75 bp | 1070 bp | 2404 |  |
| 9 | 225 bp | 800 bp | 2407, 2419, 2424 | 2313, 300101-300102, 3011, 3014L-3021 |
| 10 | 365 bp | 1070 bp | 24020101-2406, 2408-2411N, 2413, 2417, 2418, 2420-2423, 2425-2450, 2452, 2454-2456, 2458-2479 | 0210, 021701-021702, 0239, 9208, 9210, 2301-2312, 2314-2316 |
| 115,8 | 200, 245 bp | 1070 bp | 2408, 2442, 2461 | 0276, 3022 |
| 12 | 235 bp | 800 bp | 2425 | 2305 |
| 13 | 160 bp | 1070 bp | 24020101-240214, 2404-2409N, 2411N, 2413-2415, 2417, 2419, 2420, 2423-2432, 2434-2474, 2476-2479 | 0219, 0236, 0237, 0254, 2301, 2302, 2305-2316 6826 |
| 144 | 105 bp | 1070 bp | 2409N |  |
| 15 | 385 bp | 800 bp | 2414, 2415, 2451, 2453, 2457 | 0246, 0248, 0270, 0330, 3007, 3108, 3205 |
| 164,6,9 | 85, 245 bp | 1070 bp | 2417, 2423, 2441 | 2907 |
| 17 | 230 bp | 1070 bp | 2410, 2446 | 2310 |
| 18 | 270 bp | 1070 bp | 24020102L |  |
| 197 | 195 bp | 1070 bp | 24020101-240209, 240211-2411N, 2413-2415, 2417-2432, 2434-2479 | 2301-2316 |
| 206,7 | 470 bp | 1070 bp | 2411N | 0104N, 0321N, 1121N, 2307N |
| 214,10 | 125, 155 bp | 800 bp | 2431, 2472 | 9206 |
| 224,11 | 125, 270 bp | 800 bp | 2427, 2463 |  |
| 234,12 | 125, 190 bp | 800 bp | 2428, 2430, 2442, 2472 | 020109, 0250, 9206, 9222, 0309, 1106, 1118, 2312, 2603, 2606, 2621, 2630, 2633, 3208, 6805, 6815, 6820, 7406 |
| 244,13 | 135, 240 bp | 800 bp | 2419, 2428, 2444, 2460N, 2473 | 0276, 300101-300102, 3011, 3013-3021 |
| 25 | 220 bp | 1070 bp | 2413, 2418, 2424, 2456, 2465 | 021701-021702, 9208, 9210, 2301, 2303-2316, 2907 |
| 2614 | 235, 285 bp | 800 bp | 2422, 2434, 2478 |  |
| 274, 7,15 | 80, 370 bp | 800 bp | 2414, 2432 | 0246, 0248, 0270, 3205 |
| 284,16 | 145, 200 bp | 800 bp | 2435, 2473, 2477 |  |
| 2917 | 210, 240 bp | 800 bp | 2408, 2429, 2440N, 2442 | 3022 |
| 304,18 | 115, 200 bp | 800 bp | 2421, 2459, 2470 | 2303, 2907 |
| 3119 | 160, 200 bp | 800 bp | 2437, 2458, 2477 |  |
| 32 | 220 bp | 1070 bp | 2415, 2438, 2441, 2451, 2479 | 01010101-0104N, 0106-0127N, 1127, 2629, 3213, B*1512weakly, B*1519 weakly |
| 334 | 100 bp | 1070 bp | 2439, 2470 |  |
| 344,20 | 80, 195 bp | 800 bp | 2436N, 2475 |  |
| 3521 | 200, 240 bp | 800 bp | 2440N, 2445N |  |
| 364,7,22 | 115, 170 bp | 1070 bp | 2447, 2474 |  |
| 3723 | 210, 255 bp | 1070 bp | 2448N, 2454, 2468 |  |
| 384,24 | 80, 145 bp | 800 bp | 2449, 2475 |  |
| 3925 | 210, 255 bp | 800 bp | 2450, 2468 |  |
| 4026 | 225, 270 bp | 1070 bp | 2408, 2420, 2463, 2469 |  |
| 41 | 205 bp | 1070 bp | 2426, 2462 | 2629, 3105, 3213, 3310 |
| 4227 | 165, 200 bp | 1070 bp | 2424, 2458, 2467 | 2616 |
| 4328 | 160, 275 bp | 800 bp | 2443, 2469 |  |
| 4429 | 195, 250 bp | 800 bp | 2451, 2457, 2466 | 01010101-0104N, 0106-0119, 0121-0127N, 0318, 1127, 2629, 2903, 3105, 3213, 3310, 6826, 8001 |
| 457 | 240 bp | 1070 bp | 2414, 2452, 2479 | 0120, 0219, 0236, 0237, 0254, 2629, B*1512weakly, B*1519weakly |
| 464,5,7,30 | 120, 170 bp | 1070 bp | 2415, 2441, 2451, 2464, 2474 | 0250, 0273, 0293, 9222, 2610, 68020101-68020103, 6815, 6818N, 6825, 6827, 6828, 6831, 6834, 6840 |
| 4731 | 240, 285, 410 bp | 800 bp | 2423, 2426, 2455, 2478 | 0330, 3007, 3108, 3205 |
| 48 | 195 bp | 1070 bp | 2424, 2471weakly | 2301, 2302weakly, 2303-2313, 2314weakly, 2315, 2316, 2907 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*24 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*24 subtyping.

In addition, tubes number 3, 9, 12, 15, 21, 22, 23, 24, 26, 27, 28, 29, 30, 31, 34, 35, 38, 39, 43, 44 and 45 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-A alleles some non-HLA-A*24 alleles will be amplified by primer mixes 1, 2, 4 to 7, 9 to 13, 15 to 17, 19 to 21, 23 to 25, 27, 29, 30, 32, 41, 42, 44 to 48. In addition, the B*1512 and B*1519 alleles will be weakly amplified by primer mixes 32 and 45.

We assume that unknown 4th exon A*23 and A*24 nucleotide sequences are conserved within the A*23/A*24 allele group.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mixes 4, 6 and 46 may yield specific PCR products than the other A*24 primer mixes.

6Primer mix 20 may give rise to a primer oligomer artefact.

7Primer mixes 20, 27, 36, 45 and 46 may give rise to nonspecific amplifications.

8Primer mix 11: Specific PCR fragment of 200 bp in the A*2408, 2442 and the A*0276 and 3002 alleles. Specific PCR fragment of 245 bp in the A*2461 allele.

9Primer mix 16: Specific PCR fragment of 85 bp in the A*2417, 2441 and the A*2907 alleles. Specific PCR fragment of 245 bp in the A*2423 allele.

10Primer mix 21: Specific PCR fragment of 125 bp in the A*2472 and A*9206 alleles. Specific PCR fragment of 155 bp in the A*2431 allele.

11Primer mix 22: Specific PCR fragment of 125 bp in the A*2427 allele. Specific PCR fragment of 270 bp in the A*2463 allele.

12Primer mix 23: Specific PCR fragment of 125 bp in the A*2472 and A*9206 allele. Specific PCR fragment of 190 bp in the A*2428, 2430, 2442 and the A*020109, 0250, 9222, 0309, 1106, 1118, 2312, 2603, 2606, 2621, 2630, 2633, 3208, 6805, 6815, 6820, 7406 alleles.

13Primer mix 24: Specific PCR fragment of 135 bp in the A*2473 allele. Specific PCR fragment of 240 bp in the A*2419, 2428, 2444, 2460N and the A*0276, 300101-300102, 3011, 3013 to 3021 alleles.

14Primer mix 26: Specific PCR fragment of 235 bp in the A*2422 and 2434 alleles. Specific PCR fragment of 285 bp in the A*2478 allele.

15Primer mix 27: Specific PCR fragment of  80 bp in the A*2432 and A*3205 alleles. Specific PCR fragment of  370 bp in the A*2414, A*0246, A*0248 and A*0270 alleles.

16Primer mix 28:  Specific PCR fragment of 145 bp in the A*2435 allele. Specific PCR fragment of 200 bp in the A*2473 and 2477 alleles.

17Primer mix 29: Specific PCR fragment of 210 bp in the A*2408, 2429,  2442  and the A*3022 alleles. Specific PCR fragment of 240 bp in the A*2440N allele.

18Primer mix 30: Specific PCR fragment of 115 bp in the A*2421, 2470 and the A*2303 and 2907 alleles. Specific PCR fragment of 200 bp in the A*2459 allele.

19Primer 31: Specific PCR fragment of 160 bp in the A*2437 allele. Specific PCR fragment of 200 bp in the A*2458 and 2477 alleles.

20Primer mix 34: Specific PCR fragment of 80 bp in the A*2475 allele. Specific PCR fragment of 195 bp in the A*2436 N allele.

21Primermix 35: Specific PCR fragment of 200 bp in the A*2445N allele. Specific PCR fragment of 240 bp in the A*2440N allele.

22Primer mix 36: Specific PCR fragment of 115 bp in the A*2447 allele. Specific PCR fragment of 170 bp in the A*2474 allele.

23Primer mix 37: Specific PCR fragment of 210 bp in the A*2448N and 2454 alleles. Specific PCR fragment of 255 bp in the A*2468 allele.

24Primer mix 38: Specific PCR fragment of 80 bp in the A*2475 allele. Specific PCR fragment of 145 bp in the A*2449 allele.

25Primer mix 39: Specific PCR fragment of 210 bp in the A*2450 allele. Specific PCR fragment of 255 bp in the A*2468 allele.

26Primer mix 40: Specific PCR fragment of 225 bp in the A*2408 and 2420 alleles. Specific PCR fragment of 270 bp in the A*2463 and 2469 alleles.

27Primer mix 42: Specific PCR fragment of 165 bp in the A*2458 allele. Specific PCR fragment of 200 bp in the A*2424, 2467 and the A*2616 alleles.

28Primer mix 43: Specific PCR fragment o 160 bp in the A*2443 allele. Specific PCR fragment of 275 bp in the A*2469 allele.

29Primer mix 44: Specific PCR fragment of 195 bp in the A*2466 allele. Specific PCR fragment of 250 bp in the A*2451, 2457 and the A*01010101-0104N, 0106-0119, 0121-0127N, 0318, 1127, 2629, 2903, 3105, 3213, 3310, 6826 and 8001 alleles.

30Primer mix 46: Specific PCR fragment of 120 bp in the A*2415, 2441, 2451, 2464 and the A*0250, 0273, 0293, 9222, 2610, 68020101-68020103, 6815, 6818N, 6825, 6827, 6828, 6831, 6834 and 6840 alleles. Specific PCR fragment of 170 bp in the A*2474 allele.

31Primer mix 47: Specific PCR fragment of 240 bp in the A*2423 and 2455 alleles. Specific PCR fragment of 285 bp in the A*2478 allele. Specific PCR fragment of 410 bp in the A*2426 and the A*0330, 3007, 3108 and 3205 alleles.

1The sequence of the A*2401 has been shown to be in error.

The sequence of the A*2412 allele has been shown to be identical to A*2408.

The A*2416 allele has been renamed to A*3108.

2Primer mix 11: Specific PCR fragment of 200 bp in the A*2408, 2442 and the A*0276 and 3002 alleles. Specific PCR fragment of 245 bp in the A*2461 allele.

Primer mix 16: Specific PCR fragment of 85 bp in the A*2417, 2441 and the A*2907 alleles. Specific PCR fragment of 245 bp in the A*2423 allele.

Primer mix 21: Specific PCR fragment of 125 bp in the A*2472 and A*9206 alleles. Specific PCR fragment of 155 bp in the A*2431 allele.

Primer mix 22: Specific PCR fragment of 125 bp in the A*2427 allele. Specific PCR fragment of 270 bp in the A*2463 allele.

Primer mix 23: Specific PCR fragment of 125 bp in the A*2472 and A*9206 allele. Specific PCR fragment of 190 bp in the A*2428, 2430, 2442 and the A*020109, 0250, 9222, 0309, 1106, 1118, 2312, 2603, 2606, 2621, 2630, 2633, 3208, 6805, 6815, 6820, 7406 alleles.

Primer mix 24: Specific PCR fragment of 135 bp in the A*2473 allele. Specific PCR fragment of 240 bp in the A*2419, 2428, 2444, 2460N and the A*0276, 300101-300102, 3011, 3013 to 3021 alleles.

Primer mix 26: Specific PCR fragment of 235 bp in the A*2422 and 2434 alleles. Specific PCR fragment of 285 bp in the A*2478 allele.

Primer mix 27: Specific PCR fragment of  80 bp in the A*2432 and A*3205 alleles. Specific PCR fragment of  370 bp in the A*2414, A*0246, A*0248 and A*0270 alleles.

Primer mix 28:  Specific PCR fragment of 145 bp in the A*2435 allele. Specific PCR fragment of 200 bp in the A*2473 and 2477 alleles.

Primer mix 29: Specific PCR fragment of 210 bp in the A*2408, 2429,  2442  and the A*3022 alleles. Specific PCR fragment of 240 bp in the A*2440N allele.

Primer mix 30: Specific PCR fragment of 115 bp in the A*2421, 2470 and the A*2303 and 2907 alleles. Specific PCR fragment of 200 bp in the A*2459 allele.

Primer 31: Specific PCR fragment of 160 bp in the A*2437 allele. Specific PCR fragment of 200 bp in the A*2458 and 2477 alleles.

Primer mix 34: Specific PCR fragment of 80 bp in the A*2475 allele. Specific PCR fragment of 195 bp in the A*2436 N allele.

Primermix 35: Specific PCR fragment of 200 bp in the A*2445N allele. Specific PCR fragment of 240 bp in the A*2440N allele.

Primer mix 36: Specific PCR fragment of 115 bp in the A*2447 allele. Specific PCR fragment of 170 bp in the A*2474 allele.

Primer mix 37: Specific PCR fragment of 210 bp in the A*2448N and 2454 alleles. Specific PCR fragment of 255 bp in the A*2468 allele.

Primer mix 38: Specific PCR fragment of 80 bp in the A*2475 allele. Specific PCR fragment of 145 bp in the A*2449 allele.

Primer mix 39: Specific PCR fragment of 210 bp in the A*2450 allele. Specific PCR fragment of 255 bp in the A*2468 allele.

Primer mix 40: Specific PCR fragment of 225 bp in the A*2408 and 2420 alleles. Specific PCR fragment of 270 bp in the A*2463 and 2469 alleles.

Primer mix 42: Specific PCR fragment of 165 bp in the A*2458 allele. Specific PCR fragment of 200 bp in the A*2424, 2467 and the A*2616 alleles.

Primer mix 43: Specific PCR fragment o 160 bp in the A*2443 allele. Specific PCR fragment of 275 bp in the A*2469 allele.

Primer mix 44: Specific PCR fragment of 195 bp in the A*2466 allele. Specific PCR fragment of 250 bp in the A*2451, 2457 and the A*01010101-0104N, 0106-0119, 0121-0127N, 0318, 1127, 2629, 2903, 3105, 3213, 3310, 6826 and 8001 alleles.

Primer mix 46: Specific PCR fragment of 120 bp in the A*2415, 2441, 2451, 2464 and the A*0250, 0273, 0293, 9222, 2610, 68020101-68020103, 6815, 6818N, 6825, 6827, 6828, 6831, 6834 and 6840 alleles. Specific PCR fragment of 170 bp in the A*2474 allele.

Primer mix 47: Specific PCR fragment of 240 bp in the A*2423 and 2455 alleles. Specific PCR fragment of 285 bp in the A*2478 allele. Specific PCR fragment of 410 bp in the A*2426 and the A*0330, 3007, 3108 and 3205 alleles.

‘w’, might be weakly amplified.