PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*24

Prod. No:	101.422-24/24u

101.422-03/03u

Lot No:    	46R

Expiry Date:    	2015-July-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 93 primer mixes used for HLA-A*24 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*24 alleles3 | Other amplified HLA Class I alleles4 |
| --- | --- | --- | --- | --- |
| 18 | 210 bp | 800 bp | *24:02:01:01-24:02:12, 24:02:13w, 24:02:14-24:11N, 24:13:01-24:15, 24:17-24:32, 24:34-24:64, 24:66-24:75, 24:77-24:183N, 24:185N-24:209 | *23:01:01-23:54, 33:19, 80:01w-80:02w |
| 2 | 245 bp | 1070 bp | *24:03:01-24:03:02, 24:10, 24:18, 24:22, 24:33, 24:94, 24:125, 24:138, 24:167, 24:204, 24:207 | *02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300, 02:303, 23:04, 29:07, 31:29, B*18:27 |
| 3 | 210 bp | 800 bp | *24:02:01:01-24:05, 24:07-24:11N, 24:17, 24:19-24:21:02, 24:23, 24:25-24:50, 24:55-24:56, 24:58-24:63, 24:66-24:86N, 24:88-24:90N, 24:93, 24:95-24:106, 24:108-24:113, 24:115-24:132N, 24:134-24:137, 24:139-24:166, 24:168-24:187, 24:189-24:206, 24:208-24:209 | *33:19 |
| 46 | 220 bp | 1070 bp | *24:06, 24:13:01, 24:18, 24:22, 24:24, 24:94, 24:138, 24:167, 24:188, 24:204, 24:207 | *02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300, 02:303, 23:01:01-23:13, 23:14:02-23:33, 23:35-23:54, 29:07, 31:29 |
| 5 | 175 bp | 1070 bp | *24:02:01:01-24:04, 24:06-24:11N, 24:13:01-24:13:02, 24:17-24:23, 24:25-24:50, 24:54-24:56, 24:58-24:63, 24:66-24:91, 24:93, 24:95-24:113, 24:115-24:129, 24:131-24:137, 24:139-24:167, 24:169-24:187, 24:189-24:193, 24:195-24:198, 24:200-24:209 | *02:17:01w-02:17:02w, 23:14:01-23:14:02, 33:19 |
| 66 | 175 bp | 1070 bp | *24:05, 24:24, 24:199 | *23:01:01-23:13, 23:15-23:54, 29:07, 31:29, B*18:27 |
| 7 | 335 bp | 1070 bp | *24:02:01:01-24:03:02, 24:05-24:11N, 24:13:01-24:13:02, 24:17-24:18, 24:20-24:27, 24:29-24:43, 24:45N-24:50, 24:52, 24:54-24:56, 24:58-24:64, 24:66-24:88, 24:90N-24:91, 24:93-24:108, 24:110-24:117, 24:119-24:128, 24:130-24:163N, 24:165-24:209 | *23:01:01-23:26, 23:28-23:54 |
| 85 | 75 bp | 1070 bp | *24:04, 24:109 |  |
| 9 | 225 bp | 800 bp | *24:07, 24:19, 24:24, 24:112, 24:131 | *03:72, 11:88, 23:13, 29:37, 30:01:01-30:01:07, 30:11:01-30:11:02, 30:14L-30:20, 30:23-30:26, 30:30-30:31, 30:35-30:43, 30:48-30:49, 30:52-30:54, 30:56, 30:58-30:60, 30:62-30:63, 30:65, 68:45 |
| 10 | 370 bp | 1070 bp | *24:02:01:01-24:02:04, 24:02:06-24:06, 24:08-24:11N, 24:13:01-24:13:02, 24:17-24:18, 24:20-24:23, 24:25-24:50, 24:52, 24:54-24:56, 24:58-24:64, 24:66-24:91, 24:93-24:107, 24:109-24:111, 24:113-24:117, 24:119-24:130, 24:132N-24:209 | *02:10, 02:17:01-02:17:02, 02:39, 02:108, 02:110, 02:148, 02:242, 02:244, 02:268, 02:300, 02:303, 23:01:01-23:12, 23:14:01-23:21, 23:23-23:54, 31:29, 33:19 |
| 118,9 | 200 bp, 245 bp | 1070 bp | *24:08, 24:42, 24:61, 24:89 | *02:76:01-02:76:02, 02:319, 23:38N, 30:22 |
| 12 | 235 bp | 1070 bp | *24:25 | *23:05 |
| 13 | 160 bp | 1070 bp | *24:02:01:01-24:02:59, 24:04-24:09N, 24:11N, 24:13:01-24:15, 24:17, 24:19-24:20, 24:23-24:32, 24:34-24:64, 24:66-24:74, 24:76-24:93, 24:95-24:109, 24:111-24:124, 24:126-24:137, 24:139-24:157, 24:159-24:166, 24:168-24:203, 24:205-24:206, 24:209 | *02:19, 02:36-02:37, 02:54, 02:255, 23:01:01-23:02, 23:05-23:42, 23:44-23:54, 33:19, 68:26, 68:65 |
| 145 | 105 bp | 1070 bp | *24:09N |  |
| 155,10 | 105 bp, 385 bp | 800 bp | *24:14-24:15, 24:19, 24:44, 24:51, 24:53, 24:57, 24:92, 24:100, 24:109 | *02:46, 02:48, 02:70, 02:129, 03:30, 30:07, 31:08, 32:05, 33:21, 33:53 |
| 165 | 85 bp | 1070 bp | *24:17, 24:41, 24:208 | *23:53, 29:07 |
| 17 | 230 bp | 1070 bp | *24:10, 24:46 | *23:10 |
| 18 | 275 bp | 1070 bp | *24:02:01:02L |  |
| 198 | 195 bp | 1070 bp | *24:02:01:01-24:02:09, 24:02:11-24:11N, 24:13:01-24:15, 24:17-24:32, 24:34-24:64, 24:66-24:209 | *23:01:01-23:54, 33:19 |
| 207,8 | 470 bp | 1070 bp | *24:11N | *01:04N, 03:21N, 11:21N, 23:07N |
| 21 | 155 bp | 800 bp | *24:31 |  |
| 225 | 125 bp | 800 bp | *24:27 |  |
| 23 | 190 bp | 800 bp | *24:28, 24:30, 24:42, 24:89 | *02:01:09, 02:05:05, 02:06:07, 02:50, 02:76:02, 02:122, 03:09, 03:108, 11:06, 11:18, 23:12, 25:11, 26:03:01-26:03:02, 26:06, 26:21, 26:30, 29:19, 32:08, 33:24, 68:05, 68:15, 68:20, 74:06 |
| 2411 | 215 bp, 240 bp | 800 bp | *24:19, 24:24, 24:28, 24:44, 24:67, 24:89, 24:145, 24:156, 24:191 | *02:76:01-02:76:02, 02:319, 03:72, 11:88, 23:51, 26:16, 29:37, 30:01:01-30:04:02, 30:06, 30:09-30:20, 30:23-30:54, 30:56-30:66, 32:07, 68:45 |
| 25 | 220 bp | 1070 bp | *24:13:01-24:13:02, 24:18, 24:24, 24:56, 24:94, 24:177, 24:188, 24:204, 24:207 | *02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300, 02:303, 23:01:01-23:01:10, 23:03:01-23:54, 29:07, 31:29, B*18:27 |
| 265,12 | 90 bp, 220 bp | 800 bp | *24:22, 24:96, 24:160 |  |
| 275,8,13 | 80 bp, 370 bp | 800 bp | *24:14, 24:32 | *02:46, 02:48, 02:70, 02:129, 23:36, 32:05 |
| 285,8,14 | 75 bp, 155 bp | 800 bp | *24:35, 24:155N |  |
| 295,15 | 75 bp, 210 bp | 800 bp | *24:08, 24:29, 24:42, 24:89, 24:155N | *02:76:01-02:76:02, 02:319, 30:22 |
| 305,7,16 | 125 bp, 200 bp | 800 bp | *24:21:01-24:21:02, 24:59, 24:190, 24:208 | *23:03:01-23:03:02, 29:07, 31:29, B*18:27 |
| 3117 | 155 bp, 235 bp | 800 bp | *24:04, 24:37, 24:109, 24:119, 24:129 | *01:02, 01:20, 26:16, 29:37, 30:57 |
| 32 | 225 bp | 1070 bp | *24:15, 24:41, 24:51, 24:92 | *01:01:01:01-01:04N, 01:06-01:33, 01:35-01:37, 01:39-01:41, 01:43-01:71, 01:73-01:101, 01:103-01:113, 11:27, 11:38-11:39, 11:94, 23:31, 23:45, 26:29, 26:49, 32:13, 66:10, B*07:64w, B*15:12w, B*15:19w |
| 335 | 100 bp | 1070 bp | *24:39, 24:139 |  |
| 34 | 195 bp | 800 bp | *24:36N |  |
| 35 | 195 bp | 1070 bp | *24:45N, 24:88, 24:142 |  |
| 365,8 | 115 bp | 1070 bp | *24:47, 24:124 | *31:07, 31:10 |
| 37 | 210 bp | 1070 bp | *24:48N, 24:54 |  |
| 3818 | 145 bp, 170 bp, 230 bp | 800 bp | *24:49, 24:126, 24:137 |  |
| 395,19 | 75 bp, 210 bp | 800 bp | *24:50, 24:185N | *02:129 |
| 40 | 225 bp | 1070 bp | *24:08, 24:20, 24:88 |  |
| 417,20 | 200 bp, 240 bp | 1070 bp | *24:26, 24:62, 24:118, 24:132N | *01:39, 23:45, 26:29, 26:49, 31:05, 32:13, 33:10, 66:10 |
| 4221 | 170 bp, 200 bp | 1070 bp | *24:22, 24:24, 24:67, 24:160 | *02:262, 26:16, 68:45 |
| 43 | 160 bp | 800 bp | *24:43, 24:159 |  |
| 44 | 250 bp | 800 bp | *24:51, 24:57, 24:92, 24:120 | *01:01:01:01-01:04N, 01:06-01:19, 01:21-01:33, 01:35-01:65, 01:67-01:71, 01:74-01:101, 01:103-01:113, 03:18, 03:97, 03:122, 03:135, 11:27, 11:38-11:39, 11:94, 26:29, 26:49, 29:03, 29:33, 31:05, 32:13, 33:10, 66:10, 68:26, 68:65, 80:01-80:02 |
| 4522 | 200 bp, 245 bp | 1070 bp | *24:14, 24:52, 24:113, 24:132N | *01:20, 01:66, 01:73, 02:19, 02:36-02:37, 02:54, 02:255, 26:29, 26:49, 66:10, B*07:64, B*15:12, B*15:19 |
| 465 | 110 bp | 1070 bp | *24:15, 24:41, 24:51, 24:92 | *02:50, 02:73, 02:93, 02:122, 02:156, 02:172, 02:279, 11:119, 23:31, 23:45, 26:10, 32:28, 68:02:01:01-68:02:05, 68:15, 68:18N, 68:25, 68:27-68:28, 68:31, 68:34, 68:40, 68:44, 68:48-68:49N, 68:51, 68:53-68:54, 68:60-68:62, 68:64, 68:67, 68:74, 68:77-68:78, 68:80-68:82, 68:86, 68:92, B*44:136, B*51:136, C*07:204, C*12:37 |
| 4723 | 200 bp, 410 bp | 800 bp | *24:18, 24:26, 24:56, 24:123, 24:177 | *03:30, 30:07, 31:08, 32:05, 33:53 |
| 486 | 195 bp | 1070 bp | *24:24, 24:71 | *23:01:01-23:23, 23:25-23:33, 23:35-23:54, 29:07, 31:29, B*18:27 |
| 4924 | 165 bp, 215 bp, 245 bp | 1070 bp | *24:23, 24:91, 24:159 |  |
| 5025 | 235 bp, 260 bp | 1070 bp | *24:34, 24:163N |  |
| 5126 | 165 bp, 215 bp | 800 bp | *24:38, 24:140 |  |
| 528 | 240 bp | 800 bp | *24:40N |  |
| 535,27 | 125 bp, 230 bp | 1070 bp | *24:55, 24:158N |  |
| 548 | 165 bp | 1070 bp | *24:58 |  |
| 5528 | 240 bp, 275 bp | 1070 bp | *24:60N, 24:173 |  |
| 56 | 270 bp | 1070 bp | *24:63, 24:163N | *68:92 |
| 57 | 135 bp | 1070 bp | *24:64 | C*12:37 |
| 58 | 195 bp | 1070 bp | *24:66 | *23:46 |
| 5929 | 230 bp, 255 bp, 285 bp | 1070 bp | *24:68, 24:127, 24:143 |  |
| 6030 | 205 bp, 275 bp | 800 bp | *24:69, 24:183N |  |
| 615,31 | 105 bp, 280 bp | 800 bp | *24:70, 24:127, 24:141 |  |
| 625 | 125 bp | 800 bp | *24:72 | *02:106, 02:145 |
| 63 | 135 bp | 1070 bp | *24:73, 24:157 | B*18:38, B*40:197 |
| 645,8,32 | 115 bp, 170 bp | 1070 bp | *24:74, 24:124 | *31:07, 31:10 |
| 655,33 | 80 bp, 235 bp | 1070 bp | *24:75, 24:90N | *02:88N |
| 66 | 155 bp | 800 bp | *24:77 | *11:106, 23:28, 32:05 |
| 67 | 230 bp | 1070 bp | *24:79 |  |
| 6834 | 230 bp, 280 bp | 800 bp | *24:78, 24:141, 24:143 |  |
| 695,35 | 105 bp, 265 bp | 1070 bp | *24:80, 24:184 |  |
| 705,36 | 110 bp, 250 bp | 1070 bp | *24:14w, 24:81, 24:93w, 24:134 | *02:19w, 02:36-02:37w, 02:54w, 02:255w |
| 7137 | 130 bp, 205 bp, 260 bp | 800 bp | *24:82, 24:128, 24:183N | *23:29, 29:13, 31:07-31:08, 31:10, 32:42 |
| 725 | 120 bp | 1070 bp | *24:83N |  |
| 735,38 | 100 bp, 170 bp | 1070 bp | *24:84N, 24:131, 24:138, 24:188 | *02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300, 02:303 |
| 7439 | 145 bp, 200 bp | 1070 bp | *24:85, 24:180 |  |
| 757,40 | 145 bp, 290 bp | 1070 bp | *24:86N, 24:180 |  |
| 768,41 | 170 bp, 260 bp | 1070 bp | *24:76, 24:98 |  |
| 7742 | 145 bp, 210 bp | 800 bp | *24:06, 24:87, 24:122, 24:138, 24:167 | *23:02 |
| 78 | 140 bp | 800 bp | *24:57, 24:92, 24:120 | *02:50, 02:73, 02:122, 02:156, 02:245, 02:334, 02:338, 03:17:01, 11:119, 26:10, 30:61, 32:28, 33:32:01, 68:01:01:01-68:43, 68:45-68:55, 68:57-68:82, 68:84-68:90, 68:92-68:94N |
| 7943 | 195 bp, 250bp | 1070 bp | *24:14, 24:93, 24:114, 24:182 | *01:20, 01:66, 02:19, 02:36-02:37, 02:54, 02:255 |
| 8044 | 170 bp, 220 bp | 1070 bp | *24:95, 24:140, 24:142 |  |
| 815,45 | 100 bp, 215 bp | 1070 bp | *24:97, 24:133, 24:139 |  |
| 8246 | 185 bp, 225 bp | 1070 bp | *24:99, 24:125, 24:130 |  |
| 8347 | 135 bp, 275 bp | 800 bp | *24:101, 24:111, 24:173 |  |
| 845,48 | 95 bp, 210 bp, 290 bp | 1070 bp | *24:102-24:103, 24:107 |  |
| 8549 | 145 bp, 210 bp | 1070 bp | *24:104, 24:117 | *03:30, 32:05 |
| 8650 | 190 bp, 255 bp | 1070 bp | *24:105, 24:121, 24:182 |  |
| 875,51 | 90 bp, 120 bp, 195 bp | 1070 bp | *24:17, 24:41, 24:62, 24:106, 24:110, 24:136, 24:208 | *23:31, 23:45, 23:53, 29:07, 31:29 |
| 885,52 | 75 bp, 170 bp | 1070 bp | *24:116, 24:137, 24:185N |  |
| 895,53 | 75 bp, 270 bp | 1070 bp | *24:07, 24:108, 24:112, 24:131 | *23:13, B*51:129 |
| 905,54 | 95 bp, 180 bp | 800 bp | *24:115, 24:168 |  |
| 918 | 215 bp | 1070 bp | *24:135 | *23:31, 23:45, 23:52, 32:13, B*07:64, B*15:12, B*15:19 |
| 92 | 190 bp | 1070 bp | *24:144 |  |
| 938 | 270 bp | 1070 bp | *24:112 | B*51:129 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*24 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*24 subtyping.

In addition, wells number 3, 9, 15, 21 to 24, 26 to 31, 34, 38, 39, 43, 44, 47, 51, 52, 60 to 62, 66, 68, 71, 77, 78, 83 and 90 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several HLA Class I alleles 1st and/or 4th exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups.

4Due to the sharing of sequence motifs between HLA-A alleles some non-HLA-A*24 alleles will be amplified by primer mixes 1 to 7, 9 to 13, 15 to 17, 19, 20, 23 to 25, 27, 29 to 32, 36, 39, 41, 42, 44 to 48, 56, 58, 62, 64 to 66, 71, 73, 77 to 79, 85, 87, 89 and 91. In addition, the B*07:64, B*15:12 and B*15:19 alleles will be amplified by primer mixes 32w, 45 and 91, the B*18:27 allele will be amplified by mixes 2, 6, 25, 30 and 48, the B*18:38 and B*40:197 alleles will be amplified by primer mix 63, the B*44:136, B*51:136 and C*07:204 alleles will be amplified by primer mix 46, the B*51:129 allele will be amplified by primer mixes 89 and 93 and the C*12:37 allele will be amplified by primer mixes 46 and 57.

5HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

6Primer mixes 4, 6 and 48 may give rise to a lower yield of HLA-specific PCR product than the other A*24 primer mixes.

7Primer mixes 20, 30, 41 and 75 have a tendency to giving rise to primer oligomer formation.

8Primer mixes 1, 11, 19, 20, 27, 28, 36, 52, 54, 64, 76, 91 and 93 may have tendencies of unspecific amplifications, most pronounced in primer mix 91.

9Primer mix 11: Specific PCR fragment of 200 bp in the A*24:08, 24:42 and 24:89 and the A*02:76:01-02:76:02, 02:319 and 30:22 alleles. Specific PCR fragment of 245 bp in the A*24:61 and the 23:38N alleles.

10Primer mix 15: Specific PCR fragment of 105 bp in the A*24:19, 24:44, 24:100 and 24:109 alleles. Specific PCR fragment of 385 bp in the A*24:14-24:15, 24:51, 24:53, 24:57 and 24:92 and the A*02:46, 02:48, 02:70, 02:129, 03:30, 30:07, 31:08, 32:05, 33:21 and 33:53 alleles.

11Primer mix 24: Specific PCR fragment of 215 bp in the A*24:24, 24:67, 24:145, 24:156 and 24:191 and the A*03:72, 23:51, 26:16, 29:37, 30:02:01-30:04:02, 30:06, 30:09-30:10, 30:12, 30:25, 30:27-30:29, 30:32-30:34, 30:45, 30:47, 30:50-30:51, 30:57, 30:61, 30:64, 30:66 and 32:07 alleles.

Specific PCR fragment of 240 bp in the A*24:19, 24:28, 24:44 and 24:89 and the A*02:76:01-02:76:02, 02:319, 03:72, 30:31 and 30:60 alleles. Specific PCR fragment of 215 and 245 bp in the A*11:88, 30:01:01-30:01:07, 30:11:01-30:11:02, 30:13-30:20, 30:23-30:24, 30:26, 30:30, 30:35-30:44, 30:46, 30:48-30:49, 30:52-30:54, 30:56, 30:58-30:59N, 30:62-30:63, 30:65 and 68:45 alleles.

12Primer mix 26: Specific PCR fragment of 90 bp in the A*24:96 allele. Specific PCR fragment of  220 bp in the A*24:22 and 24:160 alleles.

13Primer mix 27: Specific PCR fragment of 80 bp in the A*24:32 and the A*23:36 and 32:05 alleles. Specific PCR fragment of 370 bp in the A*24:14 and in the A*24:14 and the A*02:46, 02:48, 02:70 and 02:129 alleles.

14Primer mix 28: Specific PCR fragment of 75 bp in the A*24:155N allele. Specific PCR fragment of 155 bp in the A*24:35 allele.

15Primer mix 29: Specific PCR fragment of 75 bp in the A*24:155N allele. Specific PCR fragment of 210 bp in the A*24:08, 24:29, 24:42 and 24:89 and the A*02:76:01-02:76:02, 02:319 and 30:22 alleles.

16Primer mix 30: Specific PCR fragment of 125 bp in the A*24:21:01-24:21:02 and 24:208 and in  the A*23:03:01-23:03:02, 29:07 and 31:29 and the B*18:27 alleles. Specific PCR fragment of 200 bp in the A*24:59 and 24:190 alleles.

17Primer mix 31: Specific PCR fragment of 155 bp in the A*24:37 allele. Specific PCR fragment of 235 bp in the A*24:04, 24:109, 24:119 and 24:129 and the A*01:02, 01:20, 26:16, 29:37 and 30:57 alleles.

18Primer mix 38: Specific PCR fragment of 145 bp in the A*24:49 allele. Specific PCR fragment of 170 bp in the A*24:137 allele. Specific PCR fragment of 230 bp in the A*24:126 allele.

19Primer mix 39: Specific PCR fragment of 75 bp in the A*24:185N allele. Specific PCR fragment of 210 bp in the A*24:50 and the A*02:129 alleles.

20Primer mix 41: Specific PCR fragment of 200 bp in the A*24:26 and 24:62 and the A*01:39, 23:45, 26:29, 26:49, 31:05, 32:13, 33:10 and 66:10 alleles. Specific PCR fragment of 240 bp in the A*24:118 and 24:132N alleles.

21Primer mix 42: Specific PCR fragment of 170 bp in the A*24:22 and 24:160 and the A*02:262 alleles. Specific PCR fragment of 200 bp in the A*24:24 and 24:67 and the A*26:16 and 68:45 alleles.

22Primer mix 45: Specific PCR fragment of 200 bp in the A*24:113 allele. Specific PCR fragment of 245 bp in the A*24:14, 24:52 and 24:132N and in the A*01:20, 01:66, 01:73, 02:19, 02:36-02:37, 02:54, 02:255, 26:29, 26:49 and 66:10 and the B*07:64, B*15:12 and B*15:19 alleles.

23Primer mix 47: Specific PCR fragment of 200 bp in the A*24:18, 24:56, 24:123 and 24:177 alleles. Specific PCR fragment of 410 bp in the A*24:26 and the A*03:30, 30:07, 31:08, 32:05 and 33:53 alleles.

24Primer mix 49: Specific PCR fragment of 165 bp in the A*24:159 allele. Specific PCR fragment of 215 bp in the A*24:91 allele. Specific PCR fragment of 245 bp in the A*24:23 allele.

25Primer mix 50: Specific PCR fragment of 235 bp in the A*24:34 allele. Specific PCR fragment of 260 bp in the A*24:163N allele.

26Primer mix 51: Specific PCR fragment of 165 bp in the A*24:140 allele. Specific PCR fragment of 215 bp in the A*24:38 allele.

27Primer mix 53: Specific PCR fragment of 125 bp in the A*24:158N allele. Specific PCR fragment of 230 bp in the A*24:55 allele.

28Primer mix 55: Specific PCR fragment of 240 bp in the A*24:60N allele. Specific PCR fragment of 275 bp in the A*24:173 allele.

29Primer mix 59: Specific PCR fragment of 230 bp in the A*24:143 allele. Specific PCR fragment of 255 bp in the A*24:68 allele. Specific PCR fragment of 285 bp in the A*24:127 allele.

30Primer mix 60: Specific PCR fragment of 205 bp in the A*24:183N allele. Specific PCR fragment of 275 bp in the A*24:69 allele.

31Primer mix 61: Specific PCR fragment of 105 bp in the A*24:70 allele. Specific PCR fragment of 280 bp in the A*24:127 and 24:141 alleles.

32Primer mix 64: Specific PCR fragment of 115 bp in the A*24:124 and in the A*31:07 and 31:10 alleles. Specific PCR fragment of 170 bp in the A*24:74 allele.

33Primer mix 65: Specific PCR fragment of 80 bp in the A*24:75 allele. Specific PCR fragment of 235 bp in the A*24:90N and the A*02:88N alleles.

34Primer mix 68: Specific PCR fragment of 230 bp in the A*24:143 allele. Specific PCR fragment of 285 bp in the A*24:78 and 24:141 alleles.

35Primer mix 69: Specific PCR fragment of 105 bp in the A*24:80 allele. Specific PCR fragment of 265 bp in the A*24:184 allele.

36Primer mix 70: Specific PCR fragment of 110 bp in the A*24:134 allele. Specific PCR fragment of 250 bp in the A*24:14w, 24:81 and 24:93w and the A*02:19w, 02:36-02:37w, 02:54w and 02:255w alleles.

37Primer mix 71: Specific PCR fragment of 130 bp in the A*24:128 and the A*23:29 alleles. Specific PCR fragment of 205 bp in the A*24:183N allele. Specific PCR fragment of 260 bp in the A*24:82 and the A*29:13, 31:07-31:08, 31:10, 32:42 alleles.

38Primer mix 73: Specific PCR fragment of 100 bp in the A*24:84N allele. Specific PCR fragment of 170 bp in the A*24:131, 24:138 and 24:188 and the A*02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300 and 02:303 alleles.

39Primer mix 74: Specific PCR fragment of 145 bp in the A*24:180 allele. Specific PCR fragment of 200 bp in the A*24:85 allele.

40Primer mix 75: Specific PCR fragment of 145 bp in the A*24:180 allele. Specific PCR fragment of 290 bp in the A*24:86N allele.

41Primer mix 76: Specific PCR fragment of 170 bp in the A*24:76 allele. Specific PCR fragment of 260 bp in the A*24:98 allele.

42Primer mix 77: Specific PCR fragment of 145 bp in the A*24:122 allele. Specific PCR fragment of 210 bp in the A*24:06, 24:87, 24:138, 24:167 and the A*23:02 alleles.

43Primer mix 79: Specific PCR fragment of 195 bp in the A*24:182 allele. Specific PCR fragment of 250 bp in the A*24:14, 24:93 and 24:114 and the A*01:20, 01:66, 02:19, 02:36-02:37, 02:54 and 02:255 alleles.

44Primer mix 80: Specific PCR fragment of 170 bp in the A*24:140 and 24:142 alleles. Specific PCR fragment of 220 bp in the A*24:95 allele.

45Primer mix 81: Specific PCR fragment of 100 bp in the A*24:139 allele. Specific PCR fragment of 215 bp in the A*24:97 and 24:133 alleles.

46Primer mix 82: Specific PCR fragment of 185 bp in the A*24:99 and 24:130 alleles. Specific PCR fragment of 225 bp in the A*24:125 allele.

47Primer mix 83: Specific PCR fragment of 135 bp in the A*24:111 allele. Specific PCR fragment of 275 bp in the A*24:101 and 24:173 alleles.

48Primer mix 84: Specific PCR fragment of 95 bp in the A*24:103 allele. Specific PCR fragment of 210 bp in the A*24:107 allele. Specific PCR fragment of 290 bp in the A*24:102 allele.

49Primer mix 85: Specific PCR fragment of 145 bp in the A*24:117 allele. Specific PCR fragment of 210 bp in the A*24:104 and in the A*03:30 and 32:05 alleles.

50Primer mix 86: Specific PCR fragment of 190 bp in the A*24:121 and 24:182 alleles. Specific PCR fragment of 255 bp in the A*24:105 allele.

51Primer mix 87: Specific PCR fragment of 90 bp in the A*24:17, 24:41, 24:62, 24:106 and 24:208 and the A*23:31, 23:45, 23:53, 29:07, 31:29 alleles. Specific PCR fragment of 120 bp in the A*24:110 allele. Specific PCR fragment of 195 bp in the A*24:136 alleles.

52Primer mix 88: Specific PCR fragment of 75 bp in the A**24:185N allele. Specific PCR fragment of 170 bp in the A*24:116 and 24:137 alleles.

53Primer mix 89: Specific PCR fragment of 75 bp in the A*24:07, 24:108, 24:112, 24:131 and the A*23:13 alleles. Specific PCR fragment of 270 bp in the A*24:112 and the B*51:129 alleles.

54Primer mix 90: Specific PCR fragment of 95 bp in the A*24:115 allele. Specific PCR fragment of 180 bp in the A*24:168 allele.

‘w’, might be weakly amplified.

1The sequence of the A*2401 has been shown to be in error.

The sequence of the A*2412 allele has been shown to be identical to A*24:08.

The A*2416 allele has been renamed to A*31:08.

The A*2465 allele has been renamed to A*24:13:02.

2HLA-A*24 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.9.0, July 2012.

3This lot of the HLA-A*24 subtyping kit cannot distinguish the A*24:169 allele and the A*24:02:01:01, 24:02:01:03-24:02:04, 24:02:06-24:02:09, 24:02:11-24:02:12 and  24:02:14-24:02:59 and the A*24:146-24:154, 24:161-24:162, 24:165-24:166, 24:170-24:172, 24:174-24:176, 24:178-24:179, 24:181, 24:186-24:187, 24:189, 24:192-24:193, 24:195-24:198, 24:200-24:203, 24:205-24:206 and 24:209 alleles.

4The A*24 primer set cannot separate the A*24:164 and A*33:19 alleles. These two alleles can be distinguished by the HLA-A low resolution and/or HLA-A*33 kits.

5The A*24:37 and 24:119 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 31.

6The A*24:49 and 24:126 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 38.

7The A*24:81 and 24:134 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 70.

8The A*24:82 and 24:128 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 71.

9The A*24:101 and 24:111 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 83.

10The A*24:102 and 24:103 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 84.

11The A*24:104 and 24:117 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 85.

12The A*24:105 and 24:121 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 86.

13The A*24:106 and 24:136 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 87.

14Primer mix 11: Specific PCR fragment of 200 bp in the A*24:08, 24:42 and 24:89 and the A*02:76:01-02:76:02, 02:319 and 30:22 alleles. Specific PCR fragment of 245 bp in the A*24:61 and the 23:38N alleles.

Primer mix 15: Specific PCR fragment of 105 bp in the A*24:19, 24:44, 24:100 and 24:109 alleles. Specific PCR fragment of 385 bp in the A*24:14-24:15, 24:51, 24:53, 24:57 and 24:92 and the A*02:46, 02:48, 02:70, 02:129, 03:30, 30:07, 31:08, 32:05, 33:21 and 33:53 alleles.

Primer mix 24: Specific PCR fragment of 215 bp in the A*24:24, 24:67, 24:145, 24:156 and 24:191 and the A*03:72, 23:51, 26:16, 29:37, 30:02:01-30:04:02, 30:06, 30:09-30:10, 30:12, 30:25, 30:27-30:29, 30:32-30:34, 30:45, 30:47, 30:50-30:51, 30:57, 30:61, 30:64, 30:66 and 32:07 alleles.

Specific PCR fragment of 240 bp in the A*24:19, 24:28, 24:44 and 24:89 and the A*02:76:01-02:76:02, 02:319, 03:72, 30:31 and 30:60 alleles. Specific PCR fragment of 215 and 245 bp in the A*11:88, 30:01:01-30:01:07, 30:11:01-30:11:02, 30:13-30:20, 30:23-30:24, 30:26, 30:30, 30:35-30:44, 30:46, 30:48-30:49, 30:52-30:54, 30:56, 30:58-30:59N, 30:62-30:63, 30:65 and 68:45 alleles.

Primer mix 26: Specific PCR fragment of 90 bp in the A*24:96 allele. Specific PCR fragment of  220 bp in the A*24:22 and 24:160 alleles.

Primer mix 27: Specific PCR fragment of 80 bp in the A*24:32 and the A*23:36 and 32:05 alleles. Specific PCR fragment of 370 bp in the A*24:14 and in the A*24:14 and the A*02:46, 02:48, 02:70 and 02:129 alleles.

Primer mix 28: Specific PCR fragment of 75 bp in the A*24:155N allele. Specific PCR fragment of 155 bp in the A*24:35 allele.

Primer mix 29: Specific PCR fragment of 75 bp in the A*24:155N allele. Specific PCR fragment of 210 bp in the A*24:08, 24:29, 24:42 and 24:89 and the A*02:76:01-02:76:02, 02:319 and 30:22 alleles.

Primer mix 30: Specific PCR fragment of 125 bp in the A*24:21:01-24:21:02 and 24:208 and in  the A*23:03:01-23:03:02, 29:07 and 31:29 and the B*18:27 alleles. Specific PCR fragment of 200 bp in the A*24:59 and 24:190 alleles.

Primer mix 31: Specific PCR fragment of 155 bp in the A*24:37 allele. Specific PCR fragment of 235 bp in the A*24:04, 24:109, 24:119 and 24:129 and the A*01:02, 01:20, 26:16, 29:37 and 30:57 alleles.

Primer mix 38: Specific PCR fragment of 145 bp in the A*24:49 allele. Specific PCR fragment of 170 bp in the A*24:137 allele. Specific PCR fragment of 230 bp in the A*24:126 allele.

Primer mix 39: Specific PCR fragment of 75 bp in the A*24:185N allele. Specific PCR fragment of 210 bp in the A*24:50 and the A*02:129 alleles.

Primer mix 41: Specific PCR fragment of 200 bp in the A*24:26 and 24:62 and the A*01:39, 23:45, 26:29, 26:49, 31:05, 32:13, 33:10 and 66:10 alleles. Specific PCR fragment of 240 bp in the A*24:118 and 24:132N alleles.

Primer mix 42: Specific PCR fragment of 170 bp in the A*24:22 and 24:160 and the A*02:262 alleles. Specific PCR fragment of 200 bp in the A*24:24 and 24:67 and the A*26:16 and 68:45 alleles.

Primer mix 45: Specific PCR fragment of 200 bp in the A*24:113 allele. Specific PCR fragment of 245 bp in the A*24:14, 24:52 and 24:132N and in the A*01:20, 01:66, 01:73, 02:19, 02:36-02:37, 02:54, 02:255, 26:29, 26:49 and 66:10 and the B*07:64, B*15:12 and B*15:19 alleles.

Primer mix 47: Specific PCR fragment of 200 bp in the A*24:18, 24:56, 24:123 and 24:177 alleles. Specific PCR fragment of 410 bp in the A*24:26 and the A*03:30, 30:07, 31:08, 32:05 and 33:53 alleles.

Primer mix 49: Specific PCR fragment of 165 bp in the A*24:159 allele. Specific PCR fragment of 215 bp in the A*24:91 allele. Specific PCR fragment of 245 bp in the A*24:23 allele.

Primer mix 50: Specific PCR fragment of 235 bp in the A*24:34 allele. Specific PCR fragment of 260 bp in the A*24:163N allele.

Primer mix 51: Specific PCR fragment of 165 bp in the A*24:140 allele. Specific PCR fragment of 215 bp in the A*24:38 allele.

Primer mix 53: Specific PCR fragment of 125 bp in the A*24:158N allele. Specific PCR fragment of 230 bp in the A*24:55 allele.

Primer mix 55: Specific PCR fragment of 240 bp in the A*24:60N allele. Specific PCR fragment of 275 bp in the A*24:173 allele.

Primer mix 59: Specific PCR fragment of 230 bp in the A*24:143 allele. Specific PCR fragment of 255 bp in the A*24:68 allele. Specific PCR fragment of 285 bp in the A*24:127 allele.

Primer mix 60: Specific PCR fragment of 205 bp in the A*24:183N allele. Specific PCR fragment of 275 bp in the A*24:69 allele.

Primer mix 61: Specific PCR fragment of 105 bp in the A*24:70 allele. Specific PCR fragment of 280 bp in the A*24:127 and 24:141 alleles.

Primer mix 64: Specific PCR fragment of 115 bp in the A*24:124 and in the A*31:07 and 31:10 alleles. Specific PCR fragment of 170 bp in the A*24:74 allele.

Primer mix 65: Specific PCR fragment of 80 bp in the A*24:75 allele. Specific PCR fragment of 235 bp in the A*24:90N and the A*02:88N alleles.

Primer mix 68: Specific PCR fragment of 230 bp in the A*24:143 allele. Specific PCR fragment of 285 bp in the A*24:78 and 24:141 alleles.

Primer mix 69: Specific PCR fragment of 105 bp in the A*24:80 allele. Specific PCR fragment of 265 bp in the A*24:184 allele.

Primer mix 70: Specific PCR fragment of 110 bp in the A*24:134 allele. Specific PCR fragment of 250 bp in the A*24:14w, 24:81 and 24:93w and the A*02:19w, 02:36-02:37w, 02:54w and 02:255w alleles.

Primer mix 71: Specific PCR fragment of 130 bp in the A*24:128 and the A*23:29 alleles. Specific PCR fragment of 205 bp in the A*24:183N allele. Specific PCR fragment of 260 bp in the A*24:82 and the A*29:13, 31:07-31:08, 31:10, 32:42 alleles.

Primer mix 73: Specific PCR fragment of 100 bp in the A*24:84N allele. Specific PCR fragment of 170 bp in the A*24:131, 24:138 and 24:188 and the A*02:17:01-02:17:02, 02:108, 02:110, 02:268, 02:300 and 02:303 alleles.

Primer mix 74: Specific PCR fragment of 145 bp in the A*24:180 allele. Specific PCR fragment of 200 bp in the A*24:85 allele.

Primer mix 75: Specific PCR fragment of 145 bp in the A*24:180 allele. Specific PCR fragment of 290 bp in the A*24:86N allele.

Primer mix 76: Specific PCR fragment of 170 bp in the A*24:76 allele. Specific PCR fragment of 260 bp in the A*24:98 allele.

Primer mix 77: Specific PCR fragment of 145 bp in the A*24:122 allele. Specific PCR fragment of 210 bp in the A*24:06, 24:87, 24:138, 24:167 and the A*23:02 alleles.

Primer mix 79: Specific PCR fragment of 195 bp in the A*24:182 allele. Specific PCR fragment of 250 bp in the A*24:14, 24:93 and 24:114 and the A*01:20, 01:66, 02:19, 02:36-02:37, 02:54 and 02:255 alleles.

Primer mix 80: Specific PCR fragment of 170 bp in the A*24:140 and 24:142 alleles. Specific PCR fragment of 220 bp in the A*24:95 allele.

Primer mix 81: Specific PCR fragment of 100 bp in the A*24:139 allele. Specific PCR fragment of 215 bp in the A*24:97 and 24:133 alleles.

Primer mix 82: Specific PCR fragment of 185 bp in the A*24:99 and 24:130 alleles. Specific PCR fragment of 225 bp in the A*24:125 allele.

Primer mix 83: Specific PCR fragment of 135 bp in the A*24:111 allele. Specific PCR fragment of 275 bp in the A*24:101 and 24:173 alleles.

Primer mix 84: Specific PCR fragment of 95 bp in the A*24:103 allele. Specific PCR fragment of 210 bp in the A*24:107 allele. Specific PCR fragment of 290 bp in the A*24:102 allele.

Primer mix 85: Specific PCR fragment of 145 bp in the A*24:117 allele. Specific PCR fragment of 210 bp in the A*24:104 and in the A*03:30 and 32:05 alleles.

Primer mix 86: Specific PCR fragment of 190 bp in the A*24:121 and 24:182 alleles. Specific PCR fragment of 255 bp in the A*24:105 allele.

Primer mix 87: Specific PCR fragment of 90 bp in the A*24:17, 24:41, 24:62, 24:106 and 24:208 and the A*23:31, 23:45, 23:53, 29:07, 31:29 alleles. Specific PCR fragment of 120 bp in the A*24:110 allele. Specific PCR fragment of 195 bp in the A*24:136 alleles.

Primer mix 88: Specific PCR fragment of 75 bp in the A**24:185N allele. Specific PCR fragment of 170 bp in the A*24:116 and 24:137 alleles.

Primer mix 89: Specific PCR fragment of 75 bp in the A*24:07, 24:108, 24:112, 24:131 and the A*23:13 alleles. Specific PCR fragment of 270 bp in the A*24:112 and the B*51:129 alleles.

Primer mix 90: Specific PCR fragment of 95 bp in the A*24:115 allele. Specific PCR fragment of 180 bp in the A*24:168 allele.

‘w’, might be weakly amplified.

Change in revision R01 compared to R00:

- The A*24:14, 24:93 and the A*02:19, 02:36-02:37, 02:54 and 02:255 alleles are weakly amplified in primer mix 70. This has been corrected in the Specificity and Interpretation Tables.