PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*25

Prod. No:	101. 423-06/06u

Lot No:    	37S

Expiry Date:    	2015-November-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-A*25 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*25 alleles | Other amplified HLA-A alleles3 |
| --- | --- | --- | --- | --- |
| 15 | 180 bp | 800 bp | *25:01:01-25:05, 25:07-25:17, 25:19-25:22 | *02:135, 26:01:01-26:01:20, 26:01:22-26:03:02, 26:05-26:08, 26:10-26:28, 26:29w, 26:30-26:33, 26:35-26:43:02, 26:45-26:48, 26:49w, 26:50-26:72, 26:74-26:77, 26:79-26:84, 43:01, 66:01, 66:04-66:09, 66:10w, 66:11-66:15, 66:17-66:19 |
| 24,6,7 | 75 bp | 800 bp | *25:01:01-25:14, 25:16, 25:18-25:22 | *32:01:01-32:01:13, 32:01:15-32:02, 32:04, 32:06-32:37, 32:39-32:57 |
| 34,6 | 100 bp | 800 bp | *25:01:01-25:01:07, 25:03-25:12N, 25:14, 25:16, 25:18-25:22 | *01:01:01:01-01:01:27, 01:01:29-01:04N, 01:06, 01:08-01:12, 01:14-01:16N, 01:18N-01:33, 01:35-01:70, 01:72-01:99, 01:101-01:104, 01:106-01:126, 02:346, 03:41, 11:17, 11:19, 11:40, 11:98, 11:121, 24:44w, 24:109w, 26:01:01-26:01:21, 26:01:23-26:12, 26:14-26:18, 26:20-26:29, 26:31-26:32, 26:34-26:43:02, 26:45-26:63, 26:66-26:82, 26:84, 31:03, 33:13, 36:01-36:05, 43:01, 66:05, 66:15, 74:10, 80:01:01:01-80:02 |
| 44 | 100 bp | 800 bp | *25:02 | *01:13, 01:17, 03:63, 03:88, 11:01:01-11:11, 11:13-11:16, 11:20-11:27, 11:29-11:39, 11:41-11:52Q, 11:54-11:95, 11:97, 11:99N-11:105, 11:107-11:120, 11:122-11:133, 11:135-11:139, 26:13, 26:19, 26:33, 34:01:01-34:06, 34:08, 34:10N, 66:01, 66:04, 66:06-66:11, 66:13-66:14, 66:17-66:19 |
| 5 | 425 bp | 1070 bp | *25:02 | *02:34-02:35:03, 02:56:01-02:56:02, 02:62, 02:78, 02:103, 23:13, 24:07, 24:19, 24:24, 24:112, 24:131, 26:13, 26:19, 34:01:01-34:10N, 66:01-66:02, 66:04, 66:06-66:14, 66:16-66:19, 68:01:01:01-68:02:05, 68:06-68:14, 68:16-68:19, 68:21:01-68:30, 68:32-68:56, 68:58-68:89, 68:91-68:99, 69:01 |
| 6 | 200 bp | 1070 bp | *25:03 | *01:51, 02:55, 03:24, 26:20, 32:15, 34:08, 68:71 |
| 7 | 165 bp | 800 bp | *25:04, 25:06 | *02:309, 03:01:19, 26:08-26:09, 26:48, 26:67, 26:73, 26:81, 31:03-31:04, 34:01:01-34:09 |
| 84 | 125 bp | 1070 bp | *25:05 | *26:54 |
| 9 | 135 bp | 1070 bp | *25:08 | *26:47, 66:18 |
| 10 | 170 bp | 1070 bp | *25:07 |  |
| 116 | 135 bp | 1070 bp | *25:09 | *03:01:19, 26:14, 26:18, 26:28, 26:73, 31:03-31:04, 34:03, 34:06 |
| 124 | 75 bp | 1070 bp | *25:10 | *02:309, 03:01:19, 31:03-31:04, 34:02:01-34:04, 34:07-34:10N |
| 134 | 65 bp | 1070 bp | *25:11 | *23:12, 24:30, 24:42, 32:08 |
| 14 | 165 bp | 800 bp | *25:12N |  |
| 15 | 360 bp | 1070 bp | *25:13 | *02:135, 02:309, 03:01:19, 26:30, 26:65, 31:04, 34:09, 66:02-66:03, 66:12, 66:16 |
| 164 | 80 bp | 1070 bp | *25:01:01, 25:01:03-25:12N, 25:14-25:22 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*25 SSP typings.

When the primers in a primer mix can give rise to HLA-specific PCR products of more than one length this is indicated if the size difference is more than 20 base pairs. Size differences of 20 base pairs or less are not given. For high resolution SSP kits the respective lengths of the HLA-specific PCR product(s) are given for the alleles amplified by these primer mixes.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*25 subtyping.

In addition, wells number 2, 3, 4, 7 and 14 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*25 alleles will be amplified by the A*25 primer mixes 1 to 9, 11 to 13 and 15.

4HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

5Primer mix 1 may give rise to a PCR fragment approx. 500 bp in size. This band should be disregarded in the interpretation of HLA-A*25 subtypings.

6Primer mixes 2, 3 and 11 may have tendencies of unspecific amplifications.

7Primer mix 2 may give rise to a lower yield of HLA-specific PCR product than the other A*25 primer mixes.

‘w’, may be weakly amplified.

1HLA-A*25 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.11.0, January 2013.

‘w’, may be weakly amplified.