Gel Documentation Form and Worksheet

HLA-A*25 (101.423-06/06u) 	Lot No: 50X	Expiry Date: 2017-July-01

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:________________

Reviewed By:________________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 1 may give rise to a PCR fragment approx. 500 bp in size. This band should be disregarded in the interpretation of HLA-A*25 subtypings.

Primer mixes 2, 3 and 11 may have tendencies of unspecific amplifications.

Primer mix 2 may give rise to a lower yield of HLA-specific PCR product than the other A*25 primer mixes.

Primer mix 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-A*25 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.18.0, October 2014.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 6: Specific PCR fragment of 155 bp in the A*25:19 and the A*02:309, 02:454, 03:01:19, 26:43:01-26:43:02, 31:03-31:04, 34:02:01-34:04, 34:06-34:07, 34:09 and 66:06 alleles. Specific PCR fragment of 200 bp in the A*25:03 and the A*01:51, 02:55, 03:24, 26:20, 32:15 and 68:71 alleles. Specific PCR fragment of 155 bp and 200 bp in the A*34:08 allele.

Primer mix 8: Specific PCR fragment of 100 bp in the A*25:16 allele. Specific PCR fragment of 125 bp in the A*25:05 and the A*02:454 and 26:54 alleles. Specific PCR fragment of 165 bp in the A*25:12N allele.

Primer mix 9: Specific PCR fragment of 100 bp in the A*25:16 allele. Specific PCR fragment of 145 bp in the A*25:08, and 25:19 and the A*02:309, 02:454, 03:01:19, 26:43:01-26:43:02, 26:47, 31:03-31:04, 34:02:01-34:04, 34:06-34:09, 66:06 and 66:18 alleles.

4The HLA-A*25 primer set cannot separate the A*25:18 and A*32:62 alleles. These alleles can be distinguished by the HLA-A low resolution kit and/or the HLA-A*32 high resolution kit.

The HLA-A*25 subtyping kit cannot distinguish the following silent mutations: the A*25:01:01 and 25:01:03-25:01:07alleles.

‘w’, might be weakly amplified.