PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*26

Prod. No:	101.424-12/12u

Lot No:    	40F

Expiry Date:    	2011-January-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 30 primer mixes used for HLA-A*26 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*26 alleles | Other amplified HLA-A alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 180 bp | 800 bp | 260101-260106, 2603, 2605, 260701-2608, 2610-2628, 2630-2633, 2635-2637 | 9235, 250101-2505, 4301, 6601, 6604-6608 |
| 24 | 80 bp | 800 bp | 260101-2602, 2604, 2608-2618, 2620, 2622-2629, 2632-2637 | 3313 |
| 3 | 140 bp | 1070 bp | 2602 |  |
| 4 | 255 bp | 800 bp | 2603, 2606, 2621, 2630 | 020102, 020605, 0250, 9222, 0309, 1106, 1118, 2428, 2489, 3013, 3016, 6805, 6815, 6820, 7406 |
| 5 | 180 bp | 1070 bp | 2604 |  |
| 66 | 80, 160 bp | 1070 bp | 2605, 2627 |  |
| 7 | 150 bp | 800 bp | 2608, 2628 | 2504 |
| 8 | 135 bp | 1070 bp | 2611N, 2614, 2618, 2628 | 3103, 3104, 3403, 3406 |
| 97 | 145, 190 bp | 1070 bp | 2612, 2618, 2629 | 3103 weakly, 3104 weakly, 3406 weakly, 6606 weakly |
| 104,8 | 85, 260 bp | 1070 bp | 260701-260702, 2620 | 9246 |
| 11 | 140 bp | 800 bp | 2606, 2619 | 3103 |
| 129 | 135, 240 bp | 1070 bp | 2614, 2616, 2618, 2628 | 0102, 0120, 2404, 3103, 3104, 3403, 3406 |
| 13 | 255 bp | 1070 bp | 2617 |  |
| 14 | 145 bp | 1070 bp | 2610 | 0238, 9201, 2310, 2410, 2446 |
| 154 | 100 bp | 800 bp | 2613, 2619, 2633 | 0113, 0117, 110101-1111, 1113-1116, 1120-1127, 1129-1137, 2502, 340101-3406, 3408, 6601, 6604, 6606-6608 |
| 16 | 430 bp | 1070 bp | 2613, 2619, 2633 | 0234-023503, 025601-025602, 0262, 0278, 9203, 2419, 340101-3408, 6601, 6602, 6604, 6606-6608, 680101-680202, 6806-6814, 6816-6819, 6821-6830, 6832-6835, 6837-6843, 6901 |
| 17 | 175 bp | 1070 bp | 2609 | 2506, 3103, 3104, 340101-3408 |
| 184,13 | 125, 205 bp | 800 bp | 2621, 2631 |  |
| 19 | 245 bp | 800 bp | 2622 | 0238, 9201 |
| 20 | 210 bp | 1070 bp | 2623 |  |
| 214 | 115 bp | 1070 bp | 2624 |  |
| 224 | 100 bp | 1070 bp | 2625N |  |
| 23 | 165 bp | 1070 bp | 2626 |  |
| 24 | 305 bp | 1070 bp | 2632 | 0103, 1126, 3313, 3603, 7410 |
| 25 | 360 bp | 1070 bp | 2630 | 9235, 3104, 6602, 6603 |
| 26 | 175 bp | 1070 bp | 2634 |  |
| 27 | 275 bp | 1070 bp | 2635 |  |
| 284,5 | 90 bp | 1070 bp | 2603, 2606, 2621, 2636 | 1106, 8001 |
| 29 | 330 bp | 1070 bp | 2637 |  |
| 304,11 | 125, 235 bp | 1070 bp | 2615, 2629, 2633 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*26 subtyping.

In addition, wells number 2, 4, 7, 11, 15, 18 and 19 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*26 alleles will be amplified by primer mixes 1, 4, 7, 8 to 12 and 14 to 17, 19, 24, 25 and 28.

4Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

5Primer mix 28 may yield less intense specific PCR products than the other A*26 primer mixes.

6Primer mix 6: Specific PCR fragment of 80 bp in the A*2605 allele. Specific PCR fragment of 160 bp in the A*2627 allele.

7Primer mix 9: Specific PCR fragment of 145 bp in the A*2612 and A*2618 and weakly in the A*3103, A*3104, A*3406 and A*6606 alleles. Specific PCR fragment of 190 bp in the A*2629 allele.

8Primer mix 10: Specific PCR fragment of 85 bp in the A*260701-260702 alleles. Specific PCR fragment of 260 bp in the A*2620 and in the A*9246 allele.

9Primer mix 12: Specific PCR fragment of 135 bp in the A*2614, A*2618 and A*2628 and the A*3103, A*3104, A*3403 and A*3406 alleles. Specific PCR fragment of 240 bp in the A*2616 and the A*0102,  A*0120 and A*2404 alleles.

10Primer mix 18: Specific PCR fragment of 125 bp in the A*2631 allele. Specific PCR fragment of 205 bp in the A*2621 allele.

11Primer mix 30: Specific PCR fragment of 125 bp in the A*2615 and A*2629 allele. Specific PCR fragment of 235 bp in the A*2633 allele.

1Primer mix 6: Specific PCR fragment of 80 bp in the A*2605 allele. Specific PCR fragment of 160 bp in the A*2627 allele.

Primer mix 9: Specific PCR fragment of 145 bp in the A*2612 and A*2618 and weakly in the A*3103, A*3104, A*3406 and A*6606 alleles. Specific PCR fragment of 190 bp in the A*2629 allele.

Primer mix 10: Specific PCR fragment of 85 bp in the A*260701-260702 alleles. Specific PCR fragment of 260 bp in the A*2620 and in the A*9246 allele.

Primer mix 12: Specific PCR fragment of 135 bp in the A*2614, A*2618 and A*2628 and the A*3103, A*3104, A*3403 and A*3406 alleles. Specific PCR fragment of 240 bp in the A*2616 and the A*0102,  A*0120 and A*2404 alleles.

Primer mix 18: Specific PCR fragment of 125 bp in the A*2631 allele. Specific PCR fragment of 205 bp in the A*2621 allele.

Primer mix 30: Specific PCR fragment of 125 bp in the A*2615 and A*2629 allele. Specific PCR fragment of 235 bp in the A*2633 allele.