PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*26

Prod. No:	101.424-12/12u

Lot No:    	13K

Expiry Date:    	2012-June-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32primer mixes used for HLA-A*26 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*26 alleles3 | Other amplified HLA-A alleles4 |
| --- | --- | --- | --- | --- |
| 1 | 180 bp | 800 bp | *26:01:01-26:01:13, 26:03:01-26:03:02, 26:05, 26:07:01-26:08, 26:10-26:28, 26:29w, 26:30-26:33, 26:35-26:39, 26:41-26:45 | *02:135, 25:01:01-25:05, 25:07-25:13, 31:24, 43:01, 66:01, 66:04-66:09, 66:10w, 66:11-66:13 |
| 25 | 80 bp | 800 bp | *26:01:01-26:02, 26:04, 26:08-26:18, 26:20, 26:22-26:29, 26:32-26:45 | *01:51, 33:13 |
| 3 | 140 bp | 1070 bp | *26:02 |  |
| 4 | 255 bp | 800 bp | *26:03:01-26:03:02, 26:06, 26:21, 26:30 | *03:09, 11:06, 11:18, 29:19, 30:13, 30:16, 33:24, 68:05, 68:15, 68:20, 74:06 |
| 5 | 180 bp | 1070 bp | *26:04 |  |
| 65,9 | 80 bp, 160 bp | 1070 bp | *26:05, 26:27 |  |
| 7 | 150 bp | 800 bp | *26:08, 26:28 | *25:04 |
| 8 | 135 bp | 1070 bp | *26:11N, 26:14, 26:18, 26:28 | *25:09, 31:03-31:04, 34:03, 34:06 |
| 910 | 145 bp, 190 bp | 1070 bp | *26:12, 26:18, 26:29 | *31:03-31:04, 34:06, 66:06, 66:10 |
| 105,11 | 85 bp , 260 bp | 1070 bp | *26:07:01-26:07:02, 26:20 | *01:01:13, 02:146 |
| 11 | 140 bp | 800 bp | *26:06, 26:19 | *31:03 |
| 1212 | 135 bp, 240 bp | 1070 bp | *26:14, 26:16, 26:18, 26:28 | *01:02, 01:20, 24:04, 25:09, 31:03-31:04, 34:03, 34:06 |
| 135,13 | 110 bp, 255 bp | 1070 bp | *26:17, 26:45 |  |
| 14 | 145 bp | 1070 bp | *26:10 | *02:38, 02:101, 02:154, 23:10, 24:10, 24:46 |
| 155 | 100 bp | 800 bp | *26:13, 26:19, 26:33 | *01:13, 01:17, 03:63, 11:01:01-11:11, 11:13-11:16, 11:20-11:27, 11:29-11:39, 11:41-11:59, 25:02, 34:01:01-34:06, 34:08, 66:01, 66:04, 66:06-66:11, 66:13 |
| 16 | 430 bp | 1070 bp | *26:13, 26:19, 26:33 | *02:34w, 02:35:01-02:35:03, 02:56:01w-02:56:02w, 02:62w, 02:78, 02:103w, 24:19, 34:01:01-34:08, 66:01-66:02, 66:04, 66:06-66:13, 68:01:01-68:02:04, 68:06-68:14, 68:16-68:19, 68:21-68:30, 68:32-68:35, 68:37-68:49N, 69:01 |
| 17 | 175 bp | 1070 bp | *26:09 | *25:06, 31:03-31:04, 34:01:01-34:08 |
| 185,6,14 | 125 bp, 205 bp | 1070 bp | *26:21, 26:31 |  |
| 1915 | 190 bp, 245 bp | 800 bp | *26:22, 26:40 | *01:20w, 02:38, 02:101, 66:09 |
| 207 | 210 bp | 1070 bp | *26:23 |  |
| 215,16 | 115 bp, 205 bp | 1070 bp | *26:24, 26:41 |  |
| 225,17 | 100 bp , 220 bp | 1070 bp | *26:25N, 26:38 |  |
| 23 | 165 bp | 1070 bp | *26:26 |  |
| 24 | 305 bp | 1070 bp | *26:32 | *01:03, 11:26, 33:13, 36:03, 74:10 |
| 258 | 360 bp | 1070 bp | *26:30 | *02:135, 25:13, 31:04, 66:02-66:03, 66:12 |
| 2618 | 150 bp, 175 bp | 1070 bp | *26:34, 26:42 |  |
| 27 | 275 bp | 1070 bp | *26:35 |  |
| 285 | 90 bp | 1070 bp | *26:03:01-26:03:02, 26:06, 26:21, 26:36 | *11:06, 25:11, 80:01 |
| 2919 | 155 bp, 330 bp | 1070 bp | *26:37, 26:44 |  |
| 305,20 | 125 bp, 235 bp | 1070 bp | *26:15, 26:29, 26:33 |  |
| 31 | 210 bp | 800 bp | *26:39 |  |
| 32 | 155 bp | 1070 bp | *26:43 | *31:03-31:04, 34:02-34:04, 34:06-34:08, 66:06 |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*26 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*26 subtyping.

In addition, wells number 2, 4, 7, 11, 15, 19 and 31 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3The HLA-A*26 subtyping kit cannot distinguish the A*26:01:01 to A*26:01:13, the A*26:03:01 and A*26:03:02 or the A*26:07:01 and A*26:07:02 alleles.

The A*26:17 and A*26:45 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 13.

The A*26:24 and A*26:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

The A*26:25N and A*26:38 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 22.

The A*26:37 and A*26:44 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 29.

4Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*26 alleles will be amplified by primer mixes 1, 2, 4, 7 to 12, 14 to 17, 19, 24, 25, 28 and 32.

5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Primer mix 18 may have tendencies of primer oligomer formation.

7Primer mix 20 may give rise to nonspecific amplification.

8Primer mix 25 may yield less specific PCR product than the other A*26 primer mixes.

9Primer mix 6: Specific PCR fragment of 80 bp in the A*26:05 allele. Specific PCR fragment of 160 bp in the A*26:27 allele.

10Primer mix 9: Specific PCR fragment of 145 bp in the A*26:12 and 26:18 and in the A*31:03-31:04, 34:06 and 66:06 alleles. Specific PCR fragment of 190 bp in the A*26:29 and in the A*66:10 alleles.

11Primer mix 10: Specific PCR fragment of 85 bp in the A*26:07:01-26:07:02 and in the A*02:146 alleles. Specific PCR fragment of 260 bp in the A*26:20 and in the A*01:01:13 and 02:146 alleles.

12Primer mix 12: Specific PCR fragment of 135 bp in the A*26:14, 26:18 and 26:28 and in the A*25:09, 31:03-31:04, 34:03 and 34:06 alleles. Specific PCR fragment of 240 bp in the A*26:16 and in the A*01:02, 01:20 and 24:04 alleles.

13Primer mix 13: Specific PCR fragment of 110 bp in the A*26:45 allele. Specific PCR fragment of 255 bp in the A*26:17 allele.

14Primer mix 18: Specific PCR fragment of 115 bp in the A*26:31 allele. Specific PCR fragment of 205 bp in the A*26:21 allele.

15Primer mix 19: Specific PCR fragment of 190 bp in the A*26:40 allele. Specific PCR fragment of 245 bp in the A*26:22 and in the A*01:20w, 02:38, 02:101 and 66:09 alleles.                                                                                                                                                                                                     16Primer mix 21: Specific PCR fragment of 115 bp in the A*26:24 allele. Specific PCR fragment of 205 bp in the A*26:41 allele.

17Primer mix 22: Specific PCR fragment of 100 bp in the A*26:25N allele. Specific PCR fragment of 220 bp in the A*26:38 allele.

18Primer mix 26: Specific PCR fragment of 150 bp in the A*26:42 allele. Specific PCR fragment of 175 bp in the A*26:34 allele.

19Primer mix 29: Specific PCR fragment of 155 bp in the A*26:44 allele. Specific PCR fragment of 330 bp in the A*26:37 allele.

20Primer mix 30: Specific PCR fragment of 125 bp in the A*26:15 and A*26:29 alleles. Specific PCR fragment of 235 bp in the A*26:33 allele.

‘w’, may be weakly amplified.

1The A*26:17 and A*26:45 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 13.

The A*26:24 and A*26:41 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 21.

The A*26:25N and A*26:38 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 22.

The A*26:37 and A*26:44 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 29.

2Primer mix 6: Specific PCR fragment of 80 bp in the A*26:05 allele. Specific PCR fragment of 160 bp in the A*26:27 allele.

Primer mix 9: Specific PCR fragment of 145 bp in the A*26:12 and 26:18 and in the A*31:03-31:04, 34:06 and 66:06 alleles. Specific PCR fragment of 190 bp in the A*26:29 and in the A*66:10 alleles.

Primer mix 10: Specific PCR fragment of 85 bp in the A*26:07:01-26:07:02 and in the A*02:146 alleles. Specific PCR fragment of 260 bp in the A*26:20 and in the A*01:01:13 and 02:146 alleles.

Primer mix 12: Specific PCR fragment of 135 bp in the A*26:14, 26:18 and 26:28 and in the A*25:09, 31:03-31:04, 34:03 and 34:06 alleles. Specific PCR fragment of 240 bp in the A*26:16 and in the A*01:02, 01:20 and 24:04 alleles.

Primer mix 13: Specific PCR fragment of 110 bp in the A*26:45 allele. Specific PCR fragment of 255 bp in the A*26:17 allele.

Primer mix 18: Specific PCR fragment of 115 bp in the A*26:31 allele. Specific PCR fragment of 205 bp in the A*26:21 allele.

Primer mix 19: Specific PCR fragment of 190 bp in the A*26:40 allele. Specific PCR fragment of 245 bp in the A*26:22 and in the A*01:20w, 02:38, 02:101 and 66:09 alleles.

Specific PCR fragment of 115 bp in the A*26:24 allele. Specific PCR fragment of 205 bp in the A*26:41 allele.

Primer mix 22: Specific PCR fragment of 100 bp in the A*26:25N allele. Specific PCR fragment of 220 bp in the A*26:38 allele.

Primer mix 26: Specific PCR fragment of 150 bp in the A*26:42 allele. Specific PCR fragment of 175 bp in the A*26:34 allele.

Primer mix 29: Specific PCR fragment of 155 bp in the A*26:44 allele. Specific PCR fragment of 330 bp in the A*26:37 allele.

Primer mix 30: Specific PCR fragment of 125 bp in the A*26:15 and A*26:29 alleles. Specific PCR fragment of 235 bp in the A*26:33 allele.

‘w’, may be weakly amplified.