PHOTO DOCUMENT

Olerup SSP®

Olerup SSP® HLA-A*29

Prod. No:	101.428-12/12u

Lot No:    	19K

Expiry Date:    	2012-June-01

Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 16 primer mixes used for HLA-A*29 SSP typing.

| Primer Mix | Size of spec. PCR product1 | Size of control band2 | Amplified HLA-A*29 alleles | Other amplified HLA-A alleles3 |
| --- | --- | --- | --- | --- |
| 1 | 480 bp | 800 bp | *29:01:01:01-29:01:02, 29:12, 29:15-29:17, 29:20 |  |
| 24 | 130 bp | 1070 bp | *29:01:01:02N |  |
| 3 | 440 bp | 800 bp | *29:02:01-29:02:03, 29:02:05-29:11, 29:13-29:14, 29:19, 29:21-29:22 |  |
| 4 | 165 bp | 800 bp | *29:03 | *23:03:01, 31:05, 32:13, 33:10 |
| 54,6,7 | 130 bp, 185 bp | 800 bp | *29:04, 29:22 | *02:149, 26:22, 33:22, 66:09 |
| 64,8 | 105 bp, 130 bp | 800 bp | *29:05, 29:19-29:20 | *31:24, 32:02, 32:22 |
| 74,5,9 | 105 bp, 210 bp | 1070 bp | *29:06, 29:19 | *32:12, B*08:01:07 |
| 84,6,10 | 85 bp, 260 bp | 800 bp | *29:07, 29:13 | *24:17, 24:41, 24:82, 31:07-31:08, 31:10 |
| 94,5,11 | 80 bp, 170 bp | 1070 bp | *29:08N, 29:16 | *03:27 |
| 104,12 | 95 bp, 170 bp | 800 bp | *29:09, 29:17 | *31:24 |
| 114,13 | 100 bp, 195 bp | 800 bp | *29:10, 29:14 |  |
| 124,5,6,14 | 80 bp, 185 bp | 800 bp | *29:02:04, 29:18, 29:22 | *02:149, 26:22, 32:01:01-32:03, 32:05-32:22, 33:22, 66:09, 74:01-74:13 |
| 134,5,15 | 85 bp, 165 bp | 1070 bp | *29:11-29:12 | *31:16 |
| 14 | 200 bp | 1070 bp | *29:01:01:01-29:13, 29:15-29:17, 29:19-29:22 |  |
| 15 | 240 bp | 1070 bp | *29:01:01:01-29:12, 29:14-29:18, 29:20-29:22 | *33:13 |
| 164,16 | 95 bp, 165 bp | 1070 bp | *29:15, 29:21 |  |

1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*29 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*29 subtyping.

In addition, wells number 1, 3 to 6, 8 and 10 to 12 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-A alleles a few non-HLA-A*29 alleles will be amplified by primer mixes 4 to 10, 12, 13 and 15. In addition, the B*08:01:07 allele will be amplified by primer mix 7.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mixes 7, 9, 12 and 13 may give rise to nonspecific amplifications.

6Primer mixes 5, 8 and 12 may give rise to a long, unspecific amplification product of approximately 600 bp. This should be disregarded when interpreting the A*29 typings.

7Primer mix 5: Specific PCR fragment of 130 bp in the A*29:04 allele. Specific PCR fragment of 185 bp in the A*29:22 and the A*02:149, 26:22, 33:22 and 66:09 alleles.

8Primer mix 6: Specific PCR fragment of 105 bp in the A*29:19 and 29:20 alleles. Specific PCR fragment of 130 bp in the A*29:05 and the 31:24, 32:02 and 32:22 alleles.

9Primer mix 7: Specific PCR fragment of 105 bp in the A*29:19 allele. Specific PCR fragment of 210 bp in the A*29:06 and in the in the A*32:12 and B*08:01:07 alleles.

10Primer mix 8: Specific PCR fragment of 85 bp in the A*29:07 and in the A*24:17 and 24:41 alleles. Specific PCR fragment of 260 bp in the A*29:13 and in the A*24:82, 31:07-31:08 and 31:10 alleles.

11Primer mix 9: Specific PCR fragment of 80 bp in the A*29:16 and in the A*03:27 alleles. Specific PCR fragment of 170 bp in the A*29:08N allele.

12Primer mix 10: Specific PCR fragment of 95 bp in the A*29:09 and in the A*31:24 alleles. Specific PCR fragment of 170 bp in the A*29:17 allele.

13Primer mix 11: Specific PCR fragment of 100 bp in the A*29:14 allele. Specific PCR fragment of 195 bp in the A*29:10 allele.

14Primer mix 12: Specific PCR fragment of 80 bp in the A*29:02:04 and 29:18 and in the A*32:01:01-32:03, 32:05-32:22 and 74:01-74:13 alleles. Specific PCR fragment of 185 bp in the A*29:22 and in the A*02:149, 26:22, 33:22 and 66:09 alleles.

15Primer mix 13: Specific PCR fragment of 85 bp in the A*29:11 allele. Specific PCR fragment of 165 bp in the A*29:12 and in the A*31:16 alleles.

16Primer mix 16: Specific PCR fragment of 95 bp in the A*29:15 allele. Specific PCR fragment of 165 bp in the A*29:21 allele.

1Primer mix 5: Specific PCR fragment of 130 bp in the A*29:04 allele. Specific PCR fragment of 185 bp in the A*29:22 and the A*02:149, 26:22, 33:22 and 66:09 alleles.

Primer mix 6: Specific PCR fragment of 105 bp in the A*29:19 and 29:20 alleles. Specific PCR fragment of 130 bp in the A*29:05 and the 31:24, 32:02 and 32:22 alleles.

Primer mix 7: Specific PCR fragment of 105 bp in the A*29:19 allele. Specific PCR fragment of 210 bp in the A*29:06 and in the in the A*32:12 and B*08:01:07 alleles.

Primer mix 8: Specific PCR fragment of 85 bp in the A*29:07 and in the A*24:17 and 24:41 alleles. Specific PCR fragment of 260 bp in the A*29:13 and in the A*24:82, 31:07-31:08 and 31:10 alleles.

Primer mix 9: Specific PCR fragment of 80 bp in the A*29:16 and in the A*03:27 alleles. Specific PCR fragment of 170 bp in the A*29:08N allele.

Primer mix 10: Specific PCR fragment of 95 bp in the A*29:09 and in the A*31:24 alleles. Specific PCR fragment of 170 bp in the A*29:17 allele.

Primer mix 11: Specific PCR fragment of 100 bp in the A*29:14 allele. Specific PCR fragment of 195 bp in the A*29:10 allele.

Primer mix 12: Specific PCR fragment of 80 bp in the A*29:02:04 and 29:18 and in the A*32:01:01-32:03, 32:05-32:22 and 74:01-74:13 alleles. Specific PCR fragment of 185 bp in the A*29:22 and in the A*02:149, 26:22, 33:22 and 66:09 alleles.

Primer mix 13: Specific PCR fragment of 85 bp in the A*29:11 allele. Specific PCR fragment of 165 bp in the A*29:12 and in the A*31:16 alleles.

Primer mix 16: Specific PCR fragment of 95 bp in the A*29:15 allele. Specific PCR fragment of 165 bp in the A*29:21 allele.